## STM Output
STM-03002 Rev M Electronic Data Collection form for Avive Connect Connect Sub-Assembly Release Tests and Inspection Data
Timestamp | Email Address | Work Order Number | General Instructions | Operator's Name | Serial Number | Serial Number Verification / Revision Verification | LCD Inspection | Speaker Test | GPS Satellite sign verification | Wifi Verification | "Device is synced to REALConnect” is displayed | Cellular icon appears and shows reception | Current (I) Value when wifi and cellular is connected | Verify the values are between 350 mA and 450 mA when wifi and cellular are connected | SIM sign verification | Current (I) value when device switched off | Verify the values are between 0.25 mA and 1 mA | Specify the Failure and Reason | Notes/Comments
2023-11-15 18:05 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190622 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-16 11:42 | stmstation2@avive.life | W1146 | Agree | Ireri Moreno | 2310190594 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.377 | Pass |  | n/a
2023-11-16 11:49 | stmstation1@avive.life | W1146 | Agree | Ireri Moreno | 2310190690 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-16 11:52 | stmstation2@avive.life | W1146 | Agree | Ireri Moreno | 2310190785 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-16 12:01 | stmstation2@avive.life | W1146 | Agree | Ireri Moreno | 2310190760 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-16 12:07 | stmstation1@avive.life | W1146 | Agree | Ireri Moreno | 2310190533 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-16 12:08 | stmstation2@avive.life | W1146 | Agree | Ireri Moreno | 2310190545 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-16 12:16 | stmstation1@avive.life | W1146 | Agree | Ireri Moreno | 2310190689 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-16 12:17 | stmstation2@avive.life | W1146 | Agree | Ireri Moreno | 2310190706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-16 12:24 | stmstation1@avive.life | W1146 | Agree | Ireri Moreno | 2310190536 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-16 12:25 | stmstation2@avive.life | W1146 | Agree | Ireri Moreno | 2310190714 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2023-11-16 12:32 | stmstation1@avive.life | W1146 | Agree | Ireri Moreno | 2310190753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-16 12:35 | stmstation2@avive.life | W1146 | Agree | Ireri Moreno | 2310190716 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-16 12:41 | stmstation1@avive.life | W1146 | Agree | Ireri Moreno | 2310190641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-16 12:42 | stmstation2@avive.life | W1146 | Agree | Ireri Moreno | 2310190707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-16 12:50 | stmstation1@avive.life | W1146 | Agree | Ireri Moreno | 2310190654 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-16 12:51 | stmstation2@avive.life | W1146 | Agree | Ireri Moreno | 2310190809 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-16 12:58 | stmstation1@avive.life | W1146 | Agree | Ireri Moreno | 2310190658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-16 13:14 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190657 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-16 13:15 | stmstation2@avive.life | W1146 | Agree | Ireri Moreno | 2310190662 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-16 13:38 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190679 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-16 15:57 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190752 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-16 16:02 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190807 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-16 18:28 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190825 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-16 18:36 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-16 18:41 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190595 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-16 18:44 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190610 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-16 18:49 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-16 18:50 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190615 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-16 18:55 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.365 | Pass |  | n/a
2023-11-16 18:57 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190608 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-16 19:03 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190929 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.365 | Pass |  | n/a
2023-11-16 19:05 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190550 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-16 19:46 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190599 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.365 | Pass |  | n/a
2023-11-16 19:49 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-16 19:58 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-16 20:00 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-17 09:14 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190607 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-17 09:41 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-17 09:48 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-17 09:53 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190937 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-17 09:56 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190927 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-17 10:06 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190637 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-17 10:10 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-17 10:22 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-17 10:24 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190718 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-17 10:25 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190499 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.365 | Pass |  | n/a
2023-11-17 10:27 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-17 10:32 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190561 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-17 10:35 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190720 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-17 10:59 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190616 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-17 11:02 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-17 11:04 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190830 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-17 11:10 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190534 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-17 11:23 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190805 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | N/A
2023-11-17 11:23 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190430 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.357 | Pass |  | N/A
2023-11-17 11:35 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190535 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-17 11:36 | stmstation4@avive.life | W1146 | Agree | Ananya Jain | 2310190911 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.365 | Pass |  | N/A
2023-11-17 11:36 | stmstation3@avive.life | W1146 | Agree | Ananya Jain | 2310190596 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.355 | Pass |  | N/A
2023-11-17 11:42 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190335 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-17 11:53 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190802 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-17 11:54 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190939 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-17 12:13 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190754 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-17 12:13 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190748 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-17 12:15 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190940 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.362 | Pass |  | n/a
2023-11-17 12:16 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190801 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-17 12:25 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190613 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-17 12:28 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190723 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-17 12:30 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190715 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.35 | Pass |  | N/A
2023-11-17 12:37 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190644 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.365 | Pass |  | n/a
2023-11-17 12:40 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190983 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-17 12:42 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190614 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-17 12:43 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190719 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-17 12:53 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-17 12:54 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190529 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.343 | Pass |  | n/a
2023-11-17 13:01 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190624 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.364 | Pass |  | n/a
2023-11-17 14:32 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190497 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-17 14:42 | stmstation1@avive.life | W1146 | Agree | n/a | 2310190675 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-17 14:46 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190598 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-17 14:55 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-17 15:03 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.365 | Pass |  | n/a
2023-11-17 15:08 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-17 15:45 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190835 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-17 15:53 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190544 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.365 | Pass |  | n/a
2023-11-17 16:19 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190711 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.365 | Pass |  | n/a
2023-11-17 16:29 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190639 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.365 | Pass |  | n/a
2023-11-20 11:40 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-20 11:41 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-20 11:54 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-20 11:59 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.344 | Pass |  | n/a
2023-11-20 12:08 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190603 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-20 12:09 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190738 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-20 12:16 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190602 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-20 12:20 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190482 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-20 12:33 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190732 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-20 12:34 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.378 | Pass |  | n/a
2023-11-20 12:39 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190704 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-20 12:40 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190931 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-20 13:07 | stmstation3@avive.life | W1146 | Agree | Ireri Moreno | 2310190380 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-20 13:07 | stmstation4@avive.life | W1146 | Agree | Ireri Moreno | 2310190713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.344 | Pass |  | n/a
2023-11-20 13:20 | stmstation4@avive.life | W1146 | Agree | Ireri Moreno | 2310190605 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-20 13:21 | stmstation3@avive.life | W1146 | Agree | Ireri Moreno | 2310190635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-20 13:27 | stmstation4@avive.life | W1146 | Agree | Ireri Moreno | 2310190484 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.341 | Pass |  | n/a
2023-11-20 13:32 | stmstation3@avive.life | W1146 | Agree | Ireri Moreno | 2310190828 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-20 14:43 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190551 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-20 14:43 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-20 14:49 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190590 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.359 | Pass |  | n/a
2023-11-20 14:53 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190638 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-20 15:09 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-20 15:16 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190601 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-20 15:22 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190600 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.361 | Pass |  | n/a
2023-11-20 15:23 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190735 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-20 15:26 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190552 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-20 15:26 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190632 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-20 15:28 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-20 15:33 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190963 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-20 15:43 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190780 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-20 15:47 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190834 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-20 16:07 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-20 16:11 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-20 16:43 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190979 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.344 | Pass |  | n/a
2023-11-20 16:47 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190822 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-20 16:50 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190569 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-20 16:54 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190798 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-20 17:06 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.343 | Pass |  | n/a
2023-11-20 17:11 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190823 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-20 17:19 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190592 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-20 17:32 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190866 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-20 17:43 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190570 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | No cellular , has sim
2023-11-20 17:49 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190826 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-20 17:54 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190847 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-20 17:58 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190885 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-20 18:01 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190617 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-20 18:29 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-20 18:30 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190877 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-20 18:53 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190953 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-20 18:57 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190841 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-20 19:00 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190869 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-20 19:03 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-20 19:08 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190814 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-20 19:29 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190843 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-20 19:29 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190957 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-20 19:34 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190886 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-20 19:51 | stmstation3@avive.life | W1146 | Agree | Christina Cruz | 2310190777 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-20 19:52 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190515 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-20 19:57 | stmstation4@avive.life | W1146 | Agree | Christina Cruz | 2310190522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-21 07:25 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190958 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.361 | Pass |  | n/ a
2023-11-21 07:39 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190526 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-21 07:52 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190862 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-21 07:59 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190734 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-21 08:00 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190864 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-21 08:15 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 08:18 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190848 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-21 08:42 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190879 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-21 08:43 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190587 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.344 | Pass |  | n/a
2023-11-21 09:18 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190783 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-21 09:25 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190881 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-21 09:37 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190571 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-21 09:42 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190812 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-21 09:49 | stmstation2@avive.life | W1146 | Agree | Christina Cruz | 2310190583 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.499 | Pass |  | n/a
2023-11-21 09:54 | stmstation1@avive.life | W1146 | Agree | Christina Cruz | 2310190584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-21 10:10 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 10:23 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190589 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 431 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-21 10:23 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-21 10:23 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190686 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-21 10:32 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190684 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.343 | Pass |  | n/a
2023-11-21 10:33 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.344 | Pass |  | n/a
2023-11-21 10:35 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190923 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-21 10:40 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190530 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-21 10:40 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190730 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-21 10:44 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190572 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.356 | Pass |  | n/a
2023-11-21 10:47 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190736 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-21 10:48 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190314 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-21 10:54 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.378 | Pass |  | n/a
2023-11-21 10:57 | stmstation1@avive.life | W1146 | Agree | Precious Ungco | 2310190803 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-21 10:58 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190870 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-21 10:59 | stmstation2@avive.life | W1146 | Agree | Precious Ungco | 2310190688 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-21 12:47 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190810 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-21 12:48 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190891 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-21 12:48 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190904 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-21 12:55 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190511 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-21 13:02 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190512 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-21 13:02 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190762 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-21 13:10 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190884 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.356 | Pass |  | n/a
2023-11-21 13:13 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190930 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-21 13:23 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190507 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-21 13:30 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190728 | Pass | Fail | [10 empty] | No display visual.
2023-11-21 13:41 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190817 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-21 13:42 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190820 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-21 13:44 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 13:47 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190500 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-21 13:48 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190837 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-21 13:51 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190659 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.335 | Pass |  | n/a
2023-11-21 13:54 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-21 13:54 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190845 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.361 | Pass |  | n/a
2023-11-21 13:57 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190733 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-21 14:00 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190938 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-21 14:06 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190905 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.348 | Pass |  | N/A
2023-11-21 14:08 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190842 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-21 14:14 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190509 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-21 14:16 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190932 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-21 14:18 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190888 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.361 | Pass |  | n/a
2023-11-21 14:23 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190916 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-21 14:27 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190908 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-21 14:28 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190882 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-21 14:31 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190440 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-21 14:34 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190901 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-21 14:37 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190813 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-21 14:41 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190954 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.359 | Pass |  | n/a
2023-11-21 14:43 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190868 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-21 14:46 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190525 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.366 | Pass |  | n/a
2023-11-21 14:49 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190871 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 14:50 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.366 | Pass |  | n/a
2023-11-21 14:51 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190915 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2023-11-21 14:54 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190903 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.364 | Pass |  | n/a
2023-11-21 14:56 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190917 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-21 15:00 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190906 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.378 | Pass |  | n/a
2023-11-21 15:00 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190865 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 15:02 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190919 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.356 | Pass |  | n/a
2023-11-21 15:06 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190952 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-21 15:08 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190956 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-21 15:34 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190959 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.363 | Pass |  | n/a
2023-11-21 15:35 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190921 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 15:41 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 15:43 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190890 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-21 15:59 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190902 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-21 16:05 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190960 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.356 | Pass |  | n/a
2023-11-21 16:08 | stmstation3@avive.life | W1146 | Agree | Precious Ungco | 2310190824 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-21 16:10 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190816 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-21 16:12 | stmstation4@avive.life | W1146 | Agree | Precious Ungco | 2310190907 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.358 | Pass |  | n/a
2023-11-21 16:15 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190873 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-21 16:17 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-21 16:19 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190778 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 16:21 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190874 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.37 | Pass |  | n/a
2023-11-21 16:24 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190566 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-21 16:27 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190568 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.343 | Pass |  | n/a
2023-11-21 16:27 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.344 | Pass |  | n/a
2023-11-21 16:32 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190630 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-21 16:34 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190806 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-21 16:34 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 16:40 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-21 16:45 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190547 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.356 | Pass |  | n/a
2023-11-21 16:49 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.366 | Pass |  | n/a
2023-11-21 16:55 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-21 17:01 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190556 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-21 17:06 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190531 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.359 | Pass |  | n/a
2023-11-21 17:11 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 17:16 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190811 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-21 17:23 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190549 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-21 17:30 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190804 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-21 17:40 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190685 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-21 17:44 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190683 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-21 17:50 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190889 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-21 17:53 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190844 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-22 07:18 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190846 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.348 | Pass |  | N/A
2023-11-22 07:20 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190863 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 07:32 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190542 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-22 07:36 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-22 07:50 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190829 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 07:59 | stmstation1@avive.life | W1146 | Agree | Leonora Cruz | 2310190880 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-22 08:00 | stmstation2@avive.life | W1146 | Agree | Leonora Cruz | 2310190867 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.342 | Pass |  | n/a
2023-11-22 08:16 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190827 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-22 08:18 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190527 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-22 08:27 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190819 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-22 08:29 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190781 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.34 | Pass |  | n/a
2023-11-22 08:38 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190883 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-22 08:44 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190872 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.343 | Pass |  | n/a
2023-11-22 08:49 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190861 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2023-11-22 09:10 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190912 | Pass | Fail | [10 empty] | LCD Display not responding.
2023-11-22 09:15 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190727 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 09:23 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190928 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 09:29 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190909 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-22 09:47 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 09:57 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190518 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-22 10:01 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190559 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-22 10:11 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2023-11-22 10:18 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 10:22 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190913 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.353 | Pass |  | n/a
2023-11-22 10:22 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190710 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-22 10:28 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-22 10:31 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190739 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.364 | Pass |  | n/a
2023-11-22 10:34 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190494 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-22 10:35 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-22 10:39 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190609 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-22 10:44 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.344 | Pass |  | n/a
2023-11-22 10:46 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190640 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-22 11:02 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190485 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-22 11:03 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190935 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.356 | Pass |  | n/a
2023-11-22 11:03 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190758 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-22 11:04 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-22 11:10 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190934 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-22 11:11 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-22 11:16 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190925 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.356 | Pass |  | n/a
2023-11-22 11:19 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 11:20 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190926 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-22 11:21 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190936 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.36 | Pass |  | n/a
2023-11-22 11:52 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 11:53 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.362 | Pass |  | n/a
2023-11-22 12:15 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190831 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-22 12:16 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190581 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-22 12:26 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190784 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-22 12:29 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190619 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-22 12:34 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-22 12:40 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190585 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-22 12:50 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-22 12:56 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-22 13:01 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-22 13:02 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190763 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-22 13:04 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190620 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-22 13:10 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190776 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-22 13:10 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190833 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-22 13:15 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190955 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.356 | Pass |  | n/a
2023-11-22 13:18 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.415 | Pass |  | n/a
2023-11-22 13:19 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190779 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.344 | Pass |  | n/a
2023-11-22 13:22 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190815 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-22 13:27 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-22 13:28 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190821 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-22 13:28 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190838 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 13:34 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190514 | Pass | Fail | [10 empty] | Self-test at risk
2023-11-22 13:36 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190642 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.343 | Pass |  | n/a
2023-11-22 13:39 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190621 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.343 | Pass |  | n/a
2023-11-22 13:43 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190840 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-22 13:47 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-22 13:54 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-22 14:03 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-22 14:09 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 14:10 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-22 14:11 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.418 | Pass |  | n/a
2023-11-22 14:13 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.345 | Pass |  | n/a
2023-11-22 14:21 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190498 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-22 14:27 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190914 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-22 14:34 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190521 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.357 | Pass |  | n/a
2023-11-22 14:35 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.351 | Pass |  | n/a
2023-11-22 14:57 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190910 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.355 | Pass |  | n/a
2023-11-22 15:04 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190631 | Pass | Pass | Pass | Fail | [8 empty] | GPS Satellite does not appear after several attempts.
2023-11-22 15:12 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190486 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.349 | Pass |  | n/a
2023-11-22 15:16 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190818 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-22 15:20 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190660 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.347 | Pass |  | n/a
2023-11-22 15:28 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190519 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.346 | Pass |  | n/a
2023-11-22 15:28 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190634 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.361 | Pass |  | n/a
2023-11-22 15:31 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-22 15:32 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190933 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.37 | Pass |  | n/a
2023-11-22 15:36 | stmstation3@avive.life | W1146 | Agree | Leonora Cruz | 2310190924 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.356 | Pass |  | n/a
2023-11-22 15:38 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Fail |  |  | Failed reading SIM card
2023-11-22 15:39 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190495 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.348 | Pass |  | n/a
2023-11-22 15:47 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190792 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.35 | Pass |  | n/a
2023-11-22 15:51 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190737 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.354 | Pass |  | n/a
2023-11-22 15:54 | stmstation2@avive.life | W1146 | Agree | Violeta Ayala | 2310190878 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.352 | Pass |  | n/a
2023-11-22 16:10 | stmstation3@avive.life | W1146 | Agree | Ananya Jain | 2310190631 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.354 | Pass |  | Tested on different fixture.
2023-11-22 16:12 | stmstation1@avive.life | W1146 | Agree | Violeta Ayala | 2310190567 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | No cellular, No sim
2023-11-22 16:15 | stmstation4@avive.life | W1146 | Agree | Leonora Cruz | 2310190516 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | CELLULAR NOT RESPONDING/SIM
2023-11-27 12:22 | stmstation2@avive.life | W1137 | Agree | Precious Ungco | 2305300461 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.514 | Pass |  | n/a
2023-11-27 12:28 | stmstation1@avive.life | W1137 | Agree | Precious Ungco | 2305300417 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.513 | Pass |  | n/a
2023-11-27 12:28 | stmstation3@avive.life | W1137 | Agree | Leonora Cruz | 2308110549 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.515 | Pass |  | n/a
2023-11-27 12:33 | stmstation2@avive.life | W1137 | Agree | Precious Ungco | 2307110579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.519 | Pass |  | n/a
2023-11-27 12:53 | stmstation2@avive.life | W1137 | Agree | Precious Ungco | 2305300831 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.512 | Pass |  | n/a
2023-11-27 12:56 | stmstation1@avive.life | W1137 | Agree | Precious Ungco | 2308110222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.535 | Pass |  | n/a
2023-11-27 13:24 | stmstation4@avive.life | W1137 | Agree | Leonora Cruz | 2308110049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.535 | Pass |  | n/a
2023-11-27 13:25 | stmstation3@avive.life | W1137 | Agree | Leonora Cruz | 2305300728 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.515 | Pass |  | n/a
2023-11-27 13:29 | stmstation1@avive.life | W1137 | Agree | Precious Ungco | 2307110112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.51 | Pass |  | n/a
2023-11-27 13:30 | stmstation2@avive.life | W1137 | Agree | Precious Ungco | 2308110555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.531 | Pass |  | n/a
2023-11-27 13:34 | stmstation3@avive.life | W1137 | Agree | precious ungco | 2305300816 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.514 | Pass |  | n/a
2023-11-27 13:35 | stmstation4@avive.life | W1137 | Agree | Precious UNGCO | 2308110258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.515 | Pass |  | N/A
2023-11-27 13:38 | stmstation1@avive.life | W1137 | Agree | Precious Ungco | 2308110300 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.518 | Pass |  | n/a
2023-11-27 13:41 | stmstation2@avive.life | W1137 | Agree | Precious Ungco | 2308110835 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 428 | Pass | Pass | 0.513 | Pass |  | n/a
2023-11-27 13:46 | stmstation4@avive.life | W1137 | Agree | Leonora Cruz | 2305300460 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.507 | Pass |  | n/a
2023-11-27 14:10 | stmstation1@avive.life | W1137 | Agree | Precious Ungco | 2308110718 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.523 | Pass |  | N/A
2023-11-27 14:25 | stmstation4@avive.life | W1137 | Agree | Leonora Cruz | 2305300481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.521 | Pass |  | n/a
2023-11-27 14:37 | stmstation4@avive.life | W1137 | Agree | Leonora Cruz | 2307110150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.507 | Pass |  | n/a
2023-11-27 14:43 | stmstation4@avive.life | W1137 | Agree | Leonora Cruz | 2308110093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.53 | Pass |  | n/a
2023-11-27 15:20 | stmstation1@avive.life | W1137 | Agree | Leonora Cruz | 2307110539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.508 | Pass |  | n/a
2023-11-27 15:21 | stmstation4@avive.life | W1137 | Agree | Leonora Cruz | 2308110079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.574 | Pass |  | n/a
2023-11-27 15:22 | stmstation3@avive.life | W1137 | Agree | Leonora Cruz | 2305300744 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.515 | Pass |  | n/a
2023-11-27 15:58 | stmstation1@avive.life | W1137 | Agree | Violeta Ayala | 2307110284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.518 | Pass |  | n/a
2023-11-27 16:02 | stmstation4@avive.life | W1137 | Agree | Violeta Ayala | 2305300697 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 525 | Pass |  | n/a
2023-11-27 16:06 | stmstation3@avive.life | W1137 | Agree | Violeta Ayala | 2307110273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.518 | Pass |  | n/a
2023-11-27 16:09 | stmstation2@avive.life | W1137 | Agree | Violeta Ayala | 2308110348 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.52 | Pass |  | n/a
2023-11-27 16:14 | stmstation1@avive.life | W1137 | Agree | Violeta Ayala | 2305300734 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.51 | Pass |  | n/a
2023-11-27 16:17 | stmstation4@avive.life | W1137 | Agree | Violeta Ayala | 2305300636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.511 | Pass |  | n/a
2023-11-27 16:18 | stmstation3@avive.life | W1137 | Agree | Violeta Ayala | 2305300836 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.511 | Pass |  | n/a
2023-11-27 16:24 | stmstation2@avive.life | W1137 | Agree | Violeta Ayala | 2308110854 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.533 | Pass |  | n/a
2023-11-28 10:13 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2308110471 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.546 | Pass |  | n/a
2023-11-28 10:24 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2308110114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.53 | Pass |  | n/a
2023-11-28 10:36 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.513 | Pass |  | n/a
2023-11-28 10:38 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2308110264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-28 10:53 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2307110528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.52 | Pass |  | n/a
2023-11-28 11:17 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2307110691 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.51 | Pass |  | n/a
2023-11-28 12:40 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110711 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.512 | Pass |  | n/a
2023-11-28 12:49 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2308110137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.525 | Pass |  | n/a
2023-11-28 13:06 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2308110134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.532 | Pass |  | n/a
2023-11-28 13:17 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2307110715 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.517 | Pass |  | n/a
2023-11-28 13:28 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2308110254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.521 | Pass |  | n/a
2023-11-28 13:33 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2305300859 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.512 | Pass |  | n/a
2023-11-28 13:42 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2307110611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.524 | Pass |  | n/a
2023-11-28 13:43 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110604 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.508 | Pass |  | n/a
2023-11-28 13:54 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110606 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.512 | Pass |  | n/a
2023-11-28 13:55 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2305300198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-28 14:14 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2308110011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.533 | Pass |  | n/a
2023-11-28 14:16 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.518 | Pass |  | n/a
2023-11-28 14:23 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2307110099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.509 | Pass |  | n/a
2023-11-28 14:24 | stmstation3@avive.life | W1138 | Agree | Leonora Cruz | 2308110282 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.514 | Pass |  | n/a
2023-11-28 14:32 | stmstation3@avive.life | W1138 | Agree | Leonora Cruz | 2308110564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.515 | Pass |  | n/a
2023-11-28 14:35 | stmstation4@avive.life | W1138 | Agree | Leonora Cruz | 2307110333 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-28 14:52 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.506 | Pass |  | N/A
2023-11-28 14:55 | stmstation4@avive.life | W1138 | Agree | Leonora Cruz | 2305300807 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.513 | Pass |  | n/a
2023-11-28 14:56 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2307110226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.517 | Pass |  | n/a
2023-11-28 15:04 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.513 | Pass |  | n/a
2023-11-28 15:16 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2307110078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.51 | Pass |  | n/a
2023-11-28 15:24 | stmstation4@avive.life | W1138 | Agree | Violeta Ayala | 2307110021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.511 | Pass |  | n/a
2023-11-28 15:27 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2305300434 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.514 | Pass |  | n/a
2023-11-28 15:54 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2308110487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.514 | Pass |  | N/A
2023-11-28 17:01 | stmstation2@avive.life | W1138 | Agree | Ananya Jain | 2307110336 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.517 | Pass |  | N/A
2023-11-28 17:24 | stmstation3@avive.life | W1138 | Agree | Leonora Cruz | 2308110379 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.521 | Pass |  | n/a
2023-11-28 17:31 | stmstation4@avive.life | W1138 | Agree | Leonora Cruz | 2307110878 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.509 | Pass |  | n/a
2023-11-28 17:42 | stmstation4@avive.life | W1138 | Agree | Leonora Cruz | 2305300693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.509 | Pass |  | n/a
2023-11-28 17:47 | stmstation4@avive.life | W1138 | Agree | Leonora Cruz | 2307110565 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-28 17:48 | stmstation3@avive.life | W1138 | Agree | Leonora Cruz | 2307110990 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.523 | Pass |  | n/a
2023-11-29 07:29 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2308110557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.514 | Pass |  | n/a
2023-11-29 07:30 | stmstation4@avive.life | W1138 | Agree | Leonora Cruz | 2308110443 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.511 | Pass |  | n/a
2023-11-29 07:31 | stmstation3@avive.life | W1138 | Agree | Leonora Cruz | 2308110345 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.525 | Pass |  | n/a
2023-11-29 07:36 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2308110453 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.503 | Pass |  | n/a
2023-11-29 07:42 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2307110658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.513 | Pass |  | n/a
2023-11-29 08:13 | stmstation4@avive.life | W1138 | Agree | Violeta Ayala | 2305300759 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.503 | Pass |  | n/a
2023-11-29 08:24 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2308110543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.51 | Pass |  | n/a
2023-11-29 08:32 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2307110591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.507 | Pass |  | n/a
2023-11-29 08:33 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110914 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.511 | Pass |  | n/a
2023-11-29 09:20 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2307110654 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.503 | Pass |  | n/a
2023-11-29 09:42 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.51 | Pass |  | n/a
2023-11-29 09:47 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2307110223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.512 | Pass |  | n/a
2023-11-29 10:02 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2308110399 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.529 | Pass |  | n/a
2023-11-29 10:09 | stmstation4@avive.life | W1138 | Agree | Violeta Ayala | 2308110084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.527 | Pass |  | n/a
2023-11-29 10:13 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110270 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-29 10:21 | stmstation2@avive.life | W1138 | Agree | Precious Ungco | 2308110381 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.524 | Pass |  | n/a
2023-11-29 11:03 | stmstation2@avive.life | W1138 | Agree | Precious Ungco | 2308110280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-29 11:07 | stmstation3@avive.life | W1138 | Agree | Precious Ungco | 2308110321 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.524 | Pass |  | n/a
2023-11-29 11:10 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2307110612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.52 | Pass |  | n/a
2023-11-29 11:14 | stmstation4@avive.life | W1138 | Agree | Precious Ungco | 2305300188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.271 | Pass |  | n/a
2023-11-29 11:19 | stmstation2@avive.life | W1138 | Agree | Precious Ungco | 2307110747 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.514 | Pass |  | n/a
2023-11-29 11:22 | stmstation1@avive.life | W1138 | Agree | Precious Ungco | 2308110062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.533 | Pass |  | n/a
2023-11-29 11:29 | stmstation2@avive.life | W1138 | Agree | Precious Ungco | 2308110637 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.513 | Pass |  | n/a
2023-11-29 11:37 | stmstation3@avive.life | W1138 | Agree | Precious Ungco | 2307110673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.518 | Pass |  | n/a
2023-11-29 11:44 | stmstation4@avive.life | W1138 | Agree | Precious Ungco | 2308110740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.521 | Pass |  | n/a
2023-11-29 11:52 | stmstation2@avive.life | W1138 | Agree | Precious Ungco | 2307110827 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.523 | Pass |  | n/a
2023-11-29 12:15 | stmstation1@avive.life | W1138 | Agree | Precious Ungco | 2308110135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.528 | Pass |  | n/a
2023-11-29 12:29 | stmstation4@avive.life | W1138 | Agree | Liza Ordillo | 2308110485 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.518 | Pass |  | n/a
2023-11-29 13:02 | stmstation2@avive.life | W1138 | Agree | Precious Ungco | 2308110153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.534 | Pass |  | n/a
2023-11-29 14:04 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2307110753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.544 | Pass |  | n/a
2023-11-29 14:08 | stmstation4@avive.life | W1138 | Agree | Liza Ordillo | 2308110770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.58 | Pass |  | n/a
2023-11-29 14:25 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2308110251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.582 | Pass |  | n/a
2023-11-29 14:40 | stmstation4@avive.life | W1138 | Agree | Liza Ordillo | 2307110744 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.535 | Pass |  | n/a
2023-11-29 14:44 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2308110007 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.525 | Pass |  | n/a
2023-11-29 14:52 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2308110183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.533 | Pass |  | n/a
2023-11-29 15:07 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2307110833 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.535 | Pass |  | n/a
2023-11-29 15:52 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2307110264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.556 | Pass |  | n/a
2023-11-29 15:52 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2307110179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.517 | Pass |  | n/a
2023-11-29 16:02 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2307110742 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.519 | Pass |  | n/a
2023-11-29 16:09 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2307110443 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-29 16:16 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2307110414 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.543 | Pass |  | n/a
2023-11-29 16:16 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2305300131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.572 | Pass |  | n/a
2023-11-29 16:28 | stmstation4@avive.life | W1138 | Agree | Liza Ordillo | 2307110297 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.526 | Pass |  | n/a
2023-11-29 16:39 | stmstation1@avive.life | W1138 | Agree | Christina Cruz | 2308110410 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.521 | Pass |  | n/a
2023-11-29 17:13 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2307110501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-29 17:15 | stmstation2@avive.life | W1138 | Agree | Ananya Jain | 2307110029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-29 17:17 | stmstation4@avive.life | W1138 | Agree | Violeta Ayala | 2307110305 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.653 | Pass |  | n/a
2023-11-29 17:24 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2305300750 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-29 17:25 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2307110458 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.512 | Pass |  | n/a
2023-11-29 17:31 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2307110597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.513 | Pass |  | n/a
2023-11-29 17:39 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2305300588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.512 | Pass |  | n/a
2023-11-30 06:34 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2307110364 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-30 07:00 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2307110767 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.581 | Pass |  | n/a
2023-11-30 07:18 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2307110722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.564 | Pass |  | n/a
2023-11-30 07:22 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2307110870 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.513 | Pass |  | n/a
2023-11-30 07:30 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2307110168 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.577 | Pass |  | n/a
2023-11-30 07:35 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2308110405 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.514 | Pass |  | n/a
2023-11-30 07:41 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2307110374 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.524 | Pass |  | n/a
2023-11-30 08:02 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2307110352 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.516 | Pass |  | n/a
2023-11-30 08:10 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2308110444 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.518 | Pass |  | n/a
2023-11-30 08:14 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2307110517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.52 | Pass |  | n/a
2023-11-30 08:27 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2307110483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.535 | Pass |  | n/a
2023-11-30 08:30 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2307110079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.508 | Pass |  | n/a
2023-11-30 08:40 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2307110290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.515 | Pass |  | n/a
2023-11-30 08:41 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2308110430 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.535 | Pass |  | n/a
2023-11-30 08:56 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2307110540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.525 | Pass |  | n/a
2023-11-30 09:18 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2307110119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.578 | Pass |  | n/a
2023-11-30 09:21 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2307110064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.535 | Pass |  | n/a
2023-11-30 09:22 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2305300668 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.512 | Pass |  | N/A
2023-11-30 09:36 | stmstation4@avive.life | W1138 | Agree | Liza Ordillo | 2307110237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.514 | Pass |  | n/a
2023-11-30 09:41 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2307110039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.534 | Pass |  | n/a
2023-11-30 09:48 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2308110338 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.524 | Pass |  | n/a
2023-11-30 09:54 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2308110426 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.52 | Pass |  | N/A
2023-11-30 09:57 | stmstation4@avive.life | W1138 | Agree | Liza Ordillo | 2307110628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.513 | Pass |  | n/a
2023-11-30 10:00 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2308110729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.51 | Pass |  | N/A
2023-11-30 10:07 | stmstation3@avive.life | W1138 | Agree | Liza Ordillo | 2307110257 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.517 | Pass |  | n/a
2023-11-30 10:07 | stmstation4@avive.life | W1138 | Agree | Liza Ordillo | 2308110542 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.511 | Pass |  | n/a
2023-11-30 10:18 | stmstation1@avive.life | W1138 | Agree | Liza Ordillo | 2307110795 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.516 | Pass |  | N/A
2023-11-30 10:23 | stmstation2@avive.life | W1138 | Agree | Liza Ordillo | 2308110629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.515 | Pass |  | n/a
2023-11-30 10:24 | stmstation4@avive.life | W1138 | Agree | Liza Ordillo | 2307110671 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.501 | Pass |  | N/A
2023-11-30 10:48 | stmstation2@avive.life | W1138 | Agree | Christina Cruz | 2307110843 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.518 | Pass |  | N/A
2023-11-30 10:48 | stmstation1@avive.life | W1138 | Agree | Christina Cruz | 2308110271 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.521 | Pass |  | N/A
2023-11-30 11:12 | stmstation1@avive.life | W1138 | Agree | Leonora Cruz | 2307110361 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.574 | Pass |  | n/a
2023-11-30 11:19 | stmstation2@avive.life | W1138 | Agree | Leonora Cruz | 2307110891 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.519 | Pass |  | n/a
2023-11-30 11:28 | stmstation3@avive.life | W1138 | Agree | Christina Cruz | 2308110036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 4.2 | Fail | current too high during sleep mode
2023-11-30 11:35 | stmstation2@avive.life | W1138 | Agree | Leonora Cruz | 2307110961 | Pass | Fail | [10 empty] | failed won't turn on.
2023-11-30 11:46 | stmstation2@avive.life | W1138 | Agree | Leonora Cruz | 2307110550 | Pass | Fail | [10 empty] | Screen shuts down
2023-11-30 12:06 | stmstation1@avive.life | W1138 | Agree | Leonora Cruz | 2308110290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.518 | Pass |  | N/A
2023-11-30 12:10 | stmstation3@avive.life | W1138 | Agree | Leonora Cruz | 2308110071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.54 | Pass |  | n/a
2023-11-30 12:22 | stmstation3@avive.life | W1138 | Agree | Leonora Cruz | 2308110216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.54 | Pass |  | n/a
2023-11-30 12:57 | stmstation2@avive.life | W1138 | Agree | Leonora Cruz | 2307110238 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 3.88 | Fail | out of range
2023-11-30 12:59 | stmstation3@avive.life | W1138 | Agree | Leonora Cruz | 2307110731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.512 | Pass |  | n/a
2023-11-30 13:08 | stmstation4@avive.life | W1138 | Agree | Leonora Cruz | 2308110089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.534 | Pass |  | n/a
2023-11-30 13:10 | stmstation3@avive.life | W1138 | Agree | Leonora Cruz | 2307110492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.253 | Pass |  | n/a
2023-11-30 13:11 | stmstation1@avive.life | W1138 | Agree | Leonora Cruz | 2308110272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.522 | Pass |  | N/A
2023-11-30 13:32 | stmstation2@avive.life | W1138 | Agree | Leonora Cruz | 2307110470 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.519 | Pass |  | n/a
2023-11-30 13:55 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2307110410 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.514 | Pass |  | n/a
2023-11-30 13:57 | stmstation4@avive.life | W1138 | Agree | Leonora Cruz | 2308110584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.564 | Pass |  | N/A
2023-11-30 13:58 | stmstation3@avive.life | W1138 | Agree | Leonora Cruz | 2308110227 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.764 | Pass |  | N/A
2023-11-30 14:19 | stmstation2@avive.life | W1138 | Agree | Violeta Ayala | 2307110700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.51 | Pass |  | n/a
2023-11-30 14:37 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2307110661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.529 | Pass |  | n/a
2023-11-30 15:29 | stmstation3@avive.life | W1138 | Agree | Violeta Ayala | 2308110281 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.513 | Pass |  | N/A
2023-11-30 16:38 | stmstation4@avive.life | W1138 | Agree | Violeta Ayala | 2307110831 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Fail |  |  | Fail NO SIM
2023-11-30 17:19 | stmstation3@avive.life | W1138 | Agree | Christina Cruz | 2307110961 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.517 | Pass |  | n/a
2023-11-30 17:37 | stmstation1@avive.life | W1138 | Agree | Violeta Ayala | 2307110587 | Pass | Pass | Pass | Fail | [8 empty] | No GPS
2023-12-01 06:52 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191197 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.391 | Pass |  | n/a
2023-12-01 06:56 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-01 07:05 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.415 | Pass |  | n/a
2023-12-01 07:09 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191113 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.401 | Pass |  | n/a
2023-12-01 07:15 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.42 | Pass |  | n/a
2023-12-01 07:19 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.42 | Pass |  | n/a
2023-12-01 07:33 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.434 | Pass |  | n/a
2023-12-01 07:36 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.413 | Pass |  | n/a
2023-12-01 07:42 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191103 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.434 | Pass |  | n/a
2023-12-01 07:45 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.43 | Pass |  | n/a
2023-12-01 08:02 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.43 | Pass |  | n/a
2023-12-01 08:02 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.428 | Pass |  | n/a
2023-12-01 08:11 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190997 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.417 | Pass |  | n/a
2023-12-01 08:13 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191172 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.423 | Pass |  | n/a
2023-12-01 08:19 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190893 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.433 | Pass |  | n/a
2023-12-01 08:24 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.405 | Pass |  | n/a
2023-12-01 08:28 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191063 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.424 | Pass |  | n/a
2023-12-01 08:34 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191061 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.426 | Pass |  | n/a
2023-12-01 08:38 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.43 | Pass |  | n/a
2023-12-01 08:43 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191018 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.424 | Pass |  | n/a
2023-12-01 08:49 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.433 | Pass |  | n/a
2023-12-01 08:51 | stmstation3@avive.life | W1153 | Agree | Valeria Manzo | 2310191055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-01 08:54 | stmstation4@avive.life | W1153 | Agree | Valeria Manzo | 2310191059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-01 09:00 | stmstation3@avive.life | W1153 | Agree | Valeria Manzo | 2310190693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-01 09:01 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.421 | Pass |  | n/a
2023-12-01 09:05 | stmstation4@avive.life | W1153 | Agree | Valeria Manzo | 2310191187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-01 09:06 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191200 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.45 | Pass |  | n/a
2023-12-01 09:17 | stmstation3@avive.life | W1153 | Agree | Valeria Manzo | 2310191088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-01 09:17 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.42 | Pass |  | n/a
2023-12-01 09:22 | stmstation4@avive.life | W1153 | Agree | Valeria Manzo | 2310190691 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-01 09:23 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.371 | Pass |  | n/a
2023-12-01 09:27 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190699 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.366 | Pass |  | n/a
2023-12-01 09:29 | stmstation3@avive.life | W1153 | Agree | Valeria Manzo | 2310190692 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.366 | Pass |  | n/a
2023-12-01 09:40 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-01 09:40 | stmstation4@avive.life | W1153 | Agree | Valeria Manzo | 2310191160 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-01 09:40 | stmstation3@avive.life | W1153 | Agree | Valeria Manzo | 2310191072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-01 09:48 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191166 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.418 | Pass |  | n/a
2023-12-01 09:49 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-01 09:51 | stmstation4@avive.life | W1153 | Agree | Valeria Manzo | 2310191170 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-01 09:51 | stmstation3@avive.life | W1153 | Agree | Valeria Manzo | 2310191192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-01 09:56 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-01 10:00 | stmstation3@avive.life | W1153 | Agree | Valeria Manzo | 2310191101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-01 10:01 | stmstation4@avive.life | W1153 | Agree | Valeria Manzo | 2310191044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-01 10:03 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191066 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.408 | Pass |  | n/a
2023-12-01 10:07 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-01 10:11 | stmstation4@avive.life | W1153 | Agree | Valeria Manzo | 2310191148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-01 10:12 | stmstation3@avive.life | W1153 | Agree | Valeria Manzo | 2310191169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-01 10:21 | stmstation1@avive.life | W1153 | Agree | Valeria Manzo | 2310191171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-01 10:23 | stmstation3@avive.life | W1153 | Agree | Valeria Manzo | 2310191185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-01 10:24 | stmstation4@avive.life | W1153 | Agree | Valeria Manzo | 2310191189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-01 10:37 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310191058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-01 10:40 | stmstation2@avive.life | W1153 | Agree | Christina Cruz | 2310191025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-01 10:46 | stmstation1@avive.life | W1153 | Agree | Christina Cruz | 2310190698 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-01 10:46 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.361 | Pass |  | n/a
2023-12-01 10:55 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310191181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.375 | Pass |  | n/a
2023-12-01 11:04 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310191024 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.362 | Pass |  | n/a
2023-12-01 11:04 | stmstation2@avive.life | W1153 | Agree | Valeria Manzo | 2310191168 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-01 11:06 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.384 | Pass |  | n/a
2023-12-01 11:13 | stmstation2@avive.life | W1153 | Agree | Valeria Manzo | 2310191078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-01 11:15 | stmstation1@avive.life | W1153 | Agree | Valeria Manzo | 2310191021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.369 | Pass |  | n/a
2023-12-01 11:24 | stmstation2@avive.life | W1153 | Agree | Valeria Manzo | 2310190664 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-01 11:25 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310190918 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-01 11:27 | stmstation1@avive.life | W1153 | Agree | Valeria Manzo | 2310191106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-01 11:36 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310190652 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.405 | Pass |  | n/a
2023-12-01 11:38 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310191195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.358 | Pass |  | n/a
2023-12-01 11:38 | stmstation2@avive.life | W1153 | Agree | Valeria Manzo | 2310191129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-01 11:47 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.356 | Pass |  | n/a
2023-12-01 11:48 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310191156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-01 11:49 | stmstation2@avive.life | W1153 | Agree | Valeria Manzo | 2310191183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-01 11:50 | stmstation1@avive.life | W1153 | Agree | Valeria Manzo | 2310190669 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-01 12:02 | stmstation1@avive.life | W1153 | Agree | Precious Ungco | 2310190672 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-01 12:04 | stmstation3@avive.life | W1153 | Agree | Precious Ungco | 2310191131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-01 12:08 | stmstation4@avive.life | W1153 | Agree | Precious Ungco | 2310190668 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-01 12:11 | stmstation3@avive.life | W1153 | Agree | Precious Ungco | 2310191132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.339 | Pass |  | n/a
2023-12-01 12:12 | stmstation1@avive.life | W1153 | Agree | Christina Cruz | 2310190996 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-01 12:15 | stmstation4@avive.life | W1153 | Agree | Precious Ungco | 2310191128 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-01 12:17 | stmstation2@avive.life | W1153 | Agree | Christina Cruz | 2310190920 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-01 12:18 | stmstation3@avive.life | W1153 | Agree | Precious Ungco | 2310191127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-01 12:23 | stmstation4@avive.life | W1153 | Agree | Precious Ungco | 2310191159 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-01 12:24 | stmstation3@avive.life | W1153 | Agree | Precious Ungco | 2310191158 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-01 12:32 | stmstation3@avive.life | W1153 | Agree | Precious Ungco | 2310191194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-01 12:33 | stmstation4@avive.life | W1153 | Agree | Precious Ungco | 2310191193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-01 12:34 | stmstation2@avive.life | W1153 | Agree | Christina Cruz | 2310191064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-01 12:57 | stmstation1@avive.life | W1153 | Agree | Christina Cruz | 2310190970 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-01 13:18 | stmstation4@avive.life | W1153 | Agree | Precious Ungco | 2310191151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-01 13:19 | stmstation3@avive.life | W1153 | Agree | Precious Ungco | 2310190671 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-01 13:29 | stmstation4@avive.life | W1153 | Agree | Precious Ungco | 2310190665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-01 16:05 | stmstation2@avive.life | W1153 | Agree | Christina Cruz | 2310191065 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-01 16:06 | stmstation1@avive.life | W1153 | Agree | Christina Cruz | 2310190695 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-01 16:17 | stmstation1@avive.life | W1153 | Agree | Christina Cruz | 2310191186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-04 07:00 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.413 | Pass |  | n/a
2023-12-04 07:05 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-04 07:13 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310190538 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.42 | Pass |  | n/a
2023-12-04 07:13 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.487 | Pass |  | n/a
2023-12-04 07:25 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191182 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-04 07:28 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190606 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-04 07:37 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190894 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.358 | Pass |  | n/a
2023-12-04 07:38 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191083 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.416 | Pass |  | n/a
2023-12-04 07:40 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310190988 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.427 | Pass |  | n/a
2023-12-04 07:45 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-04 07:53 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.433 | Pass |  | n/a
2023-12-04 08:04 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310190858 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-04 08:04 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-04 08:12 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.373 | Pass |  | n/a
2023-12-04 08:15 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.356 | Pass |  | n/a
2023-12-04 08:24 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310190857 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-04 08:32 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190951 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.362 | Pass |  | n/a
2023-12-04 08:33 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.386 | Pass |  | n/a
2023-12-04 08:42 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190653 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-04 08:44 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.437 | Pass |  | n/a
2023-12-04 09:01 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190696 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.36 | Pass |  | n/a
2023-12-04 09:04 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.36 | Pass |  | n/a
2023-12-04 09:14 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191162 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.426 | Pass |  | n/a
2023-12-04 09:37 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310190986 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-04 09:44 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310190700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-04 09:48 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191124 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.421 | Pass |  | n/a
2023-12-04 09:50 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310191014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-04 09:52 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191082 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.384 | Pass |  | n/a
2023-12-04 09:58 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.424 | Pass |  | n/a
2023-12-04 10:07 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310191073 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.36 | Pass |  | n/a
2023-12-04 10:09 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.42 | Pass |  | n/a
2023-12-04 10:15 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310190093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.356 | Pass |  | n/a
2023-12-04 10:28 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310190120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.36 | Pass |  | n/a
2023-12-04 10:28 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310191134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-04 10:35 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310190118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-04 10:39 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310190401 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.36 | Pass |  | n/a
2023-12-04 10:58 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.415 | Pass |  | n/a
2023-12-04 11:00 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.439 | Pass |  | n/a
2023-12-04 11:09 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.424 | Pass |  | n/a
2023-12-04 11:15 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-04 11:16 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191077 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.416 | Pass |  | n/a
2023-12-04 11:32 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190697 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-04 11:33 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310190839 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-04 11:35 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310190510 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.36 | Pass |  | n/a
2023-12-04 11:42 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191191 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.437 | Pass |  | n/a
2023-12-04 11:47 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.426 | Pass |  | n/a
2023-12-04 11:51 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310191070 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Fail |  |  |  | range to high
2023-12-04 12:02 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310190836 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-04 13:02 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.357 | Pass |  | n/a
2023-12-04 13:03 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310191090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.358 | Pass |  | n/a
2023-12-04 13:04 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.421 | Pass |  | n/a
2023-12-04 13:43 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-04 13:50 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190948 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.356 | Pass |  | N/A
2023-12-04 13:54 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.415 | Pass |  | n/a
2023-12-04 13:57 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.351 | Pass |  | N/A
2023-12-04 14:07 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.415 | Pass |  | n/a
2023-12-04 14:12 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-04 14:16 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190565 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.356 | Pass |  | N/A
2023-12-04 14:18 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.349 | Pass |  | N/A
2023-12-04 14:20 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.349 | Pass |  | N/A
2023-12-04 14:20 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310191002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.35 | Pass |  | N/A
2023-12-04 14:23 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190575 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.359 | Pass |  | N/A
2023-12-04 14:28 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190852 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.432 | Pass |  | n/a
2023-12-04 14:33 | stmstation4@avive.life | W1153 | Agree | Liza Ordillo | 2310190579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-04 14:33 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-04 14:34 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-04 14:37 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190851 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-04 14:52 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310191177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-04 14:53 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.427 | Pass |  | n/a
2023-12-04 14:53 | stmstation4@avive.life | W1153 | Agree | Leonora Cruz | 2310191009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-04 15:00 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310191095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.367 | Pass |  | n/a
2023-12-04 15:03 | stmstation4@avive.life | W1153 | Agree | Leonora Cruz | 2310191094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.362 | Pass |  | n/a
2023-12-04 15:06 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310190646 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-04 15:11 | stmstation4@avive.life | W1153 | Agree | Leonora Cruz | 2310191010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-04 15:15 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310191030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-04 15:16 | stmstation1@avive.life | W1153 | Agree | Leonora Cruz | 2310191144 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.421 | Pass |  | n/a
2023-12-04 15:19 | stmstation4@avive.life | W1153 | Agree | Leonora Cruz | 2310190651 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-04 15:20 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310190564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.42 | Pass |  | n/a
2023-12-04 15:22 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310191028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-04 15:25 | stmstation4@avive.life | W1153 | Agree | Leonora Cruz | 2310190578 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.337 | Pass |  | n/a
2023-12-04 15:28 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.359 | Pass |  | n/a
2023-12-04 15:32 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310190987 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-04 15:33 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190797 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.407 | Pass |  | n/a
2023-12-04 15:34 | stmstation4@avive.life | W1153 | Agree | Leonora Cruz | 2310190887 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-04 15:40 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310190508 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-04 15:43 | stmstation4@avive.life | W1153 | Agree | Leonora Cruz | 2310191179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-04 15:47 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310190670 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-04 15:51 | stmstation4@avive.life | W1153 | Agree | Leonora Cruz | 2310190580 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-04 15:54 | stmstation3@avive.life | W1153 | Agree | Leonora Cruz | 2310190995 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-04 15:56 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190457 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.426 | Pass |  | n/a
2023-12-05 06:42 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310190998 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.434 | Pass |  | n/a
2023-12-05 07:01 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310191000 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.446 | Pass |  | n/a
2023-12-05 07:03 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-05 07:08 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-05 07:16 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-05 07:23 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.363 | Pass |  | n/a
2023-12-05 07:29 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.421 | Pass |  | n/a
2023-12-05 07:29 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190875 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.347 | Pass |  | N/A
2023-12-05 07:36 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191041 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.389 | Pass |  | n/a
2023-12-05 07:36 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.343 | Pass |  | N/A
2023-12-05 07:38 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.434 | Pass |  | n/a
2023-12-05 07:43 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190626 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-05 07:50 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.391 | Pass |  | n/a
2023-12-05 07:54 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | N/A
2023-12-05 07:56 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190859 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.487 | Pass |  | N/A
2023-12-05 07:59 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-05 08:00 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.431 | Pass |  | n/a
2023-12-05 08:09 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190961 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.35 | Pass |  | N/A
2023-12-05 08:20 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190971 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.352 | Pass |  | N/A
2023-12-05 08:29 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190900 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-05 08:29 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190982 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.378 | Pass |  | n/a
2023-12-05 08:30 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190967 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.353 | Pass |  | N/A
2023-12-05 08:41 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190966 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.345 | Pass |  | N/A
2023-12-05 08:45 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190968 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.474 | Pass |  | n/a
2023-12-05 08:47 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190976 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-05 08:51 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190972 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.355 | Pass |  | N/A
2023-12-05 08:58 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310190975 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-05 08:59 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190949 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.421 | Pass |  | n/a
2023-12-05 08:59 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190849 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.354 | Pass |  | N/A
2023-12-05 09:11 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190969 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.356 | Pass |  | N/A
2023-12-05 09:15 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.348 | Pass |  | N/A
2023-12-05 09:22 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190974 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.349 | Pass |  | N/A
2023-12-05 09:29 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190941 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.35 | Pass |  | N/A
2023-12-05 09:40 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190947 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.434 | Pass |  | n/a
2023-12-05 09:41 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.439 | Pass |  | n/a
2023-12-05 09:48 | stmstation3@avive.life | W1153 | Agree | Liza Ordillo | 2310190950 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-05 09:52 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310190980 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-05 10:02 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-05 10:04 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190946 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 355 | Pass |  | N/A
2023-12-05 10:06 | stmstation1@avive.life | W1153 | Agree | Liza Ordillo | 2310191147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.427 | Pass |  | n/a
2023-12-05 10:15 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.402 | Pass |  | n/a
2023-12-05 10:15 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-05 10:17 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310191084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.357 | Pass |  | N/A
2023-12-05 10:23 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.355 | Pass |  | N/A
2023-12-05 10:28 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310191068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.346 | Pass |  | N/A
2023-12-05 10:38 | stmstation2@avive.life | W1153 | Agree | Liza Ordillo | 2310191139 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.429 | Pass |  | n/a
2023-12-05 10:45 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.34 | Pass |  | N/A
2023-12-05 10:54 | stmstation1@avive.life | W1153 | Agree | Leonora Cruz | 2310191165 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-05 10:55 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310190854 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-05 10:56 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190994 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.354 | Pass |  | N/A
2023-12-05 10:57 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 350 | Pass |  | N/A
2023-12-05 11:03 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.356 | Pass |  | n/a
2023-12-05 11:06 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | N/A
2023-12-05 11:07 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190850 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.354 | Pass |  | N/A
2023-12-05 11:07 | stmstation1@avive.life | W1153 | Agree | Leonora Cruz | 2310190563 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-05 11:13 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-05 11:13 | stmstation1@avive.life | W1153 | Agree | Leonora Cruz | 2310191062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-05 11:16 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190577 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.348 | Pass |  | N/A
2023-12-05 11:22 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.346 | Pass |  | N/A
2023-12-05 11:23 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310190574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-05 11:23 | stmstation1@avive.life | W1153 | Agree | Leonora Cruz | 2310190962 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-05 11:24 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310191047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.355 | Pass |  | N/A
2023-12-05 12:18 | stmstation4@avive.life | W1153 | Agree | Leonora Cruz | 2310190694 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.35 | Pass |  | N/A
2023-12-05 12:18 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-05 12:22 | stmstation1@avive.life | W1153 | Agree | Leonora Cruz | 2310190681 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-05 12:27 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.35 | Pass |  | N/A
2023-12-05 12:32 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310190682 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-05 12:34 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.351 | Pass |  | N/A
2023-12-05 12:34 | stmstation1@avive.life | W1153 | Agree | Leonora Cruz | 2310191020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-05 12:43 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190990 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.347 | Pass |  | N/A
2023-12-05 12:46 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-05 12:50 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190989 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.35 | Pass |  | N/A
2023-12-05 12:52 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-05 13:01 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190653 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.347 | Pass |  | N/A
2023-12-05 13:03 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | N/A
2023-12-05 13:04 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-05 13:10 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191007 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.355 | Pass |  | N/A
2023-12-05 13:13 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190648 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | N/A
2023-12-05 13:28 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.359 | Pass |  | N/A
2023-12-05 13:29 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190663 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.344 | Pass |  | N/A
2023-12-05 13:34 | stmstation1@avive.life | W1153 | Agree | Violeta Ayala | 2310191152 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.346 | Pass |  | N/A
2023-12-05 13:41 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310191097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.36 | Pass |  | N/A
2023-12-05 13:42 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.345 | Pass |  | N/A
2023-12-05 13:44 | stmstation1@avive.life | W1153 | Agree | Violeta Ayala | 2310191161 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.35 | Pass |  | N/A
2023-12-05 14:04 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190973 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.344 | Pass |  | N/A
2023-12-05 14:13 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190942 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.346 | Pass |  | N/A
2023-12-05 14:14 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190978 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.349 | Pass |  | N/A
2023-12-05 14:41 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190876 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.342 | Pass |  | N/A
2023-12-05 14:42 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190981 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.347 | Pass |  | N/A
2023-12-05 14:43 | stmstation1@avive.life | W1153 | Agree | Leonora Cruz | 2310190944 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-05 14:45 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191033 | Pass | Fail | [10 empty] | screen won't turn on
2023-12-05 14:52 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310190895 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.361 | Pass |  | N/A
2023-12-05 14:54 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190899 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.345 | Pass |  | N/A
2023-12-05 15:02 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310190945 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-05 15:06 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190896 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.353 | Pass |  | N/A
2023-12-05 15:08 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-05 15:15 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310191054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 432 | Pass | Pass | 0.351 | Pass |  | N/A
2023-12-05 15:19 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310190674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-05 15:26 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310190650 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-05 15:26 | stmstation4@avive.life | W1153 | Agree | Violeta Ayala | 2310190667 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.344 | Pass |  | N/A
2023-12-05 15:26 | stmstation3@avive.life | W1153 | Agree | Violeta Ayala | 2310191045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.35 | Pass |  | N/A
2023-12-05 15:33 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.361 | Pass |  | n/a
2023-12-05 15:40 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.358 | Pass |  | n/a
2023-12-05 15:46 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-05 15:52 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-05 15:59 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191167 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-05 16:13 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-05 16:22 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-05 16:32 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-05 16:38 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310190943 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-05 17:07 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310191075 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-05 17:21 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310190856 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-05 17:26 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310190964 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-05 17:27 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-05 17:34 | stmstation1@avive.life | W1153 | Agree | Leonora Cruz | 2310190898 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.354 | Pass |  | gps
2023-12-05 17:36 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310190993 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-05 17:39 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310191105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-05 17:40 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310190645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-05 17:56 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310191136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-05 17:59 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310190860 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-05 18:06 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-05 18:09 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310190985 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 8.98 | Fail | Fail sleep mode
2023-12-05 18:13 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310191196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-05 18:16 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310191142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-05 18:16 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310191141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-05 18:22 | stmstation2@avive.life | W1153 | Agree | Leonora Cruz | 2310190625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-05 18:25 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310190647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-05 18:28 | stmstation1@avive.life | W1153 | Agree | Leonora Cruz | 2310191023 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.359 | Pass |  | n/a
2023-12-05 18:33 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310190562 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-05 18:36 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310190627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-05 18:41 | stmstation4@avive.life | W1153 | Agree | Christina Cruz | 2310190965 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-06 06:20 | stmstation3@avive.life | W1153 | Agree | Precious Ungco | 2310190991 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-06 06:22 | stmstation2@avive.life | W1153 | Agree | Precious Ungco | 2310190576 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.337 | Pass |  | n/a
2023-12-06 06:27 | stmstation1@avive.life | W1153 | Agree | Precious Ungco | 2310190649 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-06 06:35 | stmstation2@avive.life | W1153 | Agree | Precious Ungco | 2310190408 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-06 06:45 | stmstation2@avive.life | W1153 | Agree | Precious Ungco | 2310191091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.356 | Pass |  | n/a
2023-12-06 06:52 | stmstation2@avive.life | W1153 | Agree | Precious Ungco | 2310191081 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-06 06:56 | stmstation3@avive.life | W1153 | Agree | Precious Ungco | 2310190999 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-06 07:05 | stmstation1@avive.life | W1153 | Agree | Precious Ungco | 2310191085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-06 07:05 | stmstation2@avive.life | W1153 | Agree | Precious Ungco | 2310190992 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-06 07:09 | stmstation3@avive.life | W1153 | Agree | Precious Ungco | 2310190573 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-06 07:15 | stmstation1@avive.life | W1153 | Agree | Precious Ungco | 2310191112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-06 08:17 | stmstation1@avive.life | W1153 | Agree | Precious Ungco | 2310191119 | Pass | Pass | Pass | Pass | Fail | [7 empty] | not connected to wifi
2023-12-06 08:21 | stmstation2@avive.life | W1153 | Agree | Precious Ungco | 2310190853 | Pass | Pass | Pass | Fail | [8 empty] | GPS sign doesn't show up. and wifi not connected
2023-12-07 09:46 | stmstation3@avive.life | W1153 | Agree | Precious Ungco | 2310191120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 1.5 | Fail | High Current during sleep
2023-12-07 09:52 | stmstation3@avive.life | W1153 | Agree | Christina Cruz | 2310191040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | Retested
2023-12-07 10:07 | stmstation2@avive.life | W1153 | Agree | Ananya Jain | 2310190653 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.365 | Pass |  | Retested
2023-12-13 08:17 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.337 | Pass |  | n/a
2023-12-13 08:22 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-13 08:30 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-13 08:36 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-13 08:39 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-13 08:43 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-13 08:49 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080041 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-13 08:52 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-13 08:56 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.337 | Pass |  | n/a
2023-12-13 09:06 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-13 09:12 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-13 09:17 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-13 09:19 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080023 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-13 09:23 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080018 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-13 09:29 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-13 09:33 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.321 | Pass |  | n/a
2023-12-13 09:36 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-13 09:44 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-13 09:46 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-13 09:49 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.337 | Pass |  | n/a
2023-12-13 09:53 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-13 10:26 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-13 10:29 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-13 10:40 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.339 | Pass |  | n/a
2023-12-13 11:33 | stmstation1@avive.life | W1156 | Agree | Christina Cruz | 2312080019 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-13 13:27 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-13 13:29 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-13 13:31 | stmstation2@avive.life | W1156 | Agree | Christina Cruz | 2312080033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-13 13:34 | stmstation1@avive.life | W1156 | Agree | Christina Cruz | 2312080039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-13 13:39 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-13 13:42 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-13 13:46 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-13 13:49 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-13 14:06 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-13 14:07 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.33 | Pass |  | n/a
2023-12-13 14:23 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-13 14:26 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-13 14:33 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-13 16:05 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-13 16:07 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-13 16:12 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-13 16:14 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-13 16:22 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-13 16:24 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-13 16:30 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-13 16:36 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080024 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-14 13:06 | stmstation2@avive.life | W1156 | Agree | Christina | 2312080183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-14 13:36 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.359 | Pass |  | n/a
2023-12-14 13:45 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080167 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-14 13:47 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080259 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.439 | Pass |  | n/a
2023-12-14 13:57 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.409 | Pass |  | n/a
2023-12-14 14:07 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.396 | Pass |  | n/a
2023-12-14 14:16 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-14 14:50 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080113 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.413 | Pass |  | n/a
2023-12-14 14:50 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080236 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-14 15:00 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.406 | Pass |  | n/a
2023-12-14 15:02 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080231 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-14 15:14 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.402 | Pass |  | n/a
2023-12-14 15:18 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080234 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.427 | Pass |  | n/a
2023-12-14 15:23 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-14 15:26 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.412 | Pass |  | n/a
2023-12-14 15:37 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-14 16:08 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080260 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-14 16:14 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080233 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-14 16:41 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-14 16:43 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080257 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-14 16:52 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-14 16:58 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-14 17:00 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080158 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-15 06:40 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080197 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.424 | Pass |  | n/a
2023-12-15 06:46 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.413 | Pass |  | n/a
2023-12-15 06:52 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-15 06:57 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080162 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-15 07:01 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-15 07:06 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.376 | Pass |  | n/a
2023-12-15 07:09 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.378 | Pass |  | n/a
2023-12-15 07:13 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080191 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.356 | Pass |  | n/a
2023-12-15 07:22 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.425 | Pass |  | n/a
2023-12-15 07:27 | stmstation4@avive.life | W1156 | Agree | Liza Ordillo | 2312080171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 07:29 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080227 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-15 07:37 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-15 07:40 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 07:56 | stmstation4@avive.life | W1156 | Agree | Liza Ordillo | 2312080088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-15 08:04 | stmstation4@avive.life | W1156 | Agree | Liza Ordillo | 2312080082 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.43 | Pass |  | n/a
2023-12-15 08:15 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080230 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.433 | Pass |  | n/a
2023-12-15 08:16 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 08:18 | stmstation4@avive.life | W1156 | Agree | Liza Ordillo | 2312080177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.375 | Pass |  | n/a
2023-12-15 08:29 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080075 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 08:30 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-15 08:36 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-15 08:37 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-15 08:44 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-15 08:45 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.36 | Pass |  | n/a
2023-12-15 08:51 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-15 08:52 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-15 08:58 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.403 | Pass |  | n/a
2023-12-15 08:59 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.415 | Pass |  | n/a
2023-12-15 09:02 | stmstation2@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 424 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-15 09:05 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-15 09:10 | stmstation2@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-15 09:13 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-15 09:14 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080083 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-15 09:56 | stmstation2@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-15 10:29 | stmstation1@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080070 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | CELLULAR RECEPTION DOESNT SHOWS UP/ DEVICE SELFTEST 2RISK
2023-12-15 10:32 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-15 10:34 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080232 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-15 10:44 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-15 10:45 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.337 | Pass |  | n/a
2023-12-15 10:46 | stmstation1@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-15 10:48 | stmstation2@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-15 10:54 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-15 10:55 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 10:56 | stmstation1@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 10:57 | stmstation2@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-15 11:05 | stmstation2@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-15 11:06 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080159 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-15 11:08 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-15 11:08 | stmstation1@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080138 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-15 11:28 | stmstation2@avive.life | W1156 | Agree | PRECIOUS UNGCO | 2312080067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-15 11:29 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-15 11:51 | stmstation3@avive.life | W1156 | Agree | Precious Ungco | 2312080112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-15 11:52 | stmstation4@avive.life | W1156 | Agree | Precious Ungco | 2312080241 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-15 12:05 | stmstation1@avive.life | W1156 | Agree | Valeria Manzo | 2312080156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-15 12:06 | stmstation3@avive.life | W1156 | Agree | Valeria Manzo | 2312080063 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-15 12:07 | stmstation4@avive.life | W1156 | Agree | Valeria Manzo | 2312080215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-15 12:18 | stmstation4@avive.life | W1156 | Agree | Valeria Manzo | 2312080239 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-15 12:19 | stmstation3@avive.life | W1156 | Agree | Valeria Manzo | 2312080053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-15 12:21 | stmstation2@avive.life | W1156 | Agree | Valeria Manzo | 2312080093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-15 12:23 | stmstation1@avive.life | W1156 | Agree | Valeria Manzo | 2312080178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-15 13:44 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.36 | Pass |  | n/a
2023-12-15 13:48 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-15 13:52 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-15 13:56 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-15 13:59 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-15 14:09 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080209 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-15 14:13 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 14:48 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080253 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-15 14:58 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.367 | Pass |  | n/a
2023-12-15 15:05 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-15 15:12 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080255 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-15 15:14 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080256 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 15:17 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-15 15:21 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 15:25 | stmstation2@avive.life | W1156 | Agree | Christina Cruz | 2312080254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 15:29 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-15 15:48 | stmstation2@avive.life | W1156 | Agree | Christina Cruz | 2312080118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-15 15:54 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080240 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-15 15:55 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080166 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-15 16:20 | stmstation2@avive.life | W1156 | Agree | Christina Cruz | 2312080054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-15 16:24 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080139 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 16:26 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080252 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 16:30 | stmstation2@avive.life | W1156 | Agree | Christina Cruz | 2312080052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.36 | Pass |  | n/a
2023-12-15 16:33 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-15 16:36 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-15 16:41 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080103 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-15 16:42 | stmstation2@avive.life | W1156 | Agree | Christina Cruz | 2312080069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-15 16:45 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 06:40 | stmstation4@avive.life | W1156 | Agree | Liza Ordillo | 2312080161 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.412 | Pass |  | n/a
2023-12-18 06:49 | stmstation4@avive.life | W1156 | Agree | Liza Ordillo | 2312080195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 06:50 | stmstation3@avive.life | W1156 | Agree | Liza Ordillo | 2312080050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.37 | Pass |  | n/a
2023-12-18 06:57 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.425 | Pass |  | n/a
2023-12-18 06:58 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.417 | Pass |  | n/a
2023-12-18 07:05 | stmstation3@avive.life | W1156 | Agree | Liza Ordillo | 2312080149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.38 | Pass |  | n/a
2023-12-18 07:06 | stmstation4@avive.life | W1156 | Agree | Liza Ordillo | 2312080076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.41 | Pass |  | n/a
2023-12-18 07:15 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080130 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 07:17 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-18 07:31 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-18 07:31 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-18 07:32 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-18 07:36 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-18 07:43 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080200 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-18 07:44 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-18 07:44 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-18 07:56 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-18 07:57 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-18 08:00 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080238 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.366 | Pass |  | n/a
2023-12-18 08:17 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-18 08:28 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080243 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-18 08:33 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-18 08:33 | stmstation3@avive.life | W1156 | Agree | Violeta Ayala | 2312080193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.348 | Pass |  | N/A
2023-12-18 08:39 | stmstation4@avive.life | W1156 | Agree | Violeta Ayala | 2312080116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.338 | Pass |  | N/A
2023-12-18 08:49 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.356 | Pass |  | n/a
2023-12-18 08:50 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080242 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 08:50 | stmstation3@avive.life | W1156 | Agree | Violeta Ayala | 2312080090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-18 08:51 | stmstation4@avive.life | W1156 | Agree | Violeta Ayala | 2312080125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-18 08:58 | stmstation3@avive.life | W1156 | Agree | Violeta Ayala | 2312080160 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.343 | Pass |  | N/A
2023-12-18 08:59 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 09:06 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-18 09:06 | stmstation4@avive.life | W1156 | Agree | Violeta Ayala | 2312080094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-18 09:14 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080212 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-18 09:32 | stmstation3@avive.life | W1156 | Agree | Violeta Ayala | 2312080085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.346 | Pass |  | N/A
2023-12-18 09:34 | stmstation4@avive.life | W1156 | Agree | Violeta Ayala | 2312080058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.336 | Pass |  | n/a
2023-12-18 09:39 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080152 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.356 | Pass |  | N/A
2023-12-18 09:39 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080073 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-18 11:18 | stmstation3@avive.life | W1156 | Agree | Violeta Ayala | 2312080213 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.349 | Pass |  | N/A
2023-12-18 11:22 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080202 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-18 11:22 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 11:26 | stmstation4@avive.life | W1156 | Agree | Violeta Ayala | 2312080079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 11:31 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-18 11:32 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080172 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-18 11:42 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080246 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-18 11:43 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080245 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-18 11:53 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.358 | Pass |  | n/a
2023-12-18 12:04 | stmstation4@avive.life | W1156 | Agree | Violeta Ayala | 2312080174 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-18 12:09 | stmstation3@avive.life | W1156 | Agree | Violeta Ayala | 2312080225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.347 | Pass |  | N/A
2023-12-18 13:59 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080066 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-18 14:00 | stmstation1@avive.life | W1156 | Agree | Leonora Cruz | 2312080170 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 14:04 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080201 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-18 14:08 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-18 14:13 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-18 14:15 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.339 | Pass |  | n/a
2023-12-18 14:20 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080228 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-18 14:24 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080235 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-18 14:28 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-18 14:29 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080128 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-18 14:35 | stmstation2@avive.life | W1156 | Agree | Leonora Cruz | 2312080119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-18 14:37 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080168 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 14:41 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 14:43 | stmstation2@avive.life | W1156 | Agree | Christina Cruz | 2312080144 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-18 14:46 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-18 14:47 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 14:54 | stmstation4@avive.life | W1156 | Agree | Christina Cruz | 2312080145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-18 14:59 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-19 06:41 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.386 | Pass |  | n/a
2023-12-19 06:51 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.374 | Pass |  | n/a
2023-12-19 06:57 | stmstation3@avive.life | W1156 | Agree | Liza Ordillo | 2312080086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.375 | Pass |  | n/a
2023-12-19 07:01 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080077 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.383 | Pass |  | n/a
2023-12-19 07:11 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.366 | Pass |  | n/a
2023-12-19 07:16 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.389 | Pass |  | n/a
2023-12-19 07:26 | stmstation3@avive.life | W1156 | Agree | Liza Ordillo | 2312080175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-19 07:29 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080165 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.435 | Pass |  | n/a
2023-12-19 07:33 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-19 08:14 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-19 08:14 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080081 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.37 | Pass |  | n/a
2023-12-19 08:26 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.425 | Pass |  | n/a
2023-12-19 08:30 | stmstation3@avive.life | W1156 | Agree | Liza Ordillo | 2312080180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.39 | Pass |  | n/a
2023-12-19 08:38 | stmstation3@avive.life | W1156 | Agree | Liza Ordillo | 2312080176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-19 09:06 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-19 10:15 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080065 | Pass | Fail | [10 empty] | No display
2023-12-19 10:31 | stmstation3@avive.life | W1156 | Agree | Christina Cruz | 2312080190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Fail |  |  | No sim display
2023-12-20 06:37 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080424 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-20 06:38 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080470 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 06:45 | stmstation4@avive.life | W1156 | Agree | Liza Ordillo | 2312080463 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-20 06:46 | stmstation3@avive.life | W1156 | Agree | Liza Ordillo | 2312080397 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-20 07:00 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080395 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 07:03 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2312080337 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-20 07:13 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080388 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-20 07:18 | stmstation1@avive.life | W1156 | Agree | Liza Ordillo | 2310190513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 07:18 | stmstation2@avive.life | W1156 | Agree | Liza Ordillo | 2312080335 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.378 | Pass |  | n/a
2023-12-20 07:20 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080339 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-20 07:21 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080457 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 07:32 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080464 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.348 | Pass |  | N/A
2023-12-20 07:37 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080432 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-20 07:38 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080431 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-20 07:44 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080414 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.348 | Pass |  | N/A
2023-12-20 07:46 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080420 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 07:46 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080465 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.336 | Pass |  | n/a
2023-12-20 07:52 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080422 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.345 | Pass |  | N/A
2023-12-20 07:55 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080426 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-20 07:56 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080428 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-20 07:56 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080412 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.345 | Pass |  | N/A
2023-12-20 08:01 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080476 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.347 | Pass |  | N/A
2023-12-20 08:04 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080459 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.35 | Pass |  | N/A
2023-12-20 08:05 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080462 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-20 08:08 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080317 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-20 08:11 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080413 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 08:12 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080469 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | N/A
2023-12-20 08:14 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080338 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-20 08:15 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080455 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.345 | Pass |  | N/A
2023-12-20 08:20 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080418 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.342 | Pass |  | N/A
2023-12-20 08:23 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080466 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-20 08:26 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080454 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.341 | Pass |  | N/A
2023-12-20 08:27 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080471 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 08:28 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080453 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-20 08:32 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080425 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.349 | Pass |  | N/A
2023-12-20 08:36 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080398 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 3 | Fail | value to high
2023-12-20 08:37 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080461 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-20 08:39 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080460 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.344 | Pass |  | N/A
2023-12-20 09:05 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080430 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.347 | Pass |  | N/A
2023-12-20 09:08 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080458 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 09:08 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080399 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 09:11 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080421 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | N/A
2023-12-20 09:15 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080417 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 09:17 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2310190506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-20 09:24 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080392 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-20 09:25 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080321 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 09:31 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080468 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 311 | Pass | Pass | 0.35 | Pass |  | N/A
2023-12-20 09:35 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080401 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.344 | Pass |  | N/A
2023-12-20 09:42 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080416 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.353 | Pass |  | N/A
2023-12-20 09:43 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080415 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-20 09:44 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080419 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-20 09:45 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080475 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.344 | Pass |  | N/A
2023-12-20 09:51 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080427 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.346 | Pass |  | N/A
2023-12-20 09:52 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080423 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-20 09:52 | stmstation4@avive.life | W1156 | Agree | Leonora Cruz | 2312080456 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-20 09:56 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080429 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.346 | Pass |  | N/A
2023-12-20 10:16 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080449 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.347 | Pass |  | N/A
2023-12-20 10:19 | stmstation2@avive.life | W1156 | Agree | Violeta Ayala | 2312080310 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.346 | Pass |  | N/A
2023-12-20 10:30 | stmstation1@avive.life | W1156 | Agree | Violeta Ayala | 2312080450 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.355 | Pass |  | N/A
2023-12-20 11:36 | stmstation3@avive.life | W1156 | Agree | Leonora Cruz | 2312080316 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-20 13:57 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080383 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 13:57 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080325 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 13:59 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080389 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-20 14:06 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080336 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 14:08 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080276 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 14:56 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080382 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-20 14:57 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080396 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-20 14:59 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080313 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-20 15:04 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080308 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-20 15:05 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080350 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-20 15:10 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080447 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 15:13 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080306 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-20 15:15 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080343 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-20 15:20 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080409 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-20 15:24 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080318 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-20 15:24 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080271 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 15:26 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080277 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-20 15:28 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080297 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 15:32 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080328 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-20 15:33 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080445 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-20 15:35 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080334 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-20 15:38 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-20 15:39 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080330 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 15:40 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080406 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-20 15:42 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.339 | Pass |  | n/a
2023-12-20 15:45 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080371 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-20 15:45 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080391 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-20 15:46 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080324 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-20 15:50 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080373 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 15:55 | stmstation2@avive.life | W1169 | Agree | Ireri Moreno | 2312080322 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-20 15:55 | stmstation1@avive.life | W1169 | Agree | Ireri Moreno | 2312080315 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 16:02 | stmstation3@avive.life | W1169 | Agree | ireri moreno | 2312080384 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-20 16:02 | stmstation4@avive.life | W1169 | Agree | ireri moreno | 2312080448 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 16:05 | stmstation1@avive.life | W1169 | Agree | ireri moreno | 2312080390 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-20 16:06 | stmstation2@avive.life | W1169 | Agree | Ireri Moreno | 2312080379 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-20 16:13 | stmstation2@avive.life | W1169 | Agree | Ireri Moreno | 2312080408 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 16:15 | stmstation1@avive.life | W1169 | Agree | ireri moreno | 2312080381 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-20 16:20 | stmstation2@avive.life | W1169 | Agree | Ireri Moreno | 2312080341 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-20 16:21 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080404 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-20 16:22 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080265 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-20 16:23 | stmstation1@avive.life | W1169 | Agree | ireri moreno | 2312080368 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 16:27 | stmstation2@avive.life | W1169 | Agree | Ireri Moreno | 2312080451 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-20 16:28 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080327 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-20 16:30 | stmstation1@avive.life | W1169 | Agree | ireri moreno | 2312080267 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 16:31 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080348 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-20 16:34 | stmstation2@avive.life | W1169 | Agree | Ireri Moreno | 2312080356 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-20 16:36 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080291 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 16:38 | stmstation1@avive.life | W1169 | Agree | ireri moreno | 2312080311 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-20 16:42 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080304 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-20 16:43 | stmstation2@avive.life | W1169 | Agree | Ireri Moreno | 2312080372 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-20 16:44 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080288 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-20 16:46 | stmstation1@avive.life | W1169 | Agree | ireri moreno | 2312080281 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-20 16:48 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080378 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.337 | Pass |  | n/a
2023-12-20 16:49 | stmstation2@avive.life | W1169 | Agree | Ireri Moreno | 2312080387 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-20 16:51 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080293 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 16:54 | stmstation1@avive.life | W1169 | Agree | ireri moreno | 2312080370 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 16:55 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080307 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 16:55 | stmstation2@avive.life | W1169 | Agree | Ireri Moreno | 2312080312 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-20 16:58 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080345 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 17:01 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080305 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.334 | Pass |  | n/a
2023-12-20 17:02 | stmstation2@avive.life | W1169 | Agree | Ireri Moreno | 2312080393 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-20 17:05 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-20 17:08 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080405 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-20 17:16 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080319 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-20 17:16 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080407 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-20 17:24 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080275 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 17:24 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080344 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-20 17:31 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080266 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-20 17:33 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080301 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 17:40 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080314 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-20 17:40 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080296 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-20 18:01 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080333 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-20 18:03 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080394 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 18:05 | stmstation1@avive.life | W1169 | Agree | Christina Cruz | 2312080354 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-20 18:08 | stmstation2@avive.life | W1169 | Agree | Christina Cruz | 2312080493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-20 18:16 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 18:20 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080375 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.357 | Pass |  | n/a
2023-12-20 18:49 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080340 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-20 18:53 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080329 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-20 18:56 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080403 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-20 18:58 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080363 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-20 19:01 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080365 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.362 | Pass |  | n/a
2023-12-20 19:05 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080270 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-20 19:07 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080247 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-20 19:14 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080411 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-20 19:22 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080268 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-20 19:30 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080351 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-20 19:32 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080607 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-20 19:36 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080485 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-20 19:40 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080278 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-20 19:43 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080446 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-20 19:45 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080358 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-20 19:54 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-21 06:33 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080302 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 06:39 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080300 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-21 06:44 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080342 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-21 06:47 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-21 06:57 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080346 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.383 | Pass |  | n/a
2023-12-21 07:00 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-21 07:05 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080292 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-21 07:10 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080299 | Pass | Fail | [10 empty] | showing white pixels
2023-12-21 07:22 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080286 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.425 | Pass |  | n/a
2023-12-21 07:27 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080510 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.425 | Pass |  | n/a
2023-12-21 07:42 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-21 07:49 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080606 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-21 07:53 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080617 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-21 07:55 | stmstation3@avive.life | W1169 | Agree | Ireri Moreno | 2312080295 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-21 07:56 | stmstation4@avive.life | W1169 | Agree | Ireri Moreno | 2312080627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-21 07:57 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080366 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.361 | Pass |  | n/a
2023-12-21 08:02 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080595 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.403 | Pass |  | n/a
2023-12-21 08:06 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080362 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-21 08:12 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080602 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-21 08:20 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-21 08:23 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080650 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-21 08:27 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080562 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-21 08:31 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080400 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-21 08:36 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080624 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-21 08:38 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080563 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-21 08:42 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080653 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-21 08:44 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080613 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-21 08:49 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-21 08:52 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080587 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-21 08:55 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080604 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.426 | Pass |  | n/a
2023-12-21 08:55 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080651 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-21 09:01 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080646 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.358 | Pass |  | n/a
2023-12-21 09:06 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080652 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-21 09:07 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080648 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.359 | Pass |  | n/a
2023-12-21 09:09 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-21 09:17 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080287 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | no cell icon
2023-12-21 09:31 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080299 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-21 09:46 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080616 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 09:49 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080361 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-21 10:03 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080618 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-21 10:08 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080367 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-21 10:28 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080263 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-21 10:31 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 10:36 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080282 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.336 | Pass |  | n/a
2023-12-21 10:57 | stmstation3@avive.life | W1169 | Agree | Violeta Ayala | 2312080331 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.349 | Pass |  | N/A
2023-12-21 10:57 | stmstation4@avive.life | W1169 | Agree | Violeta Ayala | 2312080282 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.339 | Pass |  | n/a
2023-12-21 11:05 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080283 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.344 | Pass |  | N/A
2023-12-21 11:07 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.356 | Pass |  | N/A
2023-12-21 11:33 | stmstation3@avive.life | W1169 | Agree | Violeta Ayala | 2312080615 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 11:55 | stmstation4@avive.life | W1169 | Agree | Violeta Ayala | 2312080599 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-21 12:02 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | N/A
2023-12-21 12:03 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080332 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.349 | Pass |  | N/A
2023-12-21 12:07 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080369 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 12:10 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080551 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-21 12:12 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080364 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.351 | Pass |  | N/A
2023-12-21 12:12 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.348 | Pass |  | N/A
2023-12-21 12:21 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 12:22 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080326 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.351 | Pass |  | N/A
2023-12-21 12:23 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-21 12:26 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080279 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.337 | Pass |  | n/a
2023-12-21 12:27 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080248 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.351 | Pass |  | N/A
2023-12-21 12:46 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080594 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.346 | Pass |  | N/A
2023-12-21 12:46 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080622 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.346 | Pass |  | N/A
2023-12-21 12:55 | stmstation4@avive.life | W1169 | Agree | Precious Ungco | 2312080360 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 12:56 | stmstation3@avive.life | W1169 | Agree | Precious Ungco | 2312080349 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.356 | Pass |  | n/a
2023-12-21 12:56 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080386 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.356 | Pass |  | N/A
2023-12-21 12:57 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080261 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.351 | Pass |  | N/A
2023-12-21 13:02 | stmstation3@avive.life | W1169 | Agree | Precious Ungco | 2312080352 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-21 13:03 | stmstation4@avive.life | W1169 | Agree | Precious Ungco | 2312080601 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-21 13:09 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080605 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 344 | Pass |  | N/A
2023-12-21 13:09 | stmstation3@avive.life | W1169 | Agree | Precious Ungco | 2312080547 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-21 13:09 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080610 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 349 | Pass |  | N/A
2023-12-21 13:11 | stmstation4@avive.life | W1169 | Agree | Precious Ungco | 2312080250 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.331 | Pass |  | n/a
2023-12-21 13:20 | stmstation4@avive.life | W1169 | Agree | Precious Ungco | 2312080402 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-21 13:21 | stmstation3@avive.life | W1169 | Agree | Precious Ungco | 2312080626 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-21 13:24 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080573 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 13:25 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080598 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 13:27 | stmstation3@avive.life | W1169 | Agree | Precious Ungco | 2312080269 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-21 13:29 | stmstation4@avive.life | W1169 | Agree | Precious Ungco | 2312080584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 14:02 | stmstation3@avive.life | W1169 | Agree | Precious Ungco | 2312080566 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 14:06 | stmstation4@avive.life | W1169 | Agree | Christina Cruz | 2312080564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 14:16 | stmstation4@avive.life | W1169 | Agree | Precious Ungco | 2312080567 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.357 | Pass |  | n/a
2023-12-21 14:17 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080565 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-21 14:19 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.362 | Pass |  | n/a
2023-12-21 14:24 | stmstation4@avive.life | W1169 | Agree | Precious Ungco | 2312080569 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-21 14:25 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.356 | Pass |  | n/a
2023-12-21 14:27 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080655 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-21 14:32 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-21 14:33 | stmstation3@avive.life | W1169 | Agree | Ireri Moreno | 2312080298 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.339 | Pass |  | n/a
2023-12-21 14:34 | stmstation4@avive.life | W1169 | Agree | Precious Ungco | 2312080559 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-21 14:36 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-21 14:42 | stmstation3@avive.life | W1169 | Agree | Ireri Moreno | 2312080628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-21 14:42 | stmstation4@avive.life | W1169 | Agree | Ireri Moreno | 2312080545 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-21 14:44 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-21 14:49 | stmstation3@avive.life | W1169 | Agree | Ireri Moreno | 2312080497 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-21 14:52 | stmstation4@avive.life | W1169 | Agree | Ireri Moreno | 2312080631 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 14:52 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080533 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 14:54 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080514 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-21 14:55 | stmstation3@avive.life | W1169 | Agree | Ireri Moreno | 2312080579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-21 15:21 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080630 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-21 15:23 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080357 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-21 15:27 | stmstation3@avive.life | W1169 | Agree | Ireri Moreno | 2312080516 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-21 15:28 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-21 15:29 | stmstation4@avive.life | W1169 | Agree | Ireri Moreno | 2312080480 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-21 15:31 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080512 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-21 15:34 | stmstation3@avive.life | W1169 | Agree | Ireri Moreno | 2312080577 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-21 15:35 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080534 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-21 15:36 | stmstation4@avive.life | W1169 | Agree | Ireri Moreno | 2312080219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 437 | Pass | Pass | 0.336 | Pass |  | n/a
2023-12-21 15:37 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080527 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-21 15:43 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080495 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-21 15:43 | stmstation4@avive.life | W1169 | Agree | Ireri Moreno | 2312080525 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-21 15:43 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080435 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-21 15:47 | stmstation3@avive.life | W1169 | Agree | Ireri Moreno | 2312080524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-21 15:49 | stmstation4@avive.life | W1169 | Agree | Ireri Moreno | 2312080532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-21 15:50 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.339 | Pass |  | n/a
2023-12-21 15:52 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-22 06:35 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080486 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-22 06:38 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080662 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-22 06:42 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080644 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.369 | Pass |  | n/a
2023-12-22 06:55 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080640 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-22 06:55 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080498 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-22 07:01 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080500 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.357 | Pass |  | n/a
2023-12-22 07:03 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080576 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-22 07:11 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080478 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.368 | Pass |  | n/a
2023-12-22 07:11 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.337 | Pass |  | n/a
2023-12-22 07:19 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080550 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.337 | Pass |  | n/a
2023-12-22 07:20 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080581 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-22 07:29 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080552 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-22 07:29 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-22 07:35 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080355 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-22 07:37 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.39 | Pass |  | n/a
2023-12-22 07:43 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080509 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.387 | Pass |  | n/a
2023-12-22 07:44 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080511 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-22 07:49 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080603 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.358 | Pass |  | n/a
2023-12-22 07:51 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080620 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.375 | Pass |  | n/a
2023-12-22 08:00 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-22 08:01 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080347 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-22 08:05 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-22 08:09 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080639 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.357 | Pass |  | n/a
2023-12-22 08:12 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080600 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.404 | Pass |  | n/a
2023-12-22 08:19 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080440 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-22 08:28 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.385 | Pass |  | n/a
2023-12-22 08:28 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080590 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-22 08:36 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.365 | Pass |  | n/a
2023-12-22 08:50 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-22 08:51 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-22 09:00 | stmstation2@avive.life | W1169 | Agree | Liza Ordillo | 2312080570 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-22 09:02 | stmstation1@avive.life | W1169 | Agree | Liza Ordillo | 2312080661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.425 | Pass |  | n/a
2023-12-22 09:06 | stmstation4@avive.life | W1169 | Agree | Liza Ordillo | 2312080608 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.436 | Pass |  | n/a
2023-12-22 09:08 | stmstation3@avive.life | W1169 | Agree | Liza Ordillo | 2312080515 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-22 09:17 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080477 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.343 | Pass |  | N/A
2023-12-22 09:20 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-22 09:21 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080546 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-22 09:25 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080518 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-22 09:26 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-22 09:27 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080660 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-22 09:51 | stmstation1@avive.life | W1169 | Agree | Christina Cruz | 2312080540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.359 | Pass |  | n/a
2023-12-22 09:52 | stmstation2@avive.life | W1169 | Agree | Christina Cruz | 2312080649 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.345 | Pass |  | N/A
2023-12-22 09:58 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080443 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-22 10:01 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080438 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-22 10:04 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080472 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-22 10:14 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080433 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-22 10:16 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-22 10:19 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-22 10:19 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080482 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.355 | Pass |  | N/A
2023-12-22 10:26 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080437 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-22 10:27 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080442 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-22 10:31 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.357 | Pass |  | n/a
2023-12-22 10:31 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-22 10:40 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-22 10:42 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080479 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-22 10:51 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080441 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-22 10:53 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080529 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.357 | Pass |  | n/a
2023-12-22 10:57 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080434 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-22 11:00 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080531 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-22 11:01 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080642 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-22 11:14 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080571 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-22 11:15 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080664 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-22 11:16 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080439 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-22 11:17 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080494 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.34 | Pass |  | n/a
2023-12-22 11:25 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080623 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-22 11:28 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080536 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-22 11:31 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-22 11:34 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080561 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-22 11:35 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080585 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-22 11:38 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080376 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-22 11:40 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-22 11:43 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-22 11:43 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080521 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-22 11:46 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080619 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-22 11:47 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.352 | Pass |  | n/a
2023-12-22 11:49 | stmstation4@avive.life | W1169 | Agree | Leonora Cruz | 2312080519 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-22 11:51 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080530 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-22 11:52 | stmstation3@avive.life | W1169 | Agree | Leonora Cruz | 2312080632 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-22 11:55 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080659 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-22 11:59 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080549 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-22 12:03 | stmstation4@avive.life | W1169 | Agree | Ireri Moreno | 2312080572 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-22 12:09 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.354 | Pass |  | n/a
2023-12-22 12:09 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080596 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-22 12:11 | stmstation4@avive.life | W1169 | Agree | Ireri Moreno | 2312080656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-22 12:15 | stmstation3@avive.life | W1169 | Agree | Ireri Moreno | 2312080657 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-22 12:17 | stmstation2@avive.life | W1169 | Agree | Valeria Manzo | 2312080499 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-22 12:17 | stmstation1@avive.life | W1169 | Agree | Valeria Manzo | 2312080560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-22 12:18 | stmstation4@avive.life | W1169 | Agree | Ireri Moreno | 2312080517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.344 | Pass |  | n/a
2023-12-22 13:13 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080578 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-22 13:16 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080520 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-22 13:30 | stmstation4@avive.life | W1169 | Agree | Valeria Manzo | 2312080538 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-22 13:31 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080535 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.355 | Pass |  | n/a
2023-12-22 13:34 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080444 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.343 | Pass |  | n/a
2023-12-22 13:38 | stmstation4@avive.life | W1169 | Agree | Valeria Manzo | 2312080526 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-22 13:38 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080556 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.342 | Pass |  | n/a
2023-12-22 13:40 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080575 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.341 | Pass |  | n/a
2023-12-22 13:44 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080294 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-22 13:46 | stmstation4@avive.life | W1169 | Agree | Valeria Manzo | 2312080583 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.348 | Pass |  | n/a
2023-12-22 13:50 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-22 13:53 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080593 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.346 | Pass |  | n/a
2023-12-22 13:54 | stmstation4@avive.life | W1169 | Agree | Valeria Manzo | 2312080582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.338 | Pass |  | n/a
2023-12-22 14:02 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080592 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.347 | Pass |  | n/a
2023-12-22 14:03 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.353 | Pass |  | n/a
2023-12-22 14:06 | stmstation4@avive.life | W1169 | Agree | Valeria Manzo | 2312080621 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.339 | Pass |  | n/a
2023-12-22 14:08 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080542 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-22 14:09 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080568 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.349 | Pass |  | n/a
2023-12-22 14:13 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080589 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.35 | Pass |  | n/a
2023-12-22 14:19 | stmstation1@avive.life | W1169 | Agree | Violeta Ayala | 2312080580 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.351 | Pass |  | n/a
2023-12-22 14:21 | stmstation3@avive.life | W1169 | Agree | Valeria Manzo | 2312080484 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | no cell and sim (tested 3 times 15+ min)
2023-12-22 14:23 | stmstation2@avive.life | W1169 | Agree | Violeta Ayala | 2312080274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Fail |  |  | SIM card not detected.
2023-12-22 14:24 | stmstation4@avive.life | W1169 | Agree | Valeria Manzo | 2312080654 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | no cell no sim
2023-12-22 15:31 | stmstation3@avive.life | W1169 | Agree | Christina Cruz | 2312080537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2023-12-28 12:40 | stmstation1@avive.life | W1169 | Agree | Jacob Bartee | 2312080609 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.354 | Pass |  | N/A
2024-01-05 11:09 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080939 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-05 11:15 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080436 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.381 | Pass |  | n/a
2024-01-05 11:23 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312180097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.402 | Pass |  | n/a
2024-01-05 11:24 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.371 | Pass |  | n/a
2024-01-05 11:34 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080780 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-05 11:35 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080977 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.378 | Pass |  | n/a
2024-01-05 11:36 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-05 11:42 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080958 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-05 11:44 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312180060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-05 11:47 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080912 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-05 11:56 | stmstation3@avive.life | W1175 | Agree | PRECIOUS UNGCO | 2312080950 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.375 | Pass |  | N/A
2024-01-05 11:59 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080986 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-05 12:02 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080924 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.394 | Pass |  | n/a
2024-01-05 12:05 | stmstation3@avive.life | W1175 | Agree | PRECIOUS UNGCO | 2312180105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-05 12:06 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080911 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-05 12:12 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080976 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.354 | Pass |  | n/a
2024-01-05 12:14 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080919 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-05 12:19 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312180102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.381 | Pass |  | n/a
2024-01-05 12:23 | stmstation3@avive.life | W1175 | Agree | PRECIOUS UNGCO | 2312080981 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.338 | Pass |  | N/A
2024-01-05 12:25 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080452 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-05 12:25 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080783 | Pass | Fail | [10 empty] | avive screen is blue
2024-01-05 12:35 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080884 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-05 12:36 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080971 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-05 12:52 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080988 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-05 12:55 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080984 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-05 13:02 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080985 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-05 13:08 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080989 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-05 13:11 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080899 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.399 | Pass |  | n/a
2024-01-05 13:15 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080888 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-05 13:18 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080897 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-05 13:22 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080900 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-05 13:25 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312180091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-05 13:29 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312180113 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-05 13:32 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080887 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-05 13:42 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080972 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-05 13:48 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080974 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.358 | Pass |  | n/a
2024-01-05 13:50 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312180117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-05 13:54 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080918 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-05 14:02 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312180101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-05 14:05 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080904 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-05 14:08 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080905 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-05 14:20 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312180104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-05 14:27 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080907 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-05 14:31 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312180089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-05 14:32 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080952 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-08 07:44 | stmstation2@avive.life | W1175 | Agree | Violeta Ayala | 2312080886 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-08 07:45 | stmstation1@avive.life | W1175 | Agree | Violeta Ayala | 2312080868 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-08 07:55 | stmstation3@avive.life | W1175 | Agree | Violeta Ayala | 2312080966 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-08 07:58 | stmstation4@avive.life | W1175 | Agree | Violeta Ayala | 2312080906 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-08 08:03 | stmstation1@avive.life | W1175 | Agree | Violeta Ayala | 2312080909 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-08 08:04 | stmstation2@avive.life | W1175 | Agree | Violeta Ayala | 2312080874 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-08 08:09 | stmstation3@avive.life | W1175 | Agree | Violeta Ayala | 2312080779 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-08 08:11 | stmstation4@avive.life | W1175 | Agree | Violeta Ayala | 2312080785 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-08 08:18 | stmstation1@avive.life | W1175 | Agree | Violeta Ayala | 2312080902 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-08 08:19 | stmstation2@avive.life | W1175 | Agree | Violeta Ayala | 2312080781 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-08 08:32 | stmstation4@avive.life | W1175 | Agree | Violeta Ayala | 2312080890 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.33 | Pass |  | n/a
2024-01-08 08:32 | stmstation3@avive.life | W1175 | Agree | Violeta Ayala | 2312080870 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-08 08:35 | stmstation2@avive.life | W1175 | Agree | Violeta Ayala | 2312080789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-08 08:37 | stmstation1@avive.life | W1175 | Agree | Violeta Ayala | 2312080784 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-08 08:45 | stmstation2@avive.life | W1175 | Agree | Violeta Ayala | 2312080963 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-08 08:49 | stmstation4@avive.life | W1175 | Agree | Violeta Ayala | 2312180116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-08 08:50 | stmstation3@avive.life | W1175 | Agree | Violeta Ayala | 2312180084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-08 08:50 | stmstation1@avive.life | W1175 | Agree | Violeta Ayala | 2312180098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-08 08:57 | stmstation2@avive.life | W1175 | Agree | Violeta Ayala | 2312080973 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-08 09:00 | stmstation1@avive.life | W1175 | Agree | Violeta Ayala | 2312080948 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-08 09:02 | stmstation4@avive.life | W1175 | Agree | Violeta Ayala | 2312080987 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-08 09:03 | stmstation3@avive.life | W1175 | Agree | Violeta Ayala | 2312080994 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-08 09:13 | stmstation4@avive.life | W1175 | Agree | Violeta Ayala | 2312080879 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-08 09:14 | stmstation1@avive.life | W1175 | Agree | Violeta Ayala | 2312180100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-08 09:17 | stmstation3@avive.life | W1175 | Agree | Violeta Ayala | 2312080992 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-08 09:24 | stmstation2@avive.life | W1175 | Agree | Violeta Ayala | 2312080885 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-08 09:26 | stmstation1@avive.life | W1175 | Agree | Violeta Ayala | 2312080983 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-08 09:51 | stmstation3@avive.life | W1175 | Agree | Violeta Ayala | 2312080892 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-08 09:51 | stmstation4@avive.life | W1175 | Agree | Violeta Ayala | 2312080871 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-08 10:00 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080936 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-08 10:05 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080898 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-08 10:05 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080946 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-08 10:10 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080951 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-08 10:12 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080891 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-08 10:15 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080965 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-08 10:18 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080894 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-08 10:19 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080967 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-08 10:24 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080956 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-08 10:35 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312180128 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-08 10:36 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312180133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-08 10:38 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080964 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-08 10:43 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080991 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-08 10:44 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080949 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-08 10:48 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080990 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-08 10:50 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080957 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-08 10:51 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080955 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-08 11:03 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080980 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-08 11:30 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080877 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.329 | Pass |  | n/a
2024-01-08 11:35 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312180110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-08 11:35 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080945 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-08 11:43 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312180082 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-08 11:44 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080962 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-08 11:50 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080961 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-08 11:50 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080921 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-08 11:55 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080959 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-08 11:59 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080873 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-08 12:00 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080925 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-08 12:03 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080917 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-08 12:09 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312180118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-08 12:13 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312180107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-08 12:16 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312180076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-08 12:20 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080920 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-08 12:21 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080996 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.329 | Pass |  | n/a
2024-01-08 12:24 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080928 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-08 12:27 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312180061 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-08 12:33 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080997 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-08 12:36 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080901 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-08 12:39 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080982 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-08 12:47 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080865 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-08 12:48 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080881 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-08 12:50 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080993 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-08 12:54 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312180090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-08 12:54 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080872 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-08 12:57 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080960 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-08 12:59 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312180077 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-08 13:01 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312180079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-08 13:06 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080882 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-08 14:05 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080914 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-08 14:06 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080915 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-08 14:10 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080930 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-08 14:11 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080922 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-08 14:23 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080869 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-08 14:25 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312180108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-08 14:33 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-08 14:37 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080927 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-08 14:38 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080953 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-08 14:42 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312180087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-08 14:43 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080954 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-08 14:47 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312180112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-08 14:47 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-08 14:54 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.358 | Pass |  | n/a
2024-01-08 14:55 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080931 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-08 15:05 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312180111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-08 15:07 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080926 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-08 15:14 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080507 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-08 15:15 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080916 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-08 15:17 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-08 15:28 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312180083 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-08 15:29 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312180085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-08 15:30 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312180109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-09 09:27 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080776 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-09 09:31 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080724 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-09 09:39 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080775 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-09 09:42 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080723 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-09 09:43 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-09 09:48 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080721 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-09 09:49 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-09 09:58 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080756 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-09 10:01 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080777 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-09 10:06 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.357 | Pass |  | n/a
2024-01-09 10:10 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-09 10:16 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080767 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-09 10:17 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080751 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-09 10:25 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080768 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-09 10:32 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080774 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-09 10:34 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080762 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-09 10:35 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080718 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-09 10:36 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080759 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-09 10:47 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080757 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-09 10:52 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080727 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-09 11:35 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080998 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-09 11:42 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312180099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-09 11:44 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080896 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-09 11:50 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080798 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-09 11:55 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080937 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-09 11:56 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080978 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-09 12:42 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080944 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-09 12:46 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312180096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-09 12:49 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080778 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-09 12:52 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312180093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-09 12:57 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080975 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-09 12:58 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080968 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-09 13:03 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080995 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-09 13:07 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080970 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-09 13:12 | stmstation4@avive.life | W1175 | Agree | Liza Ordillo | 2312080969 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-09 13:12 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080893 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-09 13:15 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080889 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-09 13:17 | stmstation3@avive.life | W1175 | Agree | Liza Ordillo | 2312080895 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-09 14:07 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080637 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-09 14:08 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080734 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-09 14:13 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080808 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-09 14:15 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080923 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-09 14:22 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080812 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-09 14:23 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080824 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-09 14:25 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080748 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.358 | Pass |  | n/a
2024-01-09 14:30 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312180120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-09 14:34 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312180121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-09 14:36 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080934 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-09 14:40 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312180122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-09 14:40 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080771 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-09 14:45 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312180131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 1.03 | Fail | failed limit to high
2024-01-09 14:48 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080836 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-09 14:49 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080932 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-09 14:58 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080933 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-09 15:14 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080698 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-09 15:18 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080947 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-09 15:18 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080929 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-09 15:25 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080754 | Pass | Fail | [10 empty] | screen stuck.
2024-01-09 15:27 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080820 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-09 15:34 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080696 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-09 15:36 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080730 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-09 15:36 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080669 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-09 15:52 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080694 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-09 15:52 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080825 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-09 15:59 | stmstation3@avive.life | W1175 | Agree | Leonora Cruz | 2312080697 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-09 16:00 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-09 16:01 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-09 16:03 | stmstation4@avive.life | W1175 | Agree | Leonora Cruz | 2312080750 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-09 16:09 | stmstation1@avive.life | W1175 | Agree | Leonora Cruz | 2312080692 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-09 16:10 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080769 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-10 07:03 | stmstation1@avive.life | W1175 | Agree | Precious Ungco | 2312080761 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-10 07:05 | stmstation3@avive.life | W1175 | Agree | Precious Ungco | 2312080851 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-10 07:09 | stmstation4@avive.life | W1175 | Agree | Precious Ungco | 2312080636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-10 07:12 | stmstation1@avive.life | W1175 | Agree | Precious Ungco | 2312080711 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.305 | Pass |  | n/a
2024-01-10 07:14 | stmstation2@avive.life | W1175 | Agree | Precious Ungco | 2312080244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-10 07:17 | stmstation4@avive.life | W1175 | Agree | Precious Ungco | 2312080773 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-10 07:18 | stmstation3@avive.life | W1175 | Agree | Precious Ungco | 2312080707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-10 07:22 | stmstation1@avive.life | W1175 | Agree | Precious Ungco | 2312080758 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-10 07:23 | stmstation2@avive.life | W1175 | Agree | Precious Ungco | 2312080035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-10 07:28 | stmstation4@avive.life | W1175 | Agree | Precious Ungco | 2312080728 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-10 07:30 | stmstation2@avive.life | W1175 | Agree | Precious Ungco | 2312080790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-10 07:31 | stmstation3@avive.life | W1175 | Agree | Precious Ungco | 2312080124 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-10 07:33 | stmstation1@avive.life | W1175 | Agree | Precious Ungco | 2312080726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-10 07:39 | stmstation3@avive.life | W1175 | Agree | Precious Ungco | 2312080831 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-10 07:41 | stmstation2@avive.life | W1175 | Agree | Precious Ungco | 2312080814 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.402 | Pass |  | n/a
2024-01-10 07:43 | stmstation4@avive.life | W1175 | Agree | Precious Ungco | 2312080226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.357 | Pass |  | n/a
2024-01-10 07:44 | stmstation1@avive.life | W1175 | Agree | Precious Ungco | 2312180114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-10 07:51 | stmstation3@avive.life | W1175 | Agree | Precious Ungco | 2312080320 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-10 07:52 | stmstation4@avive.life | W1175 | Agree | Precious Ungco | 2312080816 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.361 | Pass |  | n/a
2024-01-10 07:53 | stmstation1@avive.life | W1175 | Agree | Precious Ungco | 2312080819 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-10 07:54 | stmstation2@avive.life | W1175 | Agree | Precious Ungco | 2312080826 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-10 08:22 | stmstation3@avive.life | W1175 | Agree | Precious Ungco | 2312080693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-10 08:45 | stmstation4@avive.life | W1175 | Agree | Precious Ungco | 2312180132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.369 | Pass |  | n/a
2024-01-10 09:40 | stmstation4@avive.life | W1175 | Agree | Precious Ungco | 2312080690 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.34 | Pass |  | ireri moreno
2024-01-10 09:41 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080744 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-10 09:42 | stmstation1@avive.life | W1175 | Agree | ireri moreno | 2312080823 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-10 09:42 | stmstation3@avive.life | W1175 | Agree | Ireri Moreno | 2312080699 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.365 | Pass |  | n/a
2024-01-10 09:52 | stmstation4@avive.life | W1175 | Agree | Ireri Moreno | 2312080838 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-10 09:53 | stmstation3@avive.life | W1175 | Agree | Ireri Moreno | 2312080830 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-10 09:53 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080827 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-10 10:00 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080818 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-10 10:02 | stmstation4@avive.life | W1175 | Agree | Ireri Moreno | 2312080883 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-10 10:05 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312081000 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-10 10:16 | stmstation3@avive.life | W1175 | Agree | Ireri Moreno | 2312180123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Fail |  |  | no sim
2024-01-10 10:22 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312180103 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-10 10:24 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312180119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-10 10:24 | stmstation3@avive.life | W1175 | Agree | ireri Moreno | 2312080638 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-10 10:32 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-10 10:32 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080719 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-10 10:33 | stmstation3@avive.life | W1175 | Agree | ireri Moreno | 2312080722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-10 11:04 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080862 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-10 11:04 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080815 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-10 11:06 | stmstation3@avive.life | W1175 | Agree | Ireri Moreno | 2312080736 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-10 12:13 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080735 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-10 12:14 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080999 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-10 12:17 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080668 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-10 12:18 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080804 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-10 12:22 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080828 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-10 12:23 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312180080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-10 12:26 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-10 12:26 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080691 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-10 12:31 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080763 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-10 12:32 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080732 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-10 12:37 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080749 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-10 12:39 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080380 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-10 12:40 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080941 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-10 12:45 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080741 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-10 12:47 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-10 12:50 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-10 12:54 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080737 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-10 12:55 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080738 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-10 12:57 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080742 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-10 12:58 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080716 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-10 13:41 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080714 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-10 13:42 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080942 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-10 13:43 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080733 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-10 13:44 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312180106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-10 13:49 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080821 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-10 13:51 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-10 13:54 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080666 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-10 13:59 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312180094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-10 14:01 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312180092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-10 14:03 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080864 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-10 14:04 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080745 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-10 14:08 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080720 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-10 14:10 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-10 14:12 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080667 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-10 14:14 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080878 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-10 14:16 | stmstation1@avive.life | W1175 | Agree | Valeria Manzo | 2312080766 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-10 14:18 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080670 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-10 14:21 | stmstation2@avive.life | W1175 | Agree | Valeria Manzo | 2312080867 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-11 07:55 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080832 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-11 08:06 | stmstation2@avive.life | W1175 | Agree | Leonora Cruz | 2312080866 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-11 08:24 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080840 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-11 08:29 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080743 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-11 08:31 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.393 | Pass |  | n/a
2024-01-11 08:36 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080747 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-11 08:37 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080695 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-11 08:45 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080841 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-11 08:49 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080875 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-11 08:49 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080854 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-11 08:52 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080852 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-11 08:55 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080680 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-11 08:55 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080855 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-11 08:59 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080943 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-11 09:02 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080842 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-11 09:05 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080839 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-11 09:07 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080938 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-11 09:08 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080689 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-11 09:09 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080679 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-11 09:12 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080856 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-11 09:13 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080752 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-11 09:16 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080755 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-11 09:23 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080810 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.345 | Pass |  | n /a
2024-01-11 09:29 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080377 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-11 09:40 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-11 09:40 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-11 09:41 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080809 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-11 09:42 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.356 | Pass |  | n/a
2024-01-11 09:50 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080764 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-11 09:50 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080772 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-11 09:55 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080876 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-11 09:55 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080835 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-11 10:46 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080746 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-11 10:53 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080725 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-11 10:53 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080822 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-11 10:56 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080837 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.338 | Pass |  | N/A
2024-01-11 11:04 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.356 | Pass |  | n/a
2024-01-11 11:06 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-11 11:07 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080817 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-11 11:09 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-11 11:12 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080803 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-11 11:13 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080704 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-11 11:16 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-11 11:17 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312180088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-11 11:17 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-11 11:20 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080880 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-11 11:23 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080863 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-11 11:25 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080385 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-11 11:33 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080813 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-11 11:36 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080833 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-11 11:46 | stmstation2@avive.life | W1175 | Agree | Ireri Moreno | 2312080760 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-11 12:02 | stmstation1@avive.life | W1175 | Agree | Ireri Moreno | 2312080829 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | N/A
2024-01-11 12:17 | stmstation2@avive.life | W1175 | Agree | Liza Ordillo | 2312080783 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.342 | Pass |  | Reworked on Ribbon Cable
2024-01-11 12:24 | stmstation4@avive.life | W1175 | Agree | Valeria Manzo | 2312080843 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Fail |  |  | No SIM sign
2024-01-11 12:31 | stmstation3@avive.life | W1175 | Agree | Valeria Manzo | 2312080979 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.344 | Pass |  | N/A
2024-01-11 12:34 | stmstation1@avive.life | W1175 | Agree | Liza Ordillo | 2312080754 | Pass | Pass | Pass | Fail | [8 empty] | Can't connect to GPS, Cell and WIFI
2024-01-12 07:26 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312080848 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-12 07:29 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312080802 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-12 07:31 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312080681 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-12 07:39 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312080682 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-12 07:39 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312080688 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-12 07:47 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312080674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-12 07:47 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312080849 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-12 07:54 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312080673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-12 07:55 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312080683 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-12 08:01 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312080671 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-12 08:02 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312080811 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-12 08:27 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312080801 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-12 08:30 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-12 08:31 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 08:34 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312080687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-12 09:09 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312080788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-12 09:10 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180174 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-12 09:13 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312080678 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-12 09:15 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312080794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-12 09:39 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-12 09:48 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-12 09:52 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312080792 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 09:54 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-12 09:57 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180316 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-12 09:59 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312080686 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-12 10:01 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-12 10:07 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312080791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 10:07 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312080845 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-12 10:08 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312080684 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-12 10:17 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312080857 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-12 10:23 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312080677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-12 10:26 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312080675 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-12 10:29 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312080799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-12 10:37 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312080859 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-12 10:39 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312080846 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-12 10:43 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312080847 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-12 10:47 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312080685 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-12 10:54 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 10:57 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312080787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-12 11:00 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-12 11:02 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180310 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-12 11:04 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 11:09 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312080850 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-12 11:10 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180191 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-12 11:14 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312080797 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-12 11:17 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 11:18 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312080860 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-12 11:19 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-12 11:24 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180581 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-12 11:25 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 11:27 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 11:33 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 11:34 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180391 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-12 11:36 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180536 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-12 11:41 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-12 11:43 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180349 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-12 11:45 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-12 11:49 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180368 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 11:50 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180383 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-12 11:56 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180392 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-12 11:58 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180545 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-12 11:59 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180362 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-12 12:00 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180311 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-12 12:06 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180352 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-12 12:10 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-12 12:12 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-12 12:18 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-12 12:24 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312080861 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 12:27 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180023 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-12 12:29 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-12 12:31 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180547 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-12 12:36 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312080807 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.356 | Pass |  | n/a
2024-01-12 12:37 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.399 | Pass |  | n/a
2024-01-12 12:38 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 12:43 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-12 13:08 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-12 13:15 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312080858 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 13:16 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 13:28 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180313 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 13:30 | stmstation1@avive.life | W1180 | Agree | Christina Cruz | 2312180546 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-12 13:35 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312080800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 13:40 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180482 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 13:41 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180232 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-12 13:43 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180380 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-12 13:45 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180243 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-12 13:50 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-12 13:52 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180288 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-12 14:08 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180235 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-12 14:09 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180245 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 14:11 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180312 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-12 14:13 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180242 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-12 14:18 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-12 14:23 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180559 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 14:25 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312080806 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-12 14:27 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180519 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-12 14:28 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180167 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-12 14:32 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312080844 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-12 14:58 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180160 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-12 15:00 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180139 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-12 15:01 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 15:02 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-12 15:34 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312080805 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 15:37 | stmstation1@avive.life | W1180 | Agree | Christina Cruz | 2312180125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 15:41 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-12 15:50 | stmstation1@avive.life | W1180 | Agree | Christina Cruz | 2312180074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-12 15:54 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 15:57 | stmstation1@avive.life | W1180 | Agree | Christina Cruz | 2312180226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 16:03 | stmstation1@avive.life | W1180 | Agree | Christina Cruz | 2312180052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-12 16:05 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 16:10 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-12 16:11 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 16:13 | stmstation1@avive.life | W1180 | Agree | Christina Cruz | 2312180206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 16:18 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180386 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-12 16:20 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180283 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-12 16:40 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-12 16:42 | stmstation1@avive.life | W1180 | Agree | Christina Cruz | 2312180046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 16:44 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-12 16:51 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180510 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-12 16:57 | stmstation1@avive.life | W1180 | Agree | Christina Cruz | 2312180025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 08:12 | stmstation3@avive.life | W1180 | Agree | Violeta Ayala | 2312180024 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-16 08:13 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180063 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-16 08:20 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 08:22 | stmstation2@avive.life | W1180 | Agree | Violeta Ayala | 2312180017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 08:27 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-16 08:28 | stmstation3@avive.life | W1180 | Agree | Violeta Ayala | 2312180066 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 08:37 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-16 08:42 | stmstation2@avive.life | W1180 | Agree | Violeta Ayala | 2312180204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-16 08:50 | stmstation3@avive.life | W1180 | Agree | Violeta Ayala | 2312180020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 08:54 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 09:00 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180514 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-16 09:00 | stmstation2@avive.life | W1180 | Agree | Violeta Ayala | 2312180521 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-16 09:10 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180511 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-16 09:11 | stmstation2@avive.life | W1180 | Agree | Violeta Ayala | 2312080793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-16 09:13 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-16 09:31 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180504 | Pass | Fail | [10 empty] | Voltage is high; does not turn on.
2024-01-16 09:37 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180520 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-16 09:39 | stmstation2@avive.life | W1180 | Agree | Violeta Ayala | 2312180064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-16 09:47 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 09:47 | stmstation3@avive.life | W1180 | Agree | Violeta Ayala | 2312080853 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 09:52 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180065 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 10:00 | stmstation2@avive.life | W1180 | Agree | Violeta Ayala | 2312180284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-16 10:01 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-16 10:05 | stmstation3@avive.life | W1180 | Agree | Violeta Ayala | 2312180246 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 10:08 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-16 10:19 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 10:23 | stmstation2@avive.life | W1180 | Agree | Violeta Ayala | 2312180223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-16 10:25 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 10:32 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312180059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 10:33 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312180036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-16 10:36 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 10:37 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180224 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-16 10:54 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312180544 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 10:55 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 10:55 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180238 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.356 | Pass |  | n/a
2024-01-16 11:03 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312180228 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-16 11:05 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 11:07 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312180382 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-16 11:15 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-16 11:19 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312180528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-16 11:19 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312180389 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-16 11:27 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180342 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-16 11:27 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180276 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 11:28 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312180115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 11:42 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312180071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 11:43 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312180145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.377 | Pass |  | n/a
2024-01-16 11:49 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-16 11:56 | stmstation3@avive.life | W1180 | Agree | Liza Ordillo | 2312180022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 11:56 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312180508 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-16 12:01 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 12:05 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180241 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-16 12:21 | stmstation1@avive.life | W1180 | Agree | Leonora Cruz | 2312180317 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-16 12:21 | stmstation2@avive.life | W1180 | Agree | Leonora Cruz | 2312180162 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 12:28 | stmstation1@avive.life | W1180 | Agree | Leonora Cruz | 2312180045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 12:33 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180484 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-16 12:34 | stmstation2@avive.life | W1180 | Agree | Leonora Cruz | 2312180197 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 12:36 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 12:45 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-16 12:48 | stmstation1@avive.life | W1180 | Agree | Leonora Cruz | 2312180008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-16 12:48 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180294 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-16 12:52 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 12:56 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180329 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-16 12:56 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180293 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-16 12:56 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 13:02 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180325 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 13:03 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180200 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-16 13:10 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 13:12 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-16 13:14 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180279 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-16 13:14 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.401 | Pass |  | n/a
2024-01-16 13:18 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180366 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-16 13:21 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180230 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 13:26 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 13:27 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-16 13:28 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-16 13:28 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180385 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-16 13:37 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180170 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-16 13:38 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-16 13:43 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180269 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-16 13:44 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.359 | Pass |  | n/a
2024-01-16 13:45 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 13:53 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180138 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-16 13:54 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180332 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-16 13:54 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 14:00 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180075 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 14:03 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180159 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 14:04 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-16 14:06 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180319 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 14:06 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-16 14:09 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180365 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 14:12 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-16 14:15 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180267 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 14:17 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180507 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-16 14:18 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-16 14:20 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-16 14:26 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180390 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 14:26 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180530 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 14:28 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 14:30 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-16 14:41 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180381 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-16 14:44 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180526 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-16 14:49 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180124 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 14:50 | stmstation3@avive.life | W1180 | Agree | Violeta Ayala | 2312180531 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-16 14:59 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 15:02 | stmstation3@avive.life | W1180 | Agree | Violeta Ayala | 2312180009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-16 15:04 | stmstation2@avive.life | W1180 | Agree | Violeta Ayala | 2312180351 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 15:27 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180350 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-16 15:30 | stmstation3@avive.life | W1180 | Agree | Violeta Ayala | 2312180127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 15:36 | stmstation2@avive.life | W1180 | Agree | Violeta Ayala | 2312180011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-16 15:37 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 15:40 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 15:41 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180384 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.327 | Pass |  | n/a
2024-01-16 15:43 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-16 15:49 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180073 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.321 | Pass |  | n/a
2024-01-16 15:50 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180287 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-16 16:00 | stmstation1@avive.life | W1180 | Agree | Leonora Cruz | 2312180182 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 16:02 | stmstation3@avive.life | W1180 | Agree | Leonora Cruz | 2312180058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 16:02 | stmstation4@avive.life | W1180 | Agree | Leonora Cruz | 2312180184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-16 16:05 | stmstation2@avive.life | W1180 | Agree | Leonora Cruz | 2312180177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-16 16:08 | stmstation1@avive.life | W1180 | Agree | Leonora Cruz | 2312180035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 16:12 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180248 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 16:15 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-16 16:18 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 16:20 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-16 16:23 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 16:25 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-16 16:29 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-16 16:30 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180367 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-16 16:34 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180320 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 16:38 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 16:42 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180070 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-16 16:44 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 16:46 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180282 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 16:48 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-16 16:52 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180152 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 16:54 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-16 16:56 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180201 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 16:59 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180347 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 17:03 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 17:05 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-16 17:07 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180168 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-16 17:13 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 17:16 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180278 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 17:17 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180247 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 17:24 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180158 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 17:25 | stmstation2@avive.life | W1180 | Agree | Leonora Cruz | 2312180018 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-16 17:26 | stmstation1@avive.life | W1180 | Agree | Leonora Cruz | 2312180334 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-16 17:27 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180007 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 17:38 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180277 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 17:38 | stmstation1@avive.life | W1180 | Agree | Leonora Cruz | 2312180004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-16 17:40 | stmstation2@avive.life | W1180 | Agree | Leonora Cruz | 2312180281 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-16 17:44 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180330 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 17:48 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 17:48 | stmstation1@avive.life | W1180 | Agree | Leonora Cruz | 2312180561 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-16 17:54 | stmstation2@avive.life | W1180 | Agree | Leonora Cruz | 2312180318 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-16 17:57 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180331 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-16 18:00 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180165 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-16 18:04 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 18:06 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180328 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 18:09 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180081 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 18:12 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 18:15 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180549 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-16 18:17 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180485 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 18:19 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180315 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-16 18:21 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180041 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-16 18:21 | stmstation1@avive.life | W1180 | Agree | Christina Cruz | 2312180031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 05:10 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 05:12 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180233 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-17 05:15 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180236 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-17 05:16 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180343 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-17 05:17 | stmstation1@avive.life | W1180 | Agree | Valeria Manzo | 2312180229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-17 05:25 | stmstation2@avive.life | W1180 | Agree | Valeria Manzo | 2312180534 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-17 05:26 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180345 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-17 05:27 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180172 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-17 05:42 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180509 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-17 05:43 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-17 05:47 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-17 05:54 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.354 | Pass |  | n/a
2024-01-17 05:57 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180019 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-17 06:36 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-17 06:39 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180369 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-17 06:44 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-17 06:54 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 06:57 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-17 07:15 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-17 07:16 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180377 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 07:16 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180275 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-17 07:25 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180231 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-17 07:34 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180333 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-17 07:36 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-17 07:39 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180296 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-17 07:45 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180144 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-17 07:53 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180324 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.26 | Pass |  | n/a
2024-01-17 07:54 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-17 07:55 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-17 08:02 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180379 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-17 08:04 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180234 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 08:09 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180266 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 08:10 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180341 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.373 | Pass |  | n/a
2024-01-17 08:11 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-17 08:13 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-17 08:24 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180396 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-17 08:28 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180298 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-17 08:32 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180376 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-17 08:36 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180370 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-17 08:39 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-17 08:45 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180322 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.359 | Pass |  | n/a
2024-01-17 08:50 | stmstation4@avive.life | W1180 | Agree | Valeria Manzo | 2312180505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-17 08:52 | stmstation3@avive.life | W1180 | Agree | Valeria Manzo | 2312180149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 08:56 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-17 08:59 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180346 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-17 09:09 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180292 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-17 09:10 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180227 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 09:14 | stmstation1@avive.life | W1180 | Agree | Liza Ordillo | 2312180006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-17 09:28 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180291 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-17 09:32 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180344 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-17 09:37 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 09:39 | stmstation2@avive.life | W1180 | Agree | Liza Ordillo | 2312180166 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-17 09:42 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180240 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-17 09:49 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180297 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-17 09:53 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-17 09:57 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180321 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-17 09:59 | stmstation2@avive.life | W1180 | Agree | Christina Cruz | 2312180323 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 10:02 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180299 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 10:04 | stmstation4@avive.life | W1180 | Agree | Christina Cruz | 2312180055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-17 10:27 | stmstation4@avive.life | W1180 | Agree | Liza Ordillo | 2312180055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-17 10:41 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180516 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-17 10:53 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180300 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-17 10:55 | stmstation3@avive.life | W1180 | Agree | Christina Cruz | 2312180239 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-17 10:58 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-17 11:04 | stmstation4@avive.life | W1180 | Agree | Violeta Ayala | 2312180301 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-17 11:20 | stmstation1@avive.life | W1180 | Agree | Violeta Ayala | 2312180527 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Fail |  |  | No SIM sign
2024-01-19 16:53 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180417 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-19 17:12 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180988 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.341 | Pass |  | N/A
2024-01-19 17:12 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180667 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-19 17:16 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180829 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.343 | Pass |  | N/A
2024-01-19 17:17 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180936 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-22 12:04 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180867 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-22 12:05 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180721 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-22 12:13 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180815 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-22 12:16 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180866 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-22 12:16 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180747 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-22 12:18 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180678 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-22 12:31 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180748 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-22 12:31 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180728 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-22 12:31 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180730 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-22 12:31 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180777 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-22 12:42 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180862 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-22 12:42 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180739 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-22 12:45 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180817 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-22 12:46 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180727 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-22 12:54 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180733 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.354 | Pass |  | n/a
2024-01-22 12:56 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180869 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-22 12:58 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180773 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-22 13:00 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180785 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-22 13:42 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180741 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-22 13:44 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180784 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-22 13:46 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180871 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-22 13:46 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180865 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-22 13:55 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180811 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-22 13:56 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-22 14:01 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180894 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-22 14:10 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180848 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-22 14:11 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-22 14:12 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180835 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-22 14:13 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180751 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-22 15:11 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180669 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-22 15:15 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180915 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-22 15:22 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-22 15:23 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180932 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-22 15:33 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180675 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-22 15:40 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-22 15:46 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180664 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-22 15:46 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-22 15:51 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.361 | Pass |  | n/a
2024-01-22 15:58 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180699 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-22 15:59 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180922 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.354 | Pass |  | n/a
2024-01-22 16:04 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-22 16:11 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180765 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-22 16:12 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180931 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-22 16:14 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180666 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-22 16:25 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180920 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-22 16:27 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180937 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-22 16:31 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180774 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-22 16:32 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180757 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-22 16:45 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180855 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-22 16:48 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180756 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-22 16:50 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180853 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-22 16:54 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180761 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-22 16:59 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180929 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-22 17:00 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180668 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-22 17:28 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180926 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-22 17:33 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-22 17:45 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180605 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-22 17:46 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-22 17:54 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180737 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-22 17:58 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180494 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-22 18:01 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-22 18:02 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-22 18:03 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-22 18:07 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180414 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 09:04 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180798 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 09:07 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-23 09:17 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180776 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 09:25 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-23 09:29 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180409 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 09:34 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180567 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 10:22 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180846 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 10:36 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180764 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 10:37 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180688 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-23 10:39 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180671 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 10:53 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180467 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 10:55 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-23 11:02 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180500 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-23 11:06 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-23 11:12 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180468 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-23 11:18 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180464 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-23 11:23 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180685 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 11:32 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180704 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-23 11:38 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-23 11:38 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180495 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.356 | Pass |  | n/a
2024-01-23 11:40 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-23 11:51 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 11:51 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180569 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-23 11:58 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-23 12:02 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180766 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-23 12:02 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180690 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 12:03 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180694 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-23 12:11 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-23 12:16 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180778 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-23 12:16 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180767 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-23 12:16 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180609 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-23 12:24 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180696 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.341 | Pass |  | N/A
2024-01-23 12:27 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180587 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.343 | Pass |  | N/A
2024-01-23 12:28 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.334 | Pass |  | N/a
2024-01-23 12:34 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180606 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.338 | Pass |  | N/A
2024-01-23 12:36 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180692 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-23 12:36 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 429 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-23 12:39 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-23 12:42 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180533 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-23 12:45 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.346 | Pass |  | N/A
2024-01-23 12:47 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-23 12:47 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180465 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-23 12:52 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180459 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-23 12:56 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-23 12:57 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180610 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-23 12:59 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180758 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-23 13:04 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180589 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-23 13:06 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180592 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-23 13:48 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180752 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-23 13:49 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180783 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-23 13:52 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180902 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.333 | Pass |  | n//a
2024-01-23 13:57 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180856 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-23 13:58 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180854 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-23 14:01 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180683 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-23 14:13 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-23 14:20 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180608 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Fail |  |  | No SIM Sign
2024-01-23 14:27 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Fail |  |  | No SIM sign
2024-01-23 14:27 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180679 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.346 | Pass |  | N/a
2024-01-23 14:30 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180760 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-23 14:32 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 14:42 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180897 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-23 14:43 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180775 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-23 14:44 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180691 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-23 14:46 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 14:52 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180440 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-23 15:05 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180935 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-23 15:19 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180686 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 15:22 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180851 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-23 15:24 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180314 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 15:27 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180682 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 15:30 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180697 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-23 15:32 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 15:37 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180670 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.344 | Pass |  | n\a
2024-01-23 15:37 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180923 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-23 15:37 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180695 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-23 15:41 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180680 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-23 15:43 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-23 15:44 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180672 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-23 15:46 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180681 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-23 15:50 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180864 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 15:51 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180916 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-23 15:54 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180917 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-23 15:56 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.443 | Pass |  | n/a
2024-01-23 16:01 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180868 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-23 16:12 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180816 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-23 16:18 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180847 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-23 16:20 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180925 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-23 16:24 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180772 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-23 16:26 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180921 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-23 16:26 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180872 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-23 16:27 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180714 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-23 16:31 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180744 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-23 16:39 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180585 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-23 16:41 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.346 | Pass |  | N/A
2024-01-23 16:43 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180901 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-23 16:49 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180738 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 16:52 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-23 16:55 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180857 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-23 16:57 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180927 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-23 17:00 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180460 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-23 17:11 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180779 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 17:13 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180736 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-23 17:40 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180568 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-23 17:45 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-23 17:57 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180859 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-23 17:59 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 06:37 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180780 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-24 06:41 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180924 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-24 06:43 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180918 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-24 08:58 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 09:05 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.35 | Pass |  | n/a
2024-01-24 09:08 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 09:22 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180732 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 09:22 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180792 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-24 09:24 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 09:47 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180735 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-24 09:51 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180607 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-24 09:56 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180762 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 09:57 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-24 10:07 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180755 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-24 10:08 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180754 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-24 10:13 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180458 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-24 10:34 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180768 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-24 10:42 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180749 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-24 10:44 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-24 10:45 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180938 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 10:52 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180919 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 10:53 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180769 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-24 10:58 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180759 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-24 10:58 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180466 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-24 11:07 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180771 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-24 11:09 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-24 11:24 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180930 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-24 11:25 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180710 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 11:25 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180795 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-24 11:51 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180718 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-24 11:52 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180812 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-24 11:53 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180715 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.443 | Pass |  | n/a
2024-01-24 11:54 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180830 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.371 | Pass |  | n/a
2024-01-24 12:05 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180819 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-24 12:06 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180724 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.379 | Pass |  | n/a
2024-01-24 12:07 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180895 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-24 12:09 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180834 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-24 12:14 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180845 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-24 12:17 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180838 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-24 12:24 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180745 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-24 12:25 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-24 12:30 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180898 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.359 | Pass |  | n/a
2024-01-24 12:31 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180818 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.388 | Pass |  | n/a
2024-01-24 12:32 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180406 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-24 12:37 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180746 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-24 12:38 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-24 12:40 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180412 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-24 12:41 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180413 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-24 12:43 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180725 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-24 12:50 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180720 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-24 12:51 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180716 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.394 | Pass |  | n/a
2024-01-24 12:52 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-24 12:53 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-24 12:59 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180405 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-24 13:55 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-24 13:57 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-24 13:58 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-24 14:00 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-24 14:05 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180642 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-24 14:05 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180438 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-24 14:11 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180463 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-24 14:11 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180734 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-24 14:16 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180713 | Pass | Pass | Pass | Pass | Fail | [7 empty] | wifi on & off
2024-01-24 14:21 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180637 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-24 14:21 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180434 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.359 | Pass |  | n/a
2024-01-24 14:25 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180719 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-24 14:26 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180837 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-24 14:30 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 14:31 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180644 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-24 14:33 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180813 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-24 14:33 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180469 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-24 14:38 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180646 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.358 | Pass |  | n/a
2024-01-24 14:41 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180648 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-24 14:46 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180820 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-24 14:48 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180456 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-24 14:57 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180424 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-24 14:58 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180471 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-24 15:02 | stmstation3@avive.life | W1184 | Agree | Precious Ungco | 2312180426 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-24 15:12 | stmstation3@avive.life | W1184 | Agree | Ananya Jain | 2312180604 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.342 | Pass |  | N/A
2024-01-24 15:13 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-24 15:17 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180432 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.354 | Pass |  | n/a
2024-01-24 15:19 | stmstation4@avive.life | W1184 | Agree | Precious Ungco | 2312180480 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Fail |  |  | No SIM Sign
2024-01-24 15:22 | stmstation2@avive.life | W1184 | Agree | Ananya Jain | 2312180450 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-24 15:48 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180474 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-24 15:51 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180407 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-24 15:53 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180435 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-24 15:54 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180423 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-24 16:04 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180473 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 16:05 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180408 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.327 | Pass |  | n/a
2024-01-24 16:16 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180449 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-24 16:18 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180470 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-24 16:21 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180457 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-24 16:23 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180411 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-24 16:26 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180476 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-24 16:39 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180723 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-24 16:41 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180453 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.334 | Pass |  | N/A
2024-01-24 16:42 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180439 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-24 16:47 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180618 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-24 16:49 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180619 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-24 17:03 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180339 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-25 09:42 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180419 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-25 09:50 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-25 09:55 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180425 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-25 09:58 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Fail |  |  | SIM left for 30mins not working
2024-01-25 10:05 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180421 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 10:07 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180614 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.333 | Pass |  | N/A
2024-01-25 10:15 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180415 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-25 10:17 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180617 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-25 10:24 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180452 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-25 10:30 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180454 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-25 10:36 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180498 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 10:36 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180499 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.341 | Pass |  | N/A
2024-01-25 10:43 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180623 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-25 10:45 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180622 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-25 10:46 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180831 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 10:55 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-25 10:55 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180442 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 10:57 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180431 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-25 11:01 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180448 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 11:08 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.358 | Pass |  | n/a
2024-01-25 11:10 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180340 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-25 11:12 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180497 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-25 11:13 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180436 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-25 11:20 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180861 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-25 11:21 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180850 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-25 11:22 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180620 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 11:25 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180893 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 11:32 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180427 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-25 11:34 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180338 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-25 11:41 | stmstation2@avive.life | W1184 | Agree | Liza Ordillo | 2312180742 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-25 11:42 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180455 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-25 11:50 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180337 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 11:51 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180336 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-25 11:52 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180420 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.354 | Pass |  | n/a
2024-01-25 11:58 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180433 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-25 12:02 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180624 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-25 12:03 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180896 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-25 12:05 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 12:08 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-25 12:08 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-25 12:13 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180461 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-25 12:15 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180566 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-25 12:16 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180462 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-25 12:18 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180840 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-25 12:20 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180563 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-25 12:24 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180551 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-25 12:26 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180576 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-25 12:36 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180416 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-25 12:37 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180575 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-25 12:38 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180538 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-25 12:38 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180400 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-25 12:45 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180573 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-25 12:46 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180603 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-25 12:46 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-25 12:49 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180590 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-25 12:54 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180580 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-25 12:55 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180429 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-25 12:58 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 12:58 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180583 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-25 13:06 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-25 13:20 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-25 13:21 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180631 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 13:21 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-25 13:23 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180601 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.357 | Pass |  | n/a
2024-01-25 13:28 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180634 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-25 13:30 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180578 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-25 13:32 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180472 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-25 13:35 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180307 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.286 | Pass |  | n/a
2024-01-25 13:38 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180446 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 13:39 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180404 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 13:44 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180308 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.372 | Pass |  | n/a
2024-01-25 13:45 | stmstation3@avive.life | W1184 | Agree | Liza Ordillo | 2312180428 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-25 13:45 | stmstation4@avive.life | W1184 | Agree | Liza Ordillo | 2312180763 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-25 13:47 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-25 13:51 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180306 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-25 13:54 | stmstation3@avive.life | W1184 | Agree | Ireri Moreno | 2312180797 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-25 13:57 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180403 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.323 | Pass |  | n/a
2024-01-25 13:58 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180602 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 14:03 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180302 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-25 14:03 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 14:08 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180863 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 14:08 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-25 14:14 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-25 14:14 | stmstation4@avive.life | W1184 | Agree | ireri Moreno | 2312180555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 14:16 | stmstation3@avive.life | W1184 | Agree | Ireri Moreno | 2312180613 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 14:18 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180444 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 14:19 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180839 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-25 14:32 | stmstation4@avive.life | W1184 | Agree | Ireri Moreno | 2312180632 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-25 14:34 | stmstation3@avive.life | W1184 | Agree | Ireri Moreno | 2312180445 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 14:43 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180552 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-25 14:44 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180556 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-25 14:46 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180577 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-25 14:46 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-25 14:49 | stmstation3@avive.life | W1184 | Agree | Ireri Moreno | 2312180570 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 14:55 | stmstation4@avive.life | W1184 | Agree | Ireri Moreno | 2312180571 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-25 15:04 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180630 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-25 15:07 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180550 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-25 15:30 | stmstation2@avive.life | W1184 | Agree | Leonora Cruz | 2312180443 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-25 15:33 | stmstation4@avive.life | W1184 | Agree | Leonora Cruz | 2312180410 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 15:33 | stmstation3@avive.life | W1184 | Agree | Leonora Cruz | 2312180572 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-25 15:35 | stmstation1@avive.life | W1184 | Agree | Leonora Cruz | 2312180478 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-25 15:42 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180616 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-25 15:46 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180593 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-25 15:51 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180615 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-25 16:03 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180335 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-25 16:08 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180303 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-25 16:09 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180373 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.334 | Pass |  | N/a
2024-01-25 16:09 | stmstation1@avive.life | W1184 | Agree | Christina Cruz | 2312180378 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.339 | Pass |  | N/A
2024-01-25 16:11 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180309 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-25 16:17 | stmstation4@avive.life | W1184 | Agree | Christina Cruz | 2312180305 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.338 | Pass |  | N/A
2024-01-25 17:35 | stmstation3@avive.life | W1184 | Agree | Christina Cruz | 2312180304 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.336 | Pass |  | N/A
2024-01-26 06:06 | stmstation3@avive.life | W1184 | Agree | Ireri Moreno | 2312180451 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-26 06:08 | stmstation4@avive.life | W1184 | Agree | Ireri Moreno | 2312180554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-26 06:11 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180743 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-26 06:13 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180621 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-26 06:15 | stmstation3@avive.life | W1184 | Agree | Ireri Moreno | 2312180422 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-26 06:19 | stmstation4@avive.life | W1184 | Agree | Ireri Moreno | 2312180836 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-26 06:19 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180477 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-26 06:20 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180626 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-26 06:27 | stmstation4@avive.life | W1184 | Agree | Ireri Moreno | 2312180441 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-26 06:28 | stmstation2@avive.life | W1184 | Agree | Precious Ungco | 2312180447 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-26 06:28 | stmstation1@avive.life | W1184 | Agree | Precious Ungco | 2312180750 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.334 | Pass |  | n/a
2024-01-26 06:33 | stmstation4@avive.life | W1184 | Agree | Ireri Moreno | 2312180833 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-26 07:05 | stmstation2@avive.life | W1184 | Agree | Valeria Manzo | 2312180430 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | no cell 30 min and no sim
2024-01-26 08:57 | stmstation2@avive.life | W1184 | Agree | Christina Cruz | 2312180562 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-26 09:12 | stmstation1@avive.life | W1184 | Agree | Liza Ordillo | 2312180713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.346 | Pass |  | Reworked
2024-01-29 08:19 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2312180640 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-29 08:20 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2312180874 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-29 08:23 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180947 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-29 08:24 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050333 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-29 08:31 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2312180946 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-29 08:32 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2312180943 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 08:34 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180823 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-29 08:34 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180662 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-29 08:46 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2312180878 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-29 08:47 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180599 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-29 08:48 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2312180873 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-29 08:48 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050380 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-29 08:57 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2312180661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-29 08:57 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2312180957 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 09:00 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180958 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 09:00 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180974 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-29 09:08 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180876 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-29 09:08 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180881 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-29 09:11 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2401050252 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-29 09:11 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2401050119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-29 09:20 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2312180969 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-29 09:21 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2401050374 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-29 09:23 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180595 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.329 | Pass |  | n/a
2024-01-29 09:23 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180660 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 09:32 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180657 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-29 09:32 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050367 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-29 09:34 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2401050116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-29 09:34 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2312180980 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-29 09:42 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180638 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-29 09:45 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180971 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-29 10:36 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050360 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-29 10:46 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2312180956 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-29 11:50 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2312180953 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-29 11:54 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050324 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-29 12:01 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2312180598 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-29 12:03 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2312180972 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-29 12:09 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2312180949 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-29 12:10 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050361 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.38 | Pass |  | n/a
2024-01-29 12:16 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2312180650 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.366 | Pass |  | n/a
2024-01-29 12:19 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-29 12:27 | stmstation3@avive.life | W1187 | Agree | Leonora Cruz | 2401050359 | Pass | Fail | [10 empty] | screen on discolored
2024-01-29 12:36 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050159 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-29 12:39 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050362 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 12:41 | stmstation3@avive.life | W1187 | Agree | Leonora Cruz | 2401050354 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-29 12:46 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2312180987 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.389 | Pass |  | n/a
2024-01-29 12:51 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050383 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-29 12:53 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050377 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.352 | Pass |  | n/a
2024-01-29 12:56 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050168 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-29 13:03 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050381 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-29 13:04 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050243 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.404 | Pass |  | n/a
2024-01-29 13:10 | stmstation4@avive.life | W1187 | Agree | Leonora Cruz | 2401050376 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 13:10 | stmstation3@avive.life | W1187 | Agree | Leonora Cruz | 2312180882 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-29 13:14 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-29 13:18 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2312180983 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-29 13:22 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2312180658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-29 13:29 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050385 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.411 | Pass |  | n/a
2024-01-29 13:37 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2312180655 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.403 | Pass |  | n/a
2024-01-29 13:37 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-29 13:42 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-29 13:45 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-29 13:50 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050369 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-29 13:50 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180875 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 13:52 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-29 13:54 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2312180954 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-29 13:57 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050250 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-29 14:05 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180875 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-29 14:12 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180596 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-29 14:20 | stmstation4@avive.life | W1187 | Agree | Ananya Jain | 2401050359 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.389 | Pass |  | Reworked on Display
2024-01-29 14:27 | stmstation2@avive.life | W1187 | Agree | Ireri Moreno | 2312180955 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-29 14:27 | stmstation1@avive.life | W1187 | Agree | Ireri Moreno | 2401050247 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-29 14:28 | stmstation3@avive.life | W1187 | Agree | Ireri Moreno | 2401050236 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-29 14:28 | stmstation4@avive.life | W1187 | Agree | Ireri Moreno | 2401050219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 14:37 | stmstation2@avive.life | W1187 | Agree | Ireri Moreno | 2401050120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-29 14:37 | stmstation1@avive.life | W1187 | Agree | Ireri Moreno | 2401050110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-29 14:38 | stmstation3@avive.life | W1187 | Agree | Ireri Moreno | 2312180879 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-29 14:42 | stmstation4@avive.life | W1187 | Agree | Ireri Moreno | 2401050220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-29 14:49 | stmstation2@avive.life | W1187 | Agree | Ireri Moreno | 2312180952 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-29 14:49 | stmstation1@avive.life | W1187 | Agree | Ireri Moreno | 2401050249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-29 14:50 | stmstation3@avive.life | W1187 | Agree | Ireri Moreno | 2312180951 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-29 14:51 | stmstation4@avive.life | W1187 | Agree | Ireri Moreno | 2401050029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-29 15:06 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180950 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-29 15:07 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180877 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-29 15:09 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-29 15:11 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-29 15:14 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-29 15:17 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-29 15:19 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 15:25 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 15:26 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-29 15:28 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180654 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-29 15:32 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2312180967 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-29 15:35 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-29 15:38 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050172 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-29 15:39 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180689 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.329 | Pass |  | n/a
2024-01-29 15:41 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-29 15:44 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050335 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-29 15:46 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-29 15:49 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2312180968 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-29 15:56 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180966 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-29 16:00 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180975 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-29 16:03 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180639 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.331 | Pass |  | n/a
2024-01-29 16:04 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2312180880 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-29 16:07 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2312180807 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-29 16:12 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050368 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-29 16:15 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180964 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-29 16:16 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180970 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-29 16:19 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2312180659 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.357 | Pass |  | n/a
2024-01-29 16:20 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050375 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-29 16:22 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180979 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-29 16:26 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050370 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-30 09:40 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050232 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-30 09:43 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180939 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-30 09:45 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050238 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-30 09:47 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2312180653 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-30 09:52 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050239 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-30 09:52 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180999 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-30 09:57 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050113 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-30 10:00 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180852 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-30 10:01 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2312180934 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.395 | Pass |  | n/a
2024-01-30 10:03 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2312180858 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.317 | Pass |  | n/a
2024-01-30 10:07 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2312180900 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-30 10:09 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180940 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-30 10:13 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-30 10:14 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2312180810 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-30 10:19 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2401050035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-30 10:24 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-30 10:27 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2401050365 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-30 10:28 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2312180809 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-30 10:34 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2312180803 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-30 10:36 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2312180945 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-30 10:41 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2401050378 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-30 10:43 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-30 10:45 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050363 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-30 10:49 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2401050356 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-30 10:52 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050371 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-30 10:54 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-30 10:59 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050364 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-30 10:59 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050235 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-30 11:01 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180808 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-30 11:02 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2312180802 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-30 12:18 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2312180651 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-30 12:18 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2312180805 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-30 12:20 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180649 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.361 | Pass |  | n/a
2024-01-30 12:33 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050240 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-30 12:34 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050023 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-30 12:35 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2312180806 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-30 12:41 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-30 12:49 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180828 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-30 12:51 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050326 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-30 12:55 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050158 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-30 12:57 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180914 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-30 13:00 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050373 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-30 13:01 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050325 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-30 13:03 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050353 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-30 13:05 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050242 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.354 | Pass |  | n/a
2024-01-30 13:15 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050372 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-30 13:22 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2312180822 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.366 | Pass |  | n/a
2024-01-30 13:26 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.366 | Pass |  | n/a
2024-01-30 13:35 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180978 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-30 13:42 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-30 13:46 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2312180889 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-30 13:48 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-30 13:55 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180959 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-30 14:00 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-30 14:05 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-30 14:08 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-30 14:11 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2312180889 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-30 14:14 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180652 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-30 14:18 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-30 14:23 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180821 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-30 14:26 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-30 14:29 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2312180976 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-30 14:30 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180899 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-30 15:17 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2312181000 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-30 15:20 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180860 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.335 | Pass |  | n/a
2024-01-30 15:24 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050160 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-30 15:28 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-30 15:31 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050358 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-30 15:32 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-30 15:40 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050063 | Pass | Fail | [10 empty] | Screen Defect
2024-01-30 15:44 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.35 | Pass |  | N/A
2024-01-30 15:45 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2312180993 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.343 | Pass |  | N/A
2024-01-30 15:49 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050138 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-30 15:50 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180883 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.337 | Pass |  | N/A
2024-01-30 15:52 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-30 15:55 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180885 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-30 15:57 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-30 16:01 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.349 | Pass |  | n/a
2024-01-30 16:06 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-30 16:08 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050066 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-30 16:12 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180942 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-30 16:14 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-30 16:17 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050065 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-30 16:19 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-30 16:21 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180825 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-30 16:25 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.348 | Pass |  | n/a
2024-01-30 16:28 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.332 | Pass |  | n/a
2024-01-30 16:31 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050024 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-30 16:33 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-30 16:35 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050182 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-30 16:36 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-31 06:39 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050124 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-31 06:40 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-31 06:42 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050213 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-31 06:51 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-31 08:57 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050191 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-31 08:58 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-31 08:58 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-31 09:19 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-31 09:26 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050329 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-31 09:30 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-31 09:32 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2401050212 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-31 09:37 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-31 09:38 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-31 09:43 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-31 09:57 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-31 09:59 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.395 | Pass |  | n/a
2024-01-31 10:01 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2401050135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-31 10:08 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050245 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-31 10:16 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-31 12:10 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-01-31 12:55 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-31 15:05 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2312180887 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-31 15:16 | stmstation2@avive.life | W1187 | Agree | Precious Ungco | 2401050104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-31 15:17 | stmstation1@avive.life | W1187 | Agree | Precious Ungco | 2312180844 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.353 | Pass |  | n/a
2024-01-31 15:26 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2312180990 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-31 15:28 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2401050106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-31 15:34 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2312180991 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-31 15:35 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2401050154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.347 | Pass |  | n/a
2024-01-31 15:44 | stmstation2@avive.life | W1187 | Agree | Precious Ungco | 2401050328 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-31 15:45 | stmstation1@avive.life | W1187 | Agree | Precious Ungco | 2401050151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-31 15:47 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2312180995 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.333 | Pass |  | n/a
2024-01-31 15:52 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2401050086 | Pass | Fail | [10 empty] | LCD not working
2024-01-31 15:57 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2401050144 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-31 15:58 | stmstation1@avive.life | W1187 | Agree | Precious Ungco | 2401050145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-31 15:59 | stmstation2@avive.life | W1187 | Agree | Precious Ungco | 2401050147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-31 16:00 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2401050141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-31 16:07 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2401050031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-31 16:08 | stmstation2@avive.life | W1187 | Agree | Precious Ungco | 2401050190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-31 16:17 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2401050064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.393 | Pass |  | n/a
2024-01-31 16:31 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2401050068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.355 | Pass |  | n/a
2024-01-31 16:39 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2401050209 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.346 | Pass |  | n/a
2024-01-31 16:42 | stmstation1@avive.life | W1187 | Agree | Precious Ungco | 2401050246 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-31 16:43 | stmstation2@avive.life | W1187 | Agree | Precious Ungco | 2312180804 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-31 16:46 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2312180600 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-31 16:47 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2401050089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-31 16:54 | stmstation1@avive.life | W1187 | Agree | Precious Ungco | 2401050090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.337 | Pass |  | n/a
2024-01-31 16:55 | stmstation2@avive.life | W1187 | Agree | Precious Ungco | 2401050357 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-31 16:56 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2401050130 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2024-01-31 16:57 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2401050184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.329 | Pass |  | n/a
2024-01-31 17:04 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2401050366 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.34 | Pass |  | n/a
2024-01-31 17:06 | stmstation1@avive.life | W1187 | Agree | Precious Ungco | 2401050187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-31 17:07 | stmstation2@avive.life | W1187 | Agree | Precious Ungco | 2401050137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.339 | Pass |  | n/a
2024-01-31 17:08 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2401050163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.351 | Pass |  | n/a
2024-01-31 17:24 | stmstation2@avive.life | W1187 | Agree | Precious Ungco | 2401050142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-01-31 18:26 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2312180906 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.342 | Pass |  | n/a
2024-01-31 18:27 | stmstation1@avive.life | W1187 | Agree | Precious Ungco | 2312180909 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-31 18:29 | stmstation2@avive.life | W1187 | Agree | Precious Ungco | 2312180911 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-31 18:30 | stmstation4@avive.life | W1187 | Agree | Precious Ungco | 2312180965 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.336 | Pass |  | n/a
2024-01-31 18:34 | stmstation3@avive.life | W1187 | Agree | Precious Ungco | 2401050196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-01-31 18:35 | stmstation1@avive.life | W1187 | Agree | Precious Ungco | 2312180912 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.338 | Pass |  | n/a
2024-01-31 18:38 | stmstation2@avive.life | W1187 | Agree | Precious Ungco | 2312180824 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-01 07:32 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-01 07:35 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2312180793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-01 07:51 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-01 07:55 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-01 08:20 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050081 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-01 08:25 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.371 | Pass |  | n/a
2024-02-01 08:27 | stmstation3@avive.life | W1187 | Agree | Leonora Cruz | 2401050351 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-01 08:31 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-01 08:37 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.358 | Pass |  | n/a
2024-02-01 08:37 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180904 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-01 08:42 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-01 08:46 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050327 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-01 08:49 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-01 08:49 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-01 09:15 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050165 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-01 09:16 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-01 09:21 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2401050161 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.405 | Pass |  | n/a
2024-02-01 09:24 | stmstation3@avive.life | W1187 | Agree | Leonora Cruz | 2401050153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.38 | Pass |  | n/a
2024-02-01 09:38 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-01 09:48 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2312180908 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 1 | Fail | value to high
2024-02-01 10:03 | stmstation4@avive.life | W1187 | Agree | Leonora Cruz | 2401050019 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.41 | Pass |  | n/a
2024-02-01 10:12 | stmstation4@avive.life | W1187 | Agree | Leonora Cruz | 2401050234 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.354 | Pass |  | n/a
2024-02-01 10:13 | stmstation3@avive.life | W1187 | Agree | Leonora Cruz | 2401050111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.339 | Pass |  | N/A
2024-02-01 10:17 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.425 | Pass |  | n/a
2024-02-01 10:23 | stmstation3@avive.life | W1187 | Agree | Leonora Cruz | 2401050097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.337 | Pass |  | N/A
2024-02-01 10:23 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.412 | Pass |  | n/a
2024-02-01 10:25 | stmstation4@avive.life | W1187 | Agree | Leonora Cruz | 2401050067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-01 10:34 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-01 10:37 | stmstation1@avive.life | W1187 | Agree | Leonora Cruz | 2401050062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.402 | Pass |  | n/a
2024-02-01 10:48 | stmstation2@avive.life | W1187 | Agree | Leonora Cruz | 2401050061 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.356 | Pass |  | n/a
2024-02-01 13:00 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-01 13:03 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180326 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.332 | Pass |  | n/a
2024-02-01 13:04 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-01 13:06 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2312180401 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-01 13:08 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2312180372 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-01 13:12 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180260 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-01 13:15 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2312180216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-01 13:16 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2312180219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-01 13:18 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2312180257 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-01 13:21 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180944 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-01 13:24 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2312180982 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-01 13:25 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2312180656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-01 13:33 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2312180963 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-01 13:38 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2312180994 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-01 13:40 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180962 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-01 13:41 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2312180984 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-01 13:47 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2312180996 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-01 13:53 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2401050206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-01 14:04 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2312180799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-01 14:05 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2312180884 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.337 | Pass |  | N/A
2024-02-01 14:06 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2312180998 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.348 | Pass |  | N/A
2024-02-01 14:09 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180961 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.34 | Pass |  | N/A
2024-02-01 14:13 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2312180981 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-01 14:17 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2312180985 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.337 | Pass |  | N/A
2024-02-01 14:18 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2312180997 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.354 | Pass |  | N/A
2024-02-01 14:18 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-01 14:21 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2312180886 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-01 14:59 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-01 15:03 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2312180264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.349 | Pass |  | N/A
2024-02-01 15:08 | stmstation4@avive.life | W1187 | Agree | Christina Cruz | 2312180800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.338 | Pass |  | N/A
2024-02-01 15:08 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180327 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-01 15:16 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-01 15:17 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180374 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-01 15:20 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2312180402 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.335 | Pass |  | N/A
2024-02-01 15:24 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180393 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-01 15:25 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180271 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-01 15:31 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-01 15:32 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050201 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-01 15:34 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2312180986 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-01 15:37 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180263 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-01 15:40 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050103 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-01 15:44 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2312180973 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.385 | Pass |  | n/a
2024-02-01 15:46 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180977 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-01 15:52 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2401050352 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-01 15:55 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-01 15:55 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050330 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-01 16:01 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2401050355 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-01 16:05 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-01 16:06 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180209 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-01 16:12 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2312180891 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-01 16:13 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-01 16:14 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180397 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-01 16:23 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050197 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-01 16:23 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180905 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-01 16:28 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2401050174 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-01 16:36 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-01 16:37 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180663 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-01 16:38 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2401050175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-01 16:43 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-01 16:46 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-01 16:47 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2312180892 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-01 16:54 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-01 16:54 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180801 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-01 16:58 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2401050096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.352 | Pass |  | N/A
2024-02-01 17:05 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-01 17:05 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-01 17:06 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2401050080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-02 08:34 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2312180261 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-02 09:50 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2312180960 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-02 14:10 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2401050162 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.337 | Pass |  | n/A
2024-02-02 14:12 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050082 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-02 14:12 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050200 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-02 14:13 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2401050183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-02 14:16 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050167 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-02 14:16 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-02 14:17 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2401050077 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-02 14:21 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2401050073 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-02 14:24 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2312180910 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-02 14:25 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.353 | Pass |  | n/a
2024-02-02 14:26 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-02 14:27 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2312180826 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-02 14:29 | stmstation3@avive.life | W1187 | Agree | Valeria Manzo | 2401050186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-02 14:30 | stmstation4@avive.life | W1187 | Agree | Valeria Manzo | 2401050105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-02 14:31 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050331 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-02 14:33 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050334 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-02 14:34 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-02 14:36 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-02 14:37 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050075 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-02 14:38 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-02 14:41 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-02 14:45 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-02 14:47 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-02 14:48 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2312180913 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-02 14:51 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-02 14:52 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.4 | Pass |  | n/a
2024-02-02 14:53 | stmstation2@avive.life | W1187 | Agree | Christina Cruz | 2401050233 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-02 14:55 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050083 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-02 14:55 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180992 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-02 14:56 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050379 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-02 15:00 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-02 15:00 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-02 15:00 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.315 | Pass |  | n/a
2024-02-02 15:01 | stmstation1@avive.life | W1187 | Agree | Valeria Manzo | 2401050332 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-02 15:04 | stmstation2@avive.life | W1187 | Agree | Valeria Manzo | 2401050170 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-02 15:05 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-02 15:08 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-02 15:13 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050241 | Pass | Fail | [10 empty] | Not turning on
2024-02-02 15:14 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2401050204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-02 15:26 | stmstation4@avive.life | W1187 | Agree | Liza Ordillo | 2312180908 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.337 | Pass |  | Reworked
2024-02-02 15:35 | stmstation3@avive.life | W1187 | Agree | Liza Ordillo | 2401050063 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-02 15:42 | stmstation1@avive.life | W1187 | Agree | Liza Ordillo | 2312180907 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Fail |  |  | no SIM
2024-02-02 15:47 | stmstation2@avive.life | W1187 | Agree | Liza Ordillo | 2401050018 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | no cell no sim
2024-02-02 17:27 | stmstation1@avive.life | W1187 | Agree | Christina Cruz | 2401050193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-05 06:58 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050620 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 07:00 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050686 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 07:01 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050595 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-05 07:03 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050596 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 07:07 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-05 07:08 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050649 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-05 07:10 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050434 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 07:14 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050652 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 07:16 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050453 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 07:17 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050585 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-05 07:19 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050594 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 07:23 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050447 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 07:25 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050581 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 07:27 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050462 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-05 07:28 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050680 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-05 07:31 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050508 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 07:34 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050509 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-05 07:35 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050510 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-05 07:37 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050417 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 07:42 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050690 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-05 07:43 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-05 07:44 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050583 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 07:46 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050684 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 07:49 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050580 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-05 07:53 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 07:54 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050683 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-05 07:55 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050653 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-05 08:00 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 08:02 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050654 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-05 08:03 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050416 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-05 08:05 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050650 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 08:08 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050449 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 08:10 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050651 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 08:11 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050681 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 08:12 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050507 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-05 08:15 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050685 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 08:21 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.332 | Pass |  | n/a
2024-02-05 08:22 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050682 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 08:22 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050535 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 08:23 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050455 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 09:08 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050450 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-05 09:20 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050639 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-05 09:25 | stmstation3@avive.life | W1187 | Agree | Christina Cruz | 2401050086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.346 | Pass |  | Reworked
2024-02-05 09:27 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050602 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 09:35 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050266 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-05 09:40 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050603 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 09:40 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050601 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 09:52 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050297 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-05 09:52 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050600 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-05 09:53 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050675 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 10:00 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050669 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-05 10:01 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050448 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 10:06 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050632 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-05 10:07 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050640 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-05 10:11 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050670 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 10:20 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050393 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 10:22 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050292 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-05 10:22 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 10:29 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050339 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-05 10:36 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-05 10:36 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050347 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-05 10:38 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050593 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-05 10:39 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050604 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 11:18 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050293 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 11:23 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050642 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 11:26 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 11:35 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 11:38 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050631 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-05 11:48 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050671 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-05 11:49 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050637 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 11:49 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 11:50 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050598 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 12:00 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 12:10 | stmstation1@avive.life | W1188 | Agree | Valeria manzo | 2401050674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-05 12:11 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050573 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 12:15 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050343 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 12:16 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050688 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-05 12:17 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050467 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 12:20 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050630 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 12:23 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050689 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 12:24 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050456 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 12:25 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050432 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 12:29 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050425 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 12:32 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050433 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-05 12:34 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050679 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-05 12:36 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050349 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 12:38 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 12:39 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050457 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 12:42 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050473 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-05 12:44 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 12:46 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050469 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 12:48 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050346 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-05 12:50 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050451 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 12:52 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-05 12:53 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050599 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.353 | Pass |  | n/a
2024-02-05 12:56 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050458 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 12:56 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050552 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-05 12:58 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.331 | Pass |  | n/a
2024-02-05 13:07 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050638 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 13:10 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-05 13:11 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050655 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 13:13 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050634 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-05 13:17 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050007 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 13:19 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050538 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 13:22 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050298 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 13:23 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050672 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-05 13:31 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 13:34 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 13:52 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050668 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-05 13:53 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050646 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 13:54 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-05 14:01 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-05 14:04 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050556 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-05 14:07 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-05 14:10 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-05 14:14 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050566 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-05 14:16 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 14:30 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050530 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 14:40 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050476 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 14:49 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050663 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-05 14:52 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 14:53 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050311 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-05 14:56 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 15:01 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050678 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 15:03 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050648 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 15:06 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050562 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 15:09 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050667 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-05 15:12 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050572 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-05 15:15 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050575 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 15:18 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-05 15:21 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050618 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-05 15:25 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050435 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-05 15:26 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050294 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-05 15:34 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.361 | Pass |  | n/a
2024-02-05 15:36 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050533 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-05 15:38 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050531 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-05 15:43 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050617 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-05 15:55 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050622 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-05 15:58 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050569 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 16:01 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050563 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-05 16:21 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050534 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-05 16:23 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.345 | Pass |  | N/A
2024-02-05 16:23 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050559 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.337 | Pass |  | N/A
2024-02-05 16:24 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050587 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-05 16:30 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.344 | Pass |  | N/A
2024-02-05 16:30 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050424 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.342 | Pass |  | N/A
2024-02-05 16:38 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-05 16:44 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050536 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 16:46 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050664 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-05 16:49 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050567 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-06 06:59 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.453 | Pass |  | n/a
2024-02-06 07:06 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-06 07:08 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050414 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-06 07:14 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.398 | Pass |  | n/a
2024-02-06 07:15 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-06 07:18 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.387 | Pass |  | n/a
2024-02-06 07:21 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.371 | Pass |  | n/a
2024-02-06 07:24 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050500 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.394 | Pass |  | n/a
2024-02-06 07:27 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050496 | Pass | Fail | [10 empty] | Failed screen won't turn on.
2024-02-06 07:39 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.397 | Pass |  | n/a
2024-02-06 07:47 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050570 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.391 | Pass |  | n/a
2024-02-06 07:57 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050551 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-06 08:05 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050291 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-06 08:08 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050666 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.423 | Pass |  | n/a
2024-02-06 08:14 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.408 | Pass |  | n/a
2024-02-06 08:15 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050571 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.357 | Pass |  | n/a
2024-02-06 08:22 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050568 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.401 | Pass |  | n/a
2024-02-06 08:23 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050565 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.376 | Pass |  | n/a
2024-02-06 08:26 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050662 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-06 08:29 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.463 | Pass |  | n/a
2024-02-06 08:31 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050561 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.418 | Pass |  | n/a
2024-02-06 08:37 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050657 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-06 08:38 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050420 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.407 | Pass |  | n/a
2024-02-06 08:43 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-06 08:44 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050485 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-06 08:45 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050626 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.361 | Pass |  | n/a
2024-02-06 08:56 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050405 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-06 08:56 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050423 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-06 08:59 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050542 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 09:05 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050419 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-06 09:05 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050659 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-06 09:07 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050421 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-06 09:15 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050403 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-06 09:16 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050402 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-06 09:20 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050422 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-06 09:22 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-06 09:24 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-06 09:25 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050406 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-06 09:26 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050281 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-06 09:28 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050283 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-06 09:29 | stmstation3@avive.life | W1188 | Agree | Precious Ungco | 2401050284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-06 09:33 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050499 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-06 09:38 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-06 09:41 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-06 09:45 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 09:45 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050279 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 09:45 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050577 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-06 09:48 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-06 09:52 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.353 | Pass |  | n/a
2024-02-06 09:55 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-06 09:57 | stmstation4@avive.life | W1188 | Agree | Precious Ungco | 2401050492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.331 | Pass |  | n/a
2024-02-06 09:58 | stmstation3@avive.life | W1188 | Agree | Precious Ungco | 2401050621 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 10:02 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050624 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-06 10:03 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050660 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-06 10:14 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050514 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 10:21 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050494 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-06 10:22 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050412 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-06 10:23 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 10:24 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-06 10:28 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050550 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-06 10:28 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 10:40 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050578 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-06 10:40 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050436 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-06 10:44 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050511 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-06 10:45 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050418 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-06 10:50 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-06 10:53 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050431 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-06 10:57 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050576 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-06 10:59 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050546 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-06 11:01 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050544 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-06 11:02 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050519 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-06 11:03 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050415 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-06 11:08 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-06 11:12 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050408 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-06 11:13 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-06 11:14 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050520 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-06 11:16 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050549 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-06 11:22 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-06 11:23 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050426 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-06 11:23 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-06 11:23 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050547 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 11:29 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050615 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 11:30 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050623 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-06 11:35 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050404 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-06 11:41 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-06 11:45 | stmstation1@avive.life | W1188 | Agree | Liza Ordillo | 2401050278 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-06 11:47 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050399 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-06 11:52 | stmstation3@avive.life | W1188 | Agree | Liza Ordillo | 2401050400 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-06 11:54 | stmstation2@avive.life | W1188 | Agree | Liza Ordillo | 2401050427 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-06 11:56 | stmstation4@avive.life | W1188 | Agree | Liza Ordillo | 2401050028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-06 12:45 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050407 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-06 12:46 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050398 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 12:48 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 12:51 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050475 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-06 12:53 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-06 12:54 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-06 12:58 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050471 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-06 13:02 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050480 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-06 13:03 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 13:07 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050484 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-06 13:12 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050479 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-06 13:12 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050518 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-06 13:14 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050411 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-06 13:22 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050498 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-06 13:26 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050310 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 13:27 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050397 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-06 13:30 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050305 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-06 13:32 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-06 13:33 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050477 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-06 13:36 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-06 13:38 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-06 13:40 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050497 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-06 13:43 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-06 13:44 | stmstation2@avive.life | W1188 | Agree | Valeria Manzo | 2401050345 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-06 13:45 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050314 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 13:45 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050342 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-06 14:06 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 14:07 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050413 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-06 14:24 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2401050503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-06 14:24 | stmstation2@avive.life | W1188 | Agree | Christina Cruz | 2401050487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-06 14:27 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-06 14:27 | stmstation4@avive.life | W1188 | Agree | Christina Cruz | 2401050529 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-06 15:25 | stmstation2@avive.life | W1188 | Agree | Ananya Jain | 2401050496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.334 | Pass |  | Reworked on Screen
2024-02-07 06:35 | stmstation2@avive.life | W1188 | Agree | Leonora Cruz | 2401050470 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-07 06:41 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.374 | Pass |  | n/a
2024-02-07 06:43 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.436 | Pass |  | n/a
2024-02-07 06:53 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.381 | Pass |  | n/a
2024-02-07 06:54 | stmstation2@avive.life | W1188 | Agree | Leonora Cruz | 2401050313 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.364 | Pass |  | n/a
2024-02-07 07:21 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-07 07:25 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.413 | Pass |  | n/a
2024-02-07 07:37 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050286 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-07 07:57 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050527 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-07 08:04 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.355 | Pass |  | n/a
2024-02-07 08:14 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050472 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.409 | Pass |  | n/a
2024-02-07 08:15 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050321 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 08:21 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050317 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.378 | Pass |  | n/a
2024-02-07 08:24 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050288 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 08:25 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050322 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-07 08:29 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050320 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-07 08:39 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050316 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-07 08:42 | stmstation2@avive.life | W1188 | Agree | Leonora Cruz | 2401050515 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.675 | Pass |  | n/a
2024-02-07 08:45 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-07 08:50 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050287 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.373 | Pass |  | n/a
2024-02-07 08:51 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050495 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-07 09:00 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050478 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 09:04 | stmstation2@avive.life | W1188 | Agree | Leonora Cruz | 2401050482 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.404 | Pass |  | n/a
2024-02-07 09:10 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 09:12 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.357 | Pass |  | n/a
2024-02-07 09:13 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.366 | Pass |  | n/a
2024-02-07 09:36 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-07 09:36 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050338 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 09:37 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050224 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-07 09:37 | stmstation2@avive.life | W1188 | Agree | Leonora Cruz | 2401050521 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.414 | Pass |  | n/a
2024-02-07 10:11 | stmstation2@avive.life | W1188 | Agree | Leonora Cruz | 2401050454 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.406 | Pass |  | n/a
2024-02-07 10:14 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-07 10:19 | stmstation3@avive.life | W1188 | Agree | Leonora Cruz | 2401050525 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 10:20 | stmstation1@avive.life | W1188 | Agree | Leonora Cruz | 2401050307 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.407 | Pass |  | n/a
2024-02-07 10:22 | stmstation2@avive.life | W1188 | Agree | Leonora Cruz | 2401050052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 10:25 | stmstation4@avive.life | W1188 | Agree | Leonora Cruz | 2401050258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-07 10:37 | stmstation3@avive.life | W1188 | Agree | Precious Ungco | 2401050344 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-07 10:39 | stmstation4@avive.life | W1188 | Agree | Precious Ungco | 2401050382 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-07 10:42 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050516 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-07 10:43 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050303 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-07 10:48 | stmstation3@avive.life | W1188 | Agree | Precious Ungco | 2401050299 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 10:48 | stmstation4@avive.life | W1188 | Agree | Precious Ungco | 2401050336 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.33 | Pass |  | n/a
2024-02-07 10:50 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050340 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 10:51 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050309 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-07 10:58 | stmstation3@avive.life | W1188 | Agree | Precious Ungco | 2401050229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 11:00 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.329 | Pass |  | n/a
2024-02-07 11:01 | stmstation4@avive.life | W1188 | Agree | Precious Ungco | 2401050255 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-07 11:03 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050308 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-07 11:32 | stmstation3@avive.life | W1188 | Agree | Precious Ungco | 2401050277 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 11:32 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050268 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-07 11:33 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050269 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 11:34 | stmstation4@avive.life | W1188 | Agree | Precious Ungco | 2401050276 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-07 11:40 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-07 11:43 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 11:44 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050387 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 11:44 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050526 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-07 11:46 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-07 11:50 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-07 11:50 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 11:50 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050337 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 11:52 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050263 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-07 11:58 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050256 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.332 | Pass |  | n/a
2024-02-07 12:00 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050395 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-07 12:01 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050392 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-07 12:02 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050230 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 12:07 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050261 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-07 12:07 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050300 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-07 12:07 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-07 12:10 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.316 | Pass |  | n/a
2024-02-07 12:13 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050306 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 12:14 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-07 12:16 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050262 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 12:18 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050396 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 12:21 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050341 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-07 12:22 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050315 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-07 12:23 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050319 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 12:24 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050259 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 12:29 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050390 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-07 12:29 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 12:30 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050318 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-07 12:32 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050394 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 12:35 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.331 | Pass |  | n/a
2024-02-07 12:35 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-07 12:36 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050275 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-07 12:37 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050302 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-07 12:40 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 12:41 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-07 12:41 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050228 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-07 12:47 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050260 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-07 12:50 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050389 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 12:51 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050267 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-07 12:53 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.331 | Pass |  | n/a
2024-02-07 12:53 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-07 12:55 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 12:55 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050295 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-07 12:59 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050227 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-07 13:00 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050296 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.332 | Pass |  | n/a
2024-02-07 13:01 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050253 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 13:04 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050323 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-07 13:05 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050410 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-07 13:07 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 13:07 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050301 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 13:11 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 13:14 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050391 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-07 13:16 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2312180353 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-07 13:19 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050304 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-07 13:22 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050409 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 13:22 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050350 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-07 13:24 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050512 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 13:27 | stmstation3@avive.life | W1188 | Agree | Valeria Manzo | 2401050312 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-07 13:29 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 13:33 | stmstation3@avive.life | W1188 | Agree | Precious Ungco | 2401050231 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 344 | Pass |  | n/a
2024-02-07 13:35 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2312180364 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-07 13:38 | stmstation4@avive.life | W1188 | Agree | Valeria Manzo | 2401050058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.313 | Pass |  | n/a
2024-02-07 13:42 | stmstation3@avive.life | W1188 | Agree | Precious Ungco | 2401050005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-07 13:45 | stmstation4@avive.life | W1188 | Agree | Precious Ungco | 2401050009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-07 13:46 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-07 13:48 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2312180358 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-07 13:55 | stmstation1@avive.life | W1188 | Agree | Precious Ungco | 2401050041 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 13:56 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-07 13:57 | stmstation3@avive.life | W1188 | Agree | Precious Ungco | 2401050053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-07 13:58 | stmstation4@avive.life | W1188 | Agree | Precious Ungco | 2401050270 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-07 14:03 | stmstation2@avive.life | W1188 | Agree | Precious Ungco | 2401050388 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 14:48 | stmstation1@avive.life | W1188 | Agree | Valeria Manzo | 2401050401 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-07 15:25 | stmstation1@avive.life | W1188 | Agree | Christina Cruz | 2312180249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-08 13:57 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050271 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Fail |  |  | No Sim
2024-02-08 14:16 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2312180357 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Fail |  |  | No Sim
2024-02-08 14:21 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050257 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Fail |  |  | No Sim
2024-02-08 14:24 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2312180250 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Fail |  |  | No Sim
2024-02-08 16:20 | stmstation3@avive.life | W1188 | Agree | Christina Cruz | 2401050265 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-09 11:34 | stmstation1@avive.life | W1192 | Agree | Valeria Manzo | 2401050848 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-09 11:39 | stmstation4@avive.life | W1192 | Agree | Valeria Manzo | 2401050763 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-09 11:39 | stmstation3@avive.life | W1192 | Agree | Valeria Manzo | 2401050991 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-09 11:41 | stmstation2@avive.life | W1192 | Agree | Valeria Manzo | 2401050770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-09 11:44 | stmstation1@avive.life | W1192 | Agree | Valeria Manzo | 2401050911 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-09 11:46 | stmstation4@avive.life | W1192 | Agree | Valeria Manzo | 2401050902 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-09 11:47 | stmstation3@avive.life | W1192 | Agree | Valeria Manzo | 2401050899 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-09 11:50 | stmstation2@avive.life | W1192 | Agree | Valeria Manzo | 2401050897 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-09 11:51 | stmstation1@avive.life | W1192 | Agree | Valeria Manzo | 2401050895 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-09 11:54 | stmstation4@avive.life | W1192 | Agree | Valeria Manzo | 2401290049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-09 11:59 | stmstation1@avive.life | W1192 | Agree | Valeria Manzo | 2401050890 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-09 12:00 | stmstation2@avive.life | W1192 | Agree | Valeria Manzo | 2401290045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-09 12:02 | stmstation3@avive.life | W1192 | Agree | Valeria Manzo | 2401290054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-09 12:03 | stmstation4@avive.life | W1192 | Agree | Valeria Manzo | 2401290015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-09 12:06 | stmstation1@avive.life | W1192 | Agree | Valeria Manzo | 2401290018 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-09 12:08 | stmstation2@avive.life | W1192 | Agree | Valeria Manzo | 2401290050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-09 12:11 | stmstation4@avive.life | W1192 | Agree | Valeria Manzo | 2401050789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-09 12:11 | stmstation3@avive.life | W1192 | Agree | Valeria Manzo | 2401290048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-09 12:14 | stmstation2@avive.life | W1192 | Agree | Valeria Manzo | 2401050856 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-09 12:20 | stmstation4@avive.life | W1192 | Agree | Valeria Manzo | 2401050844 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-09 12:21 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050885 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-09 12:22 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050852 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-09 12:28 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050854 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-09 12:28 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050853 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-09 12:30 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050893 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-09 12:31 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-09 13:53 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-09 13:55 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401050894 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-09 13:58 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050847 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-09 14:00 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050842 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-09 14:11 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050845 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-09 14:12 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050908 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-09 14:13 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050874 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-09 14:14 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050850 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-09 14:21 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050898 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-09 14:22 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050896 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.355 | Pass |  | n/a
2024-02-09 14:23 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050906 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-09 14:24 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050849 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-09 14:32 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-09 14:33 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050904 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-09 14:34 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050801 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-09 14:35 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050738 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-09 14:42 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050767 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-09 14:43 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050804 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-09 14:44 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050760 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-09 14:45 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050872 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-09 14:56 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050710 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-09 14:58 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050719 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-09 15:00 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050768 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-09 15:01 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050868 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 09:12 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050764 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-12 09:13 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050900 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-12 09:14 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050754 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-12 09:15 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050756 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-12 09:23 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050739 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-12 09:24 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050745 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-12 09:25 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050752 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-12 09:26 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050751 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-12 09:33 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050839 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-12 09:35 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401290036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-12 09:37 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.332 | Pass |  | n/a
2024-02-12 09:38 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050873 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 09:47 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-12 09:48 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-12 09:49 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-12 09:49 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401290037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-12 09:57 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401290026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-12 09:58 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050822 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 09:59 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401290025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-12 10:00 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050714 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-12 10:07 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050759 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-12 10:08 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050871 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-12 10:09 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050855 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-12 10:10 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050915 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-12 10:43 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050840 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.376 | Pass |  | n/a
2024-02-12 10:44 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050846 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-12 10:45 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050724 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 10:46 | stmstation4@avive.life | W1192 | Agree | Precious Ungco | 2401050901 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 10:52 | stmstation2@avive.life | W1192 | Agree | Precious Ungco | 2401050802 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-12 10:54 | stmstation3@avive.life | W1192 | Agree | Precious Ungco | 2401050905 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-12 10:55 | stmstation1@avive.life | W1192 | Agree | Precious Ungco | 2401050857 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 13:25 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050881 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-12 13:26 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050887 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-12 13:29 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401050883 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 13:33 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401050889 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-12 13:35 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050892 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-12 13:37 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401290027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-12 13:39 | stmstation3@avive.life | W1192 | Agree | Ireri Moreno | 2401050730 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-12 13:42 | stmstation4@avive.life | W1192 | Agree | Ireri Moreno | 2401050750 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-12 13:44 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-12 13:46 | stmstation3@avive.life | W1192 | Agree | Ireri Moreno | 2401050716 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 13:51 | stmstation4@avive.life | W1192 | Agree | Ireri Moreno | 2401050737 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-12 13:51 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050711 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-12 13:54 | stmstation3@avive.life | W1192 | Agree | Ireri Moreno | 2401290047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-12 13:57 | stmstation4@avive.life | W1192 | Agree | Ireri Moreno | 2401050757 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-12 13:59 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-12 14:01 | stmstation3@avive.life | W1192 | Agree | Ireri Moreno | 2401290023 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-12 14:05 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401290032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-12 14:08 | stmstation3@avive.life | W1192 | Agree | Ireri Moreno | 2401050755 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-12 14:08 | stmstation4@avive.life | W1192 | Agree | Ireri Moreno | 2401290002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-12 14:18 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050758 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-12 14:19 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050961 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 14:20 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401290038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-12 14:27 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050882 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-12 14:32 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.356 | Pass |  | n/a
2024-02-12 14:36 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050888 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-12 14:38 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050466 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-12 14:38 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050749 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-12 14:45 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050959 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-12 14:51 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050808 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-12 14:51 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050429 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-12 14:54 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401290001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-12 15:02 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050762 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-12 15:02 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050439 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-12 15:03 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050967 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-12 15:09 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050718 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-12 15:16 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401290011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-12 15:18 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050798 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-12 15:19 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401290021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-12 15:19 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050907 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 15:31 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050428 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-12 15:31 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050886 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-12 15:31 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050869 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-12 15:32 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050820 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-12 15:43 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050965 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-12 15:43 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050841 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-12 15:46 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050824 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-12 15:50 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050704 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-12 15:53 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050867 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-12 15:57 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050903 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-12 15:57 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-12 15:59 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050891 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.416 | Pass |  | n/a
2024-02-12 16:07 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050858 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-12 16:08 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401050843 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-12 16:15 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401290012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-12 16:17 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401050705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-12 16:21 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401290014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-12 16:23 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-12 16:25 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401290019 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-12 16:30 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401050461 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-12 16:30 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-12 16:34 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401050720 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-12 16:35 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050715 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-12 16:40 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401050698 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-12 16:41 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050465 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-12 16:46 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050765 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-12 16:49 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401290041 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-12 16:51 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050982 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-12 16:52 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401050468 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-12 16:58 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050766 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-12 17:01 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050463 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-12 17:02 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401290046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-13 08:41 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050805 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-13 08:43 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050960 | Pass | Fail | [10 empty] | no LCD
2024-02-13 08:44 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050957 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-13 08:46 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050797 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-13 08:55 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050940 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-13 09:01 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050949 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-13 09:07 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-13 09:14 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-13 09:15 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050732 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-13 09:18 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050860 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-13 09:19 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050859 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-13 09:27 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050944 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-13 09:29 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050952 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-13 09:33 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-13 09:37 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050973 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-13 09:49 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050774 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-13 10:06 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050919 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 10:11 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050936 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-13 10:11 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050947 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-13 10:13 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-13 10:25 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050780 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-13 10:25 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050777 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-13 10:29 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050775 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-13 10:29 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050735 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-13 11:03 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050934 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-13 11:03 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050933 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-13 11:04 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050863 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-13 11:04 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050927 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-13 11:16 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050922 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-13 11:17 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050918 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 11:21 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050610 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-13 11:21 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050459 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-13 11:29 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050943 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 11:32 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050938 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-13 11:33 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-13 11:42 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050606 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 11:44 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050877 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-13 11:50 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050939 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-13 11:50 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050727 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-13 11:58 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050917 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-13 12:36 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050778 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 12:38 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050807 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-13 12:41 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050795 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.359 | Pass |  | n/a
2024-02-13 12:45 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 12:50 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050776 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-13 12:56 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050785 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-13 12:59 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050784 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-13 13:10 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 13:11 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-13 13:16 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050761 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 13:19 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050733 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-13 13:23 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050921 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-13 13:27 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050912 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 14:27 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050931 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-13 14:29 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401290053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-13 14:37 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050926 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 14:41 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050779 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 14:44 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050430 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-13 14:54 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050769 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-13 14:56 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050950 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-13 15:14 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050942 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-13 15:14 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050951 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-13 15:23 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050920 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-13 15:23 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050792 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-13 15:33 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050444 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-13 15:36 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050875 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-13 15:43 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050870 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-13 15:44 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050876 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-13 15:48 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050953 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.362 | Pass |  | n/a
2024-02-13 15:54 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050925 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-13 15:58 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050928 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.354 | Pass |  | n/a
2024-02-13 15:59 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050464 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-13 16:08 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.651 | Pass |  | n/a
2024-02-13 16:15 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050605 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-13 16:24 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050972 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.294 | Pass |  | n/a
2024-02-13 16:25 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050609 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-13 16:32 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050440 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-13 16:36 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050736 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-14 09:36 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050728 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-14 09:37 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050923 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-14 09:41 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401050441 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-14 09:42 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050924 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-14 09:50 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050932 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-14 09:55 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401050916 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-14 09:56 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-14 09:58 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401050744 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.331 | Pass |  | n/a
2024-02-14 10:33 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050830 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-14 10:34 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050970 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-14 10:38 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050823 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-14 10:44 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050826 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-14 10:48 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050833 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-14 10:51 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050817 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-14 10:55 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050810 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-14 10:57 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050946 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-14 11:05 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050960 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.34 | Pass |  | rework performed
2024-02-14 11:06 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050836 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-14 11:14 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050948 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-14 11:18 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050721 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-14 11:22 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050743 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-14 11:22 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050818 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-14 11:40 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050941 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-14 11:47 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050813 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-14 11:48 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050748 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-14 11:50 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050723 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-14 11:53 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050831 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Fail |  |  | no sIM
2024-02-14 12:02 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050832 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-14 12:04 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050974 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-14 12:10 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050812 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-14 12:12 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050816 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-14 12:15 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050699 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-14 12:24 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050607 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-14 12:25 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050835 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.428 | Pass |  | n/a
2024-02-14 12:25 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050614 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-14 12:35 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401050809 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-14 13:15 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-14 13:16 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050837 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-14 13:18 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401050958 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-14 13:25 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050781 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-14 13:27 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050966 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.4 | Pass |  | n/a
2024-02-14 13:32 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050691 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-14 13:39 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050692 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-14 13:41 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050954 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.399 | Pass |  | n/a
2024-02-14 13:42 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050861 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-14 13:47 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050783 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-14 13:47 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050592 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-14 13:56 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050864 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-14 13:58 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050829 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-14 15:59 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050971 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-14 16:02 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050590 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-14 16:13 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401050588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-14 16:17 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050880 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-14 16:38 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050742 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-14 16:42 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401050955 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-15 06:42 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050956 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-15 06:52 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050913 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-15 06:52 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050878 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-15 06:56 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050746 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-15 07:02 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-15 07:06 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050747 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-15 07:10 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-15 07:11 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050828 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-15 07:22 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050935 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-15 07:27 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050460 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-15 07:30 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-15 07:35 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-15 07:40 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050825 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-15 07:56 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050862 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-15 07:59 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-15 08:04 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050879 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-15 08:08 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050910 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-15 08:17 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050811 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-15 08:17 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050694 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-15 08:22 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050914 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-15 08:48 | stmstation4@avive.life | W1192 | Agree | Leonora Cruz | 2401050866 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-15 08:51 | stmstation3@avive.life | W1192 | Agree | Leonora Cruz | 2401050930 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-15 09:34 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2401050909 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-15 09:37 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2312180371 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.356 | Pass |  | n/a
2024-02-15 09:37 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2312180270 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-15 09:43 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2312180398 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-15 09:44 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2312180268 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-15 09:44 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2312180399 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-15 09:49 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2312180265 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-15 09:53 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2312180355 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-15 09:54 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2312180348 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-15 09:58 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2312180262 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-15 10:02 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2312180356 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-15 10:41 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050771 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-15 11:00 | stmstation2@avive.life | W1192 | Agree | Leonora Cruz | 2401050438 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-15 11:02 | stmstation1@avive.life | W1192 | Agree | Leonora Cruz | 2312180394 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-15 11:03 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2312180220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-15 11:09 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050616 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-15 11:15 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050608 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-15 11:26 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050945 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-15 11:28 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050619 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-15 11:42 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050697 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-15 11:49 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-15 11:55 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2312180395 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-15 11:55 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2312180190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-15 12:00 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050725 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-15 12:05 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050741 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-15 12:10 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2312180215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-15 12:11 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050827 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-15 12:17 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2312180361 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-15 12:26 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2312180255 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-15 12:28 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050819 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-15 12:34 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050486 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-15 14:09 | stmstation2@avive.life | W1192 | Agree | Liza Ordillo | 2401050695 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-15 14:09 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050821 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-15 14:09 | stmstation4@avive.life | W1192 | Agree | Liza Ordillo | 2401050693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.332 | Pass |  | n/a
2024-02-15 14:19 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.355 | Pass |  | n/a
2024-02-15 14:23 | stmstation1@avive.life | W1192 | Agree | Liza Ordillo | 2401050937 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-15 14:39 | stmstation3@avive.life | W1192 | Agree | Liza Ordillo | 2401050834 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-16 15:02 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050831 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.347 | Pass |  | Reworked
2024-02-16 15:10 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050696 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.343 | Pass |  | n.a
2024-02-16 15:10 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401050815 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-16 15:10 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401050969 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-16 15:13 | stmstation1@avive.life | W1192 | Agree | Christina Cruz | 2401050446 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-16 15:19 | stmstation4@avive.life | W1192 | Agree | Christina Cruz | 2401050838 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-16 15:24 | stmstation3@avive.life | W1192 | Agree | Christina Cruz | 2401050865 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-16 16:04 | stmstation2@avive.life | W1192 | Agree | Christina Cruz | 2401050589 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | Fail Cellular & no sim
2024-02-20 06:58 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290376 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 07:01 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290366 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-20 07:11 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290297 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-20 07:17 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290508 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-20 07:20 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290354 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-20 07:31 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290367 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-20 07:32 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290344 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-20 07:34 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-20 07:38 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290353 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-20 07:38 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290348 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-20 07:46 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290383 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-20 07:47 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290390 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-20 07:48 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290292 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-20 07:50 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290340 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-20 08:03 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290339 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-20 08:05 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 08:06 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290295 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-20 08:07 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290355 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-20 08:23 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290342 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-20 08:23 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290287 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-20 08:25 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-20 08:25 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-20 08:42 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.388 | Pass |  | n/a
2024-02-20 08:43 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290288 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-20 08:46 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290200 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-20 08:50 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290436 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-20 08:58 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290327 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-20 08:59 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-20 09:04 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290338 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-20 09:04 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290165 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-20 09:14 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290161 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-20 09:15 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290431 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-20 09:20 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290197 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-20 09:22 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290433 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 09:27 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-20 09:27 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-20 09:30 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.331 | Pass |  | n/a
2024-02-20 09:36 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290139 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-20 09:37 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290434 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-20 09:42 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 09:46 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290381 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-20 09:49 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290430 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-20 09:51 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290432 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-20 09:56 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290456 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-20 10:03 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 10:39 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 10:40 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290388 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-20 10:40 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-20 10:41 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290352 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-20 10:50 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290346 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 10:54 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-20 11:00 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-20 11:00 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290350 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-20 11:02 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290389 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-20 11:11 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290384 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-20 11:11 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-20 11:27 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290382 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-20 11:28 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290494 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-20 11:38 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-20 11:42 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290162 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-20 13:45 | stmstation4@avive.life | W1214 | Agree | Liza Ordillo | 2401290216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-20 13:46 | stmstation3@avive.life | W1214 | Agree | Liza Ordillo | 2401290419 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-20 13:50 | stmstation1@avive.life | W1214 | Agree | Liza Ordillo | 2401290285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-20 13:53 | stmstation2@avive.life | W1214 | Agree | Liza Ordillo | 2401290225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.347 | Pass |  | N/A
2024-02-20 15:04 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290372 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-20 15:05 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290358 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-20 15:06 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290299 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-20 15:11 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290265 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 15:16 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290406 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.341 | Pass |  | N/A
2024-02-20 15:16 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290270 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-20 15:18 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290360 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.335 | Pass |  | N/a
2024-02-20 15:26 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290293 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.364 | Pass |  | N/A
2024-02-20 15:32 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290331 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-20 15:36 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290364 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 15:38 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290375 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 15:56 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290369 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-20 15:58 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290268 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-20 16:06 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290416 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-20 16:08 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290124 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-20 16:11 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290269 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-20 16:14 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290298 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-20 16:16 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290374 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-20 16:19 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290351 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-20 16:22 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290349 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-20 16:27 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290300 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-20 16:37 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-21 06:45 | stmstation3@avive.life | W1214 | Agree | Precious Ungco | 2401290378 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-21 06:46 | stmstation4@avive.life | W1214 | Agree | Precious Ungco | 2401290380 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 06:47 | stmstation1@avive.life | W1214 | Agree | Precious Ungco | 2401290371 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-21 06:50 | stmstation2@avive.life | W1214 | Agree | Precious Ungco | 2401290126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 07:03 | stmstation3@avive.life | W1214 | Agree | Precious Ungco | 2401290504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-21 07:04 | stmstation4@avive.life | W1214 | Agree | Precious Ungco | 2401290401 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-21 07:05 | stmstation1@avive.life | W1214 | Agree | Precious Ungco | 2401290296 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 07:06 | stmstation2@avive.life | W1214 | Agree | Precious Ungco | 2401290289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 07:16 | stmstation3@avive.life | W1214 | Agree | Precious Ungco | 2401290428 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-21 07:17 | stmstation4@avive.life | W1214 | Agree | Precious Ungco | 2401290043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 07:21 | stmstation1@avive.life | W1214 | Agree | Precious Ungco | 2401290370 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-21 07:22 | stmstation2@avive.life | W1214 | Agree | Precious Ungco | 2401290439 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 07:27 | stmstation3@avive.life | W1214 | Agree | Precious Ungco | 2401290435 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 07:29 | stmstation4@avive.life | W1214 | Agree | Precious Ungco | 2401290322 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-21 07:30 | stmstation1@avive.life | W1214 | Agree | Precious Ungco | 2401290294 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-21 07:31 | stmstation2@avive.life | W1214 | Agree | Precious Ungco | 2401290324 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-21 07:39 | stmstation3@avive.life | W1214 | Agree | Precious Ungco | 2401290341 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.332 | Pass |  | n/a
2024-02-21 07:40 | stmstation4@avive.life | W1214 | Agree | Precious Ungco | 2401290427 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-21 07:42 | stmstation1@avive.life | W1214 | Agree | Precious Ungco | 2401290131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.328 | Pass |  | n/a
2024-02-21 07:43 | stmstation2@avive.life | W1214 | Agree | Precious Ungco | 2401290429 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 07:51 | stmstation3@avive.life | W1214 | Agree | Precious Ungco | 2401290379 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-21 07:52 | stmstation4@avive.life | W1214 | Agree | Precious Ungco | 2401290345 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.331 | Pass |  | n/a
2024-02-21 07:55 | stmstation1@avive.life | W1214 | Agree | Precious Ungco | 2401290356 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-21 07:56 | stmstation2@avive.life | W1214 | Agree | Precious Ungco | 2401290368 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-21 08:02 | stmstation3@avive.life | W1214 | Agree | Precious Ungco | 2401290336 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-21 08:03 | stmstation4@avive.life | W1214 | Agree | Precious Ungco | 2401290335 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-21 08:05 | stmstation1@avive.life | W1214 | Agree | Precious Ungco | 2401290363 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-21 08:06 | stmstation2@avive.life | W1214 | Agree | Precious Ungco | 2401290418 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-21 08:14 | stmstation1@avive.life | W1214 | Agree | Precious Ungco | 2401290347 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-21 08:15 | stmstation2@avive.life | W1214 | Agree | Precious Ungco | 2401290507 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-21 08:16 | stmstation3@avive.life | W1214 | Agree | Precious Ungco | 2401290361 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-21 08:17 | stmstation4@avive.life | W1214 | Agree | Precious Ungco | 2401290359 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-21 08:23 | stmstation1@avive.life | W1214 | Agree | Precious Ungco | 2401290317 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 08:25 | stmstation2@avive.life | W1214 | Agree | Precious Ungco | 2401290458 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-21 08:42 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-21 08:44 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-21 08:46 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290305 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-21 08:49 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290159 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-21 08:50 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-21 08:50 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-21 09:02 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-21 09:03 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-21 09:09 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.276 | Pass |  | n/a
2024-02-21 09:15 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-21 09:16 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290082 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 09:20 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-21 09:23 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290077 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 09:26 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290257 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-21 09:30 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290233 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-21 09:34 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290247 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-21 09:36 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290160 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-21 09:42 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290061 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-21 09:43 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290227 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.355 | Pass |  | n/a
2024-02-21 09:45 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290260 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-21 09:52 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 09:53 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290279 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-21 09:54 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290261 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-21 09:58 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.331 | Pass |  | n/a
2024-02-21 10:00 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290066 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 10:04 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-21 10:19 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-21 10:21 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-21 10:29 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-21 10:31 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-21 10:33 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290377 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-21 10:37 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-21 10:39 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290311 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-21 10:44 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290228 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-21 10:46 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290263 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-21 10:49 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290266 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-21 10:49 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-21 10:52 | stmstation2@avive.life | W1214 | Agree | Christina Cruz | 2401290239 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-21 10:53 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-21 12:42 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290276 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-21 12:46 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290158 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-21 12:47 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.366 | Pass |  | n/a
2024-02-21 16:37 | stmstation1@avive.life | W1214 | Agree | Christina Cruz | 2401290240 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-21 16:40 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-21 16:41 | stmstation3@avive.life | W1214 | Agree | Christina Cruz | 2401290210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-22 09:09 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290195 | Pass | Fail | [10 empty] | No display, low current reading, .113a
2024-02-22 11:55 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290291 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-22 12:00 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290403 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-22 12:01 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290302 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-22 12:07 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-22 12:10 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290262 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-22 12:17 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290113 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-22 12:19 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-22 12:21 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290070 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-22 12:23 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290282 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-22 12:36 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290303 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-22 12:36 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-22 12:37 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290307 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-22 12:38 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-22 12:48 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-22 12:48 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290075 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-22 12:49 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-22 12:49 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290201 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.404 | Pass |  | n/a
2024-02-22 13:04 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-22 13:07 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290230 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-22 13:09 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290250 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-22 13:16 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290128 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-22 13:17 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-22 14:29 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290278 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-22 14:31 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 06:38 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-23 06:38 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290256 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-23 06:41 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290212 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 06:42 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 06:47 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-23 06:48 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 06:49 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290243 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 06:53 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290413 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 06:55 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290202 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-23 06:58 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290281 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-23 06:59 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290231 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-23 06:59 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290343 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-23 07:01 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290234 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-23 07:05 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290277 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-23 07:07 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 07:09 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-23 07:11 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 07:13 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 07:16 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-23 07:18 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290167 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-23 07:19 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290236 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-23 07:20 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290241 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-23 07:25 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-23 07:26 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 07:26 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290182 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-23 07:29 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-23 07:35 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 07:38 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-23 07:41 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290235 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 07:44 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 07:45 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 07:46 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290209 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 07:51 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 07:52 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290275 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 07:53 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 07:54 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290310 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 07:59 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 07:59 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290308 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 08:03 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-23 08:07 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 08:08 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 08:10 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 08:11 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 08:15 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-23 08:17 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-23 08:21 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-23 08:22 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290065 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 08:25 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290315 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.359 | Pass |  | n/a
2024-02-23 08:25 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290319 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-23 08:27 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 08:28 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290286 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-23 08:33 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290267 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 08:35 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-23 08:38 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290081 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-23 08:38 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-23 08:40 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-23 08:43 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 08:47 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-23 08:49 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 08:51 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290073 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-23 08:52 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290252 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 08:53 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290245 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-23 08:56 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290213 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 09:34 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290420 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-23 09:36 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2312180354 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 09:49 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.356 | Pass |  | n/a
2024-02-23 09:50 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 09:51 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-23 10:01 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2401290145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-23 10:01 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 10:12 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2312180375 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-23 10:12 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2312180256 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 10:13 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-23 10:14 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2312180359 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-23 10:24 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 10:25 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290232 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.382 | Pass |  | n/a
2024-02-23 10:25 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2312180254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 10:28 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290306 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 10:33 | stmstation2@avive.life | W1214 | Agree | Leonora Cruz | 2312180363 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.384 | Pass |  | n/a
2024-02-23 10:34 | stmstation1@avive.life | W1214 | Agree | Leonora Cruz | 2401290259 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 10:38 | stmstation3@avive.life | W1214 | Agree | Leonora Cruz | 2401290309 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 10:38 | stmstation4@avive.life | W1214 | Agree | Leonora Cruz | 2401290068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-23 10:46 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-23 10:47 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290301 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-23 10:51 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-23 10:53 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-23 10:54 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 10:57 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 10:59 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-23 11:04 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-23 11:05 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 11:09 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290063 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-23 11:10 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-23 11:11 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-23 11:15 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290191 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-23 11:17 | stmstation1@avive.life | W1214 | Agree | Valeria Manzo | 2401290255 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-23 11:21 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290248 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-23 11:23 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-23 11:29 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290357 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-23 11:32 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 11:40 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2401290080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-23 11:41 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290312 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-23 11:46 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2401290284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-23 11:48 | stmstation4@avive.life | W1214 | Agree | Valeria Manzo | 2312180259 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-23 11:49 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2312180251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-23 11:52 | stmstation3@avive.life | W1214 | Agree | Valeria Manzo | 2401290084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-23 11:56 | stmstation2@avive.life | W1214 | Agree | Valeria Manzo | 2312180360 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-23 12:23 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2401290373 | Pass | Fail | [10 empty] | Power on, Frozen on Avive screen
2024-02-23 13:18 | stmstation3@avive.life | W1214 | Agree | Ananya Jain | 2401290246 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Fail |  |  | No SIM Sign
2024-02-23 14:51 | stmstation4@avive.life | W1214 | Agree | Christina Cruz | 2312180258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-26 09:26 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401290326 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-26 09:31 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290332 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-26 09:31 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-26 09:33 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401050977 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-26 09:41 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401050980 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-26 09:41 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401050979 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-26 09:42 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-26 09:49 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-26 09:52 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-26 09:59 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401290387 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-26 09:59 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-26 10:45 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401050975 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-26 10:47 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401050964 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-26 10:48 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-26 10:49 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401290175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.329 | Pass |  | n/a
2024-02-26 10:57 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290405 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-26 10:57 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401290325 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-26 11:06 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290330 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-26 11:08 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-26 11:08 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290329 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-26 11:13 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-26 11:20 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-26 11:22 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290407 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-26 11:25 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290174 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-26 11:29 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290452 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-26 11:30 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-26 11:35 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290103 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-26 11:36 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-26 11:40 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290170 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-26 11:41 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290451 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-26 11:47 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290253 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-26 11:48 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290898 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.353 | Pass |  | n/a
2024-02-26 11:55 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290916 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-26 11:56 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290781 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-26 12:01 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290783 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.355 | Pass |  | n/a
2024-02-26 12:02 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290316 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-26 12:07 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290907 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-26 12:08 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290942 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-26 12:12 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290909 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-26 12:17 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290453 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-26 12:27 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-26 12:28 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-26 12:30 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290168 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-26 12:34 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-26 12:34 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290333 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-26 12:37 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290386 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-26 12:40 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-26 12:42 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-26 12:42 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-26 12:45 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-26 12:49 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-26 12:49 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290172 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-26 12:51 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290417 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-26 12:51 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-26 12:56 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-26 12:57 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290408 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.364 | Pass |  | n/a
2024-02-26 12:57 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290385 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.362 | Pass |  | n/a
2024-02-26 12:57 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.355 | Pass |  | n/a
2024-02-26 13:39 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290404 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-26 13:39 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-26 13:41 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290365 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-26 13:41 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-26 13:45 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290224 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-26 13:47 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-26 13:47 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290415 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-26 13:48 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401290219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-26 13:51 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-26 13:52 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290454 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-26 13:52 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290362 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-26 13:54 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401290402 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-26 14:05 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401290133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-26 14:06 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290459 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.356 | Pass |  | n/a
2024-02-26 14:06 | stmstation1@avive.life | W1217 | Agree | Precious Ungco | 2401290110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-26 14:07 | stmstation2@avive.life | W1217 | Agree | Precious Ungco | 2401290491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-26 14:20 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401290438 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-26 14:21 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-26 14:21 | stmstation1@avive.life | W1217 | Agree | Precious Ungco | 2401290492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-26 14:22 | stmstation2@avive.life | W1217 | Agree | Precious Ungco | 2401290460 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-26 14:30 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401290125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-26 14:30 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290440 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-26 14:31 | stmstation1@avive.life | W1217 | Agree | Precious Ungco | 2401290450 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-26 14:31 | stmstation2@avive.life | W1217 | Agree | Precious Ungco | 2401290881 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-26 14:43 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401290142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.36 | Pass |  | n/a
2024-02-26 14:43 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290138 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-26 14:44 | stmstation1@avive.life | W1217 | Agree | Precious Ungco | 2401290314 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-26 14:44 | stmstation2@avive.life | W1217 | Agree | Precious Ungco | 2401290313 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-26 14:54 | stmstation2@avive.life | W1217 | Agree | Precious Ungco | 2401290320 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-26 14:55 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401290135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-26 14:55 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290323 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-26 14:56 | stmstation1@avive.life | W1217 | Agree | Precious Ungco | 2401290111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-26 15:05 | stmstation2@avive.life | W1217 | Agree | Precious Ungco | 2401290151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-26 15:05 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401290107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-26 15:05 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290457 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-26 15:06 | stmstation1@avive.life | W1217 | Agree | Precious Ungco | 2401290148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-27 08:40 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401050983 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-27 08:42 | stmstation1@avive.life | W1217 | Agree | Precious Ungco | 2401290016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-27 08:45 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-27 08:45 | stmstation2@avive.life | W1217 | Agree | Precious Ungco | 2401290035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-27 08:52 | stmstation2@avive.life | W1217 | Agree | Precious Ungco | 2401050962 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-27 08:53 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290007 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-27 08:54 | stmstation1@avive.life | W1217 | Agree | Precious Ungco | 2401290220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-27 08:54 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401050963 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-27 09:14 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401050990 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.333 | Pass |  | n/a
2024-02-27 09:15 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290024 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-27 09:15 | stmstation1@avive.life | W1217 | Agree | Precious Ungco | 2401050996 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-27 09:16 | stmstation2@avive.life | W1217 | Agree | Precious Ungco | 2401290006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-27 09:23 | stmstation2@avive.life | W1217 | Agree | Precious Ungco | 2401290008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-27 09:24 | stmstation1@avive.life | W1217 | Agree | Precious Ungco | 2401290030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-27 09:24 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401290334 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-27 09:25 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401290221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-27 10:14 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401050994 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-27 10:17 | stmstation4@avive.life | W1217 | Agree | Precious Ungco | 2401050992 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-27 10:23 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290337 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.334 | Pass |  | n/a
2024-02-27 10:24 | stmstation3@avive.life | W1217 | Agree | Precious Ungco | 2401050995 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-27 10:29 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290328 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-27 10:36 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-27 10:36 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.358 | Pass |  | n/a
2024-02-27 10:39 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-27 10:47 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290321 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-27 10:47 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290933 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-27 10:48 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290437 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-27 10:57 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290144 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-27 11:01 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290455 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-27 11:03 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290929 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-27 11:07 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290318 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-27 11:13 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290913 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-27 11:14 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290166 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-27 11:14 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290934 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-27 11:16 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290926 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-27 11:29 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290935 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-27 11:30 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290928 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-27 11:30 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290938 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-27 11:31 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290931 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-27 11:43 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290901 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-27 11:44 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290883 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-27 11:45 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290930 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.359 | Pass |  | n/a
2024-02-27 11:45 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290937 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-27 11:50 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290892 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-27 11:52 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290936 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-27 11:56 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290891 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-27 11:58 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290882 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-27 12:00 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-27 12:02 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290924 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-27 12:03 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290950 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-27 12:08 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290932 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-27 12:09 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290943 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-27 12:09 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290906 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-27 12:12 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401050988 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-27 12:16 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290889 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-27 12:16 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290925 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-27 12:16 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290914 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.356 | Pass |  | n/a
2024-02-27 12:22 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290897 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-27 12:25 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290941 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-27 12:26 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290888 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-27 12:27 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290887 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-27 12:31 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290945 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.353 | Pass |  | n/a
2024-02-27 12:52 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-27 12:53 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401290412 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-27 12:53 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290478 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-27 13:08 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401050976 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-27 13:09 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290410 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-27 13:09 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401050999 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-27 13:17 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401050993 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-27 13:22 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-27 13:23 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-27 13:25 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401290010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-27 13:26 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401050984 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-27 13:36 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401050987 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-27 13:37 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290409 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-27 13:37 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401050989 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-27 13:41 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-27 13:48 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290649 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-27 13:48 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290414 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-27 13:48 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401051000 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-27 13:51 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401050968 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-27 14:05 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-28 07:18 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290486 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-28 07:18 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290884 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-28 07:22 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-28 07:22 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.332 | Pass |  | n/a
2024-02-28 07:27 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-28 07:28 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290469 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-28 07:31 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-28 07:32 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290890 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-28 07:35 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290886 | Pass | Fail | [10 empty] | "manufacturing default" on screen
2024-02-28 07:38 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.337 | Pass |  | n/a
2024-02-28 07:39 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-28 07:42 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-28 07:45 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290468 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-28 07:46 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290650 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-28 07:49 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-28 07:50 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-28 07:53 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-28 07:56 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290885 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.353 | Pass |  | n/a
2024-02-28 08:00 | stmstation2@avive.life | W1217 | Agree | Valeria Manzo | 2401290642 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-28 08:01 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-28 08:02 | stmstation4@avive.life | W1217 | Agree | Valeria Manzo | 2401290676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-28 08:04 | stmstation3@avive.life | W1217 | Agree | Valeria Manzo | 2401290484 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-28 08:06 | stmstation1@avive.life | W1217 | Agree | Valeria Manzo | 2401290680 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-28 10:52 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290621 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-28 10:54 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290949 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-28 10:57 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.339 | Pass |  | n/a
2024-02-28 10:57 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401290625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.342 | Pass |  | n/a
2024-02-28 11:13 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290618 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-28 11:22 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401290896 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-28 11:23 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290908 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-28 11:24 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290630 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-28 11:26 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-28 12:15 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-28 12:15 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-28 12:15 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290944 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-28 12:16 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401290706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-28 12:27 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-28 12:28 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290622 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-28 12:28 | stmstation2@avive.life | W1217 | Agree | Liza Ordillo | 2401290644 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-28 12:29 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-28 12:42 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290720 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-28 12:43 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290752 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-28 12:43 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.335 | Pass |  | n/a
2024-02-28 12:55 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290947 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-28 12:56 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290946 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-28 12:56 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-28 13:06 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-28 13:07 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-28 13:11 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-28 13:13 | stmstation4@avive.life | W1217 | Agree | Liza Ordillo | 2401290748 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-28 13:18 | stmstation1@avive.life | W1217 | Agree | Liza Ordillo | 2401290766 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-28 13:24 | stmstation3@avive.life | W1217 | Agree | Liza Ordillo | 2401290710 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-28 13:27 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290714 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-28 13:29 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290480 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.359 | Pass |  | n/a
2024-02-28 13:32 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-28 14:05 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290728 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.35 | Pass |  | n/a
2024-02-28 14:06 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290477 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.455 | Pass |  | n/a
2024-02-28 14:06 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-28 14:07 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290648 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-28 14:15 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.352 | Pass |  | n/a
2024-02-28 14:19 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-28 14:19 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401050986 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.378 | Pass |  | n/a
2024-02-28 14:32 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401050998 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-28 14:33 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.343 | Pass |  | n/a
2024-02-28 14:33 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290704 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-28 14:39 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401050985 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-29 09:26 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-29 09:26 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.347 | Pass |  | n/a
2024-02-29 09:27 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-29 09:32 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290465 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-29 09:39 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290721 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.365 | Pass |  | n/a
2024-02-29 09:39 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290751 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-29 09:41 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.351 | Pass |  | n/a
2024-02-29 09:42 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.341 | Pass |  | n/a
2024-02-29 09:57 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290679 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-29 09:59 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401050981 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.338 | Pass |  | n/a
2024-02-29 10:03 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290678 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-29 10:19 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-29 10:20 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290886 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-29 10:23 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.336 | Pass |  | n/a
2024-02-29 10:24 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.348 | Pass |  | n/a
2024-02-29 13:36 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2312180218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-29 13:37 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2401290705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.34 | Pass |  | n/a
2024-02-29 13:38 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290660 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.346 | Pass |  | n/a
2024-02-29 13:41 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2312180214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.344 | Pass |  | n/a
2024-02-29 13:50 | stmstation3@avive.life | W1217 | Agree | Leonora Cruz | 2401290488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-29 13:52 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2401290656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-02-29 13:52 | stmstation4@avive.life | W1217 | Agree | Leonora Cruz | 2312180213 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.355 | Pass |  | n/a
2024-02-29 13:55 | stmstation2@avive.life | W1217 | Agree | Leonora Cruz | 2401290675 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.349 | Pass |  | n/a
2024-02-29 14:02 | stmstation1@avive.life | W1217 | Agree | Leonora Cruz | 2312180210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.347 | Pass |  | n/a
2024-03-15 16:13 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | Test PCBA from Paramit MA
2024-03-15 16:15 | stmstation1@avive.life | WTest | Agree | Ananya Jain | 2402290027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.355 | Pass |  | Test Boards from Paramit MA
2024-03-15 16:21 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.359 | Pass |  | Test PCBA from Paramit MA
2024-03-15 16:29 | stmstation4@avive.life | WTest | Agree | Ananya jain | 2402290025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | Test PCBA from Paramit MA
2024-03-15 16:35 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.354 | Pass |  | Test PCBA from Paramit MA
2024-03-15 16:41 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | Test PCBA from Paramit MA
2024-03-15 16:48 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.357 | Pass |  | Test PCBA from Paramit MA
2024-03-15 16:53 | stmstation1@avive.life | WTest | Agree | Ananya JAin | 2402290028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.355 | Pass |  | Test PCBA from Paramit MA
2024-03-18 11:10 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.343 | Pass |  | Test PCBAs from Paramit MA
2024-03-18 11:18 | stmstation1@avive.life | WTest | Agree | Ananya Jain | 2402290021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.348 | Pass |  | Low= 0.179
2024-03-18 11:20 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.351 | Pass |  | Low=0.175
2024-03-18 11:28 | stmstation1@avive.life | WTest | Agree | Ananya Jain | 2402290051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.359 | Pass |  | Low=0.184
2024-03-18 11:31 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 358 | Pass | Pass | 0.348 | Pass |  | Low=0.175
2024-03-18 11:46 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.347 | Pass |  | Low=0.189
2024-03-18 11:47 | stmstation1@avive.life | WTest | Agree | Ananya Jain | 2402290022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 338 | Pass | Pass | 0.344 | Pass |  | Low= 0.177
2024-03-18 11:58 | stmstation1@avive.life | WTest | Agree | Ananya Jain | 2402290035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.353 | Pass |  | Test PCBAs from Paramit MA
2024-03-18 12:18 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.344 | Pass |  | Test PCBA from Paramit MA
2024-03-18 12:27 | stmstation1@avive.life | WTest | Agree | Ananya Jain | 2402290015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.346 | Pass |  | Test PCBAs from Paramit MA
2024-03-18 12:29 | stmstation4@avive.life | WTEst | Agree | Ananya Jain | 2402290055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.333 | Pass |  | Test PCBA from Paramit MA
2024-03-18 13:36 | stmstation4@avive.life | Wtest | Agree | Ananya Jain | 2402290053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.343 | Pass |  | Test PCBA from Paramit MA
2024-03-18 13:44 | stmstation4@avive.life | Wtest | Agree | Ananya Jain | 2402290017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.352 | Pass |  | Test PCBA from Paramit MA
2024-03-18 13:52 | stmstation4@avive.life | WTest | Agree | Ananya JAin | 2402290049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.336 | Pass |  | Test PCBA from Paramit MA
2024-03-18 13:58 | stmstation1@avive.life | WTest | Agree | Ananya jain | 2402290033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.357 | Pass |  | Test PCBAs from Paramit MA
2024-03-18 14:14 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.342 | Pass |  | Test PCBA from Paramit MA
2024-03-18 14:25 | stmstation4@avive.life | Wtest | Agree | Ananya Jain | 2402290008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.353 | Pass |  | Test PCBA from Paramit MA
2024-03-18 14:33 | stmstation4@avive.life | WTest | Agree | Ananya Jain | 2402290013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.365 | Pass |  | Test PCBA from Paramit MA
2024-04-01 15:25 | stmstation2@avive.life | W1240 | Agree | Christina Cruz | 2402130241 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-01 16:15 | stmstation1@avive.life | W1240 | Agree | Christina Cruz | 2402130203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-01 17:14 | stmstation3@avive.life | W1240 | Agree | Christina Cruz | 2402130039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-01 17:17 | stmstation4@avive.life | W1240 | Agree | Christina Cruz | 2402130004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.355 | Pass |  | n/a
2024-04-02 09:28 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2401290871 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-02 09:39 | stmstation3@avive.life | W1240 | Agree | Liza Ordillo | 2401290976 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-02 09:43 | stmstation4@avive.life | W1240 | Agree | Liza Ordillo | 2402130280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 09:45 | stmstation1@avive.life | W1240 | Agree | Christina Cruz | 2402130219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 09:46 | stmstation2@avive.life | W1240 | Agree | Christina Cruz | 2402130243 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-02 10:00 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130245 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-02 10:07 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 10:11 | stmstation4@avive.life | W1240 | Agree | Liza Ordillo | 2401290974 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 10:12 | stmstation3@avive.life | W1240 | Agree | Liza Ordillo | 2401290798 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 10:16 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 10:18 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130202 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-02 10:19 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-02 10:20 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-02 10:23 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2401290864 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-02 10:25 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 10:27 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.36 | Pass |  | n/a
2024-04-02 10:31 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130248 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-02 10:31 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2401290801 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-02 10:35 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2401290857 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-02 10:39 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2401290865 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-02 10:40 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2401290862 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 10:42 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-02 10:42 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2401290879 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-02 10:45 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130286 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.354 | Pass |  | n/a
2024-04-02 10:45 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-02 10:48 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-02 10:53 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-02 10:53 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2401290880 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-02 10:54 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130144 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 10:56 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-02 10:58 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130007 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 11:00 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-02 11:01 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2401290799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 11:05 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2401290797 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.353 | Pass |  | n/a
2024-04-02 11:06 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130283 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-02 11:06 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-02 11:09 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-02 11:13 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2401290874 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 11:17 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130292 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.356 | Pass |  | n/a
2024-04-02 11:18 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2401290802 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-02 11:18 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130246 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.34 | Pass |  | n/a
2024-04-02 11:18 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130291 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-02 11:22 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2401290808 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-02 11:24 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2401290807 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.356 | Pass |  | n/a
2024-04-02 11:25 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130295 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-02 11:27 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2401290806 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-02 11:28 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 11:31 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2401290813 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-02 11:32 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2401290804 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-02 11:35 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2401290805 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.356 | Pass |  | n/a
2024-04-02 11:38 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2401290800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.354 | Pass |  | n/a
2024-04-02 11:41 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2401290803 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 12:42 | stmstation1@avive.life | W1240 | Agree | Christina Cruz | 2402130114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.485 | Pass |  | n/a
2024-04-02 12:48 | stmstation2@avive.life | W1240 | Agree | Christina Cruz | 2402130184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-02 12:53 | stmstation3@avive.life | W1240 | Agree | Christina Cruz | 2402130147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 12:54 | stmstation4@avive.life | W1240 | Agree | Christina Cruz | 2402130115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-02 13:02 | stmstation1@avive.life | W1240 | Agree | Christina Cruz | 2402130187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 13:05 | stmstation2@avive.life | W1240 | Agree | Christina Cruz | 2401290866 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 13:18 | stmstation4@avive.life | W1240 | Agree | Christina Cruz | 2402130017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 13:21 | stmstation1@avive.life | W1240 | Agree | Christina Cruz | 2402130204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-02 13:27 | stmstation2@avive.life | W1240 | Agree | Christina Cruz | 2402130279 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.363 | Pass |  | n/a
2024-04-02 13:36 | stmstation3@avive.life | W1240 | Agree | Christina Cruz | 2401290861 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-02 13:43 | stmstation1@avive.life | W1240 | Agree | Christina Cruz | 2402130042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-02 13:43 | stmstation2@avive.life | W1240 | Agree | Christina Cruz | 2402130287 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 13:57 | stmstation3@avive.life | W1240 | Agree | Christina Cruz | 2402130284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-02 14:03 | stmstation1@avive.life | W1240 | Agree | Christina Cruz | 2402130281 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-02 14:21 | stmstation3@avive.life | W1240 | Agree | Christina Cruz | 2402130111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-02 14:23 | stmstation2@avive.life | W1240 | Agree | Christina Cruz | 2402130123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 14:31 | stmstation1@avive.life | W1240 | Agree | Christina Cruz | 2402130194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-02 14:43 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 14:44 | stmstation3@avive.life | W1240 | Agree | Liza Ordillo | 2402130131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 14:49 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-02 14:51 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130128 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-02 14:52 | stmstation3@avive.life | W1240 | Agree | Liza Ordillo | 2402130188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-02 15:00 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-02 15:01 | stmstation3@avive.life | W1240 | Agree | Liza Ordillo | 2402130112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-02 15:02 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-02 15:52 | stmstation4@avive.life | W1240 | Agree | Liza Ordillo | 2402130148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.35 | Pass |  | N/A
2024-04-02 17:09 | stmstation3@avive.life | W1240 | Agree | Christina Cruz | 2402130297 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-02 17:12 | stmstation2@avive.life | W1240 | Agree | Christina Cruz | 2402130109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 17:17 | stmstation3@avive.life | W1240 | Agree | Christina Cruz | 2401290868 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-02 17:17 | stmstation2@avive.life | W1240 | Agree | Christina Cruz | 2402130052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-02 17:17 | stmstation1@avive.life | W1240 | Agree | Christina Cruz | 2402130085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.357 | Pass |  | n/a
2024-04-02 17:24 | stmstation1@avive.life | W1240 | Agree | Christina Cruz | 2401290863 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-02 17:28 | stmstation2@avive.life | W1240 | Agree | Christina Cruz | 2402130125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.353 | Pass |  | n/a
2024-04-02 17:32 | stmstation3@avive.life | W1240 | Agree | Christina Cruz | 2402130192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-03 06:52 | stmstation1@avive.life | W1240 | Agree | Precious Ungco | 2402130010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.355 | Pass |  | n/a
2024-04-03 06:53 | stmstation2@avive.life | W1240 | Agree | Precious Ungco | 2402130217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.339 | Pass |  | n/a
2024-04-03 06:55 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-03 06:56 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-03 07:02 | stmstation1@avive.life | W1240 | Agree | Precious Ungco | 2402130018 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-03 07:03 | stmstation2@avive.life | W1240 | Agree | Precious Ungco | 2402130012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.339 | Pass |  | n/a
2024-04-03 07:05 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-03 07:05 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-03 07:12 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-03 07:13 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-03 07:13 | stmstation2@avive.life | W1240 | Agree | Precious Ungco | 2402130009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-03 07:15 | stmstation1@avive.life | W1240 | Agree | Precious Ungco | 2402130013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-03 07:22 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.363 | Pass |  | n/a
2024-04-03 07:23 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-03 07:23 | stmstation1@avive.life | W1240 | Agree | Precious Ungco | 2402130185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-03 07:24 | stmstation2@avive.life | W1240 | Agree | Precious Ungco | 2402130137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-03 07:32 | stmstation1@avive.life | W1240 | Agree | Precious Ungco | 2402130036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-03 07:33 | stmstation2@avive.life | W1240 | Agree | Precious Ungco | 2402130001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-03 07:33 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2401290875 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-03 07:34 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-03 07:44 | stmstation1@avive.life | W1240 | Agree | Precious Ungco | 2402130129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-03 07:45 | stmstation2@avive.life | W1240 | Agree | Precious Ungco | 2401290977 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-03 07:47 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2401290858 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-03 07:55 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-03 07:55 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130191 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-03 07:57 | stmstation1@avive.life | W1240 | Agree | Precious Ungco | 2402130136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-03 07:57 | stmstation2@avive.life | W1240 | Agree | Precious Ungco | 2402130126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-03 08:55 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-03 08:56 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130130 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-03 08:59 | stmstation1@avive.life | W1240 | Agree | Precious Ungco | 2402130031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-03 09:00 | stmstation2@avive.life | W1240 | Agree | Precious Ungco | 2402130200 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-03 09:05 | stmstation1@avive.life | W1240 | Agree | Precious Ungco | 2402130134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.359 | Pass |  | n/a
2024-04-03 09:06 | stmstation2@avive.life | W1240 | Agree | Precious Ungco | 2402130242 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-03 09:10 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-03 09:12 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-03 09:16 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-03 09:17 | christina@avive.life | W1240 | Agree | Liza Ordillo | 2401290878 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-03 09:17 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-03 09:18 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-03 09:27 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-03 09:27 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-03 09:29 | christina@avive.life | W1240 | Agree | Liza Ordillo | 2401290867 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-03 09:32 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-03 09:33 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-03 09:35 | christina@avive.life | W1240 | Agree | Liza Ordillo | 2401290872 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-03 09:36 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130113 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-03 09:39 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-03 09:44 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-03 09:44 | christina@avive.life | W1240 | Agree | Liza Ordillo | 2402130199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.366 | Pass |  | n/a
2024-04-03 09:45 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-03 09:53 | stmstation3@avive.life | W1240 | Agree | Precious Ungco | 2402130030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-03 09:53 | christina@avive.life | W1240 | Agree | Liza Ordillo | 2402130019 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.353 | Pass |  | n/a
2024-04-03 09:53 | stmstation4@avive.life | W1240 | Agree | Precious Ungco | 2402130195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-03 10:01 | christina@avive.life | W1240 | Agree | Liza Ordillo | 2402130198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-03 10:11 | stmstation3@avive.life | W1240 | Agree | Liza Ordillo | 2402130197 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-03 10:13 | stmstation4@avive.life | W1240 | Agree | Liza Ordillo | 2402130122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.353 | Pass |  | n/a
2024-04-03 10:17 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.353 | Pass |  | n/a
2024-04-03 10:25 | stmstation2@avive.life | W1240 | Agree | Liza Ordillo | 2402130138 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-03 10:29 | stmstation1@avive.life | W1240 | Agree | Liza Ordillo | 2402130290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-03 10:32 | stmstation4@avive.life | W1240 | Agree | Liza Ordillo | 2402130294 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-04 10:29 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2402130174 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 10:30 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2402130175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-04 10:30 | stmstation3@avive.life | W1241 | Agree | Liza Ordillo | 2401290818 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-04 10:34 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2402130238 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-04 10:36 | stmstation5@avive.life | W1241 | Agree | Liza Ordillo | 2402130046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-04 10:42 | stmstation3@avive.life | W1241 | Agree | Liza Ordillo | 2402130105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-04 10:42 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2401290815 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-04 10:43 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2402130043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-04 10:45 | stmstation5@avive.life | W1241 | Agree | Liza Ordillo | 2402130054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 10:47 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2402130100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-04 10:58 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2401290792 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-04 10:59 | stmstation3@avive.life | W1241 | Agree | Liza Ordillo | 2402130057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-04 11:04 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2401290837 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 11:06 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2402130120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-04 11:10 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2402130178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-04 11:10 | stmstation5@avive.life | W1241 | Agree | Liza Ordillo | 2401290817 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 11:15 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2401290831 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-04 11:16 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2402130116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-04 11:25 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2401290796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-04 11:27 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2402130101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 11:27 | stmstation3@avive.life | W1241 | Agree | Liza Ordillo | 2402130102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-04 11:28 | stmstation5@avive.life | W1241 | Agree | Liza Ordillo | 2402130139 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-04 11:33 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2402130107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-04 11:39 | stmstation3@avive.life | W1241 | Agree | Liza Ordillo | 2402130090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-04 11:39 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2402130099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-04 11:39 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2401290828 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-04 11:44 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2402130053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-04 11:47 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2402130225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-04 11:58 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2401290823 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.368 | Pass |  | n/a
2024-04-04 12:06 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2401290985 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-04 12:06 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2402130226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-04 12:07 | stmstation1@avive.life | W1241 | Agree | Leonora Cruz | 2401290824 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-04 12:07 | stmstation2@avive.life | W1241 | Agree | Leonora Cruz | 2401290832 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-04 12:08 | stmstation3@avive.life | W1241 | Agree | Leonora Cruz | 2402130056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.36 | Pass |  | n/a
2024-04-04 12:21 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2402130222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 12:26 | stmstation3@avive.life | W1241 | Agree | Leonora Cruz | 2402130098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-04 12:26 | stmstation2@avive.life | W1241 | Agree | Leonora Cruz | 2402130097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.544 | Pass |  | n/a
2024-04-04 12:27 | stmstation1@avive.life | W1241 | Agree | Leonora Cruz | 2402130104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-04 12:31 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2402130259 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.355 | Pass |  | n/a
2024-04-04 12:32 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2402130086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.355 | Pass |  | n/a
2024-04-04 12:39 | stmstation2@avive.life | W1241 | Agree | Leonora Cruz | 2402130229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-04 12:39 | stmstation1@avive.life | W1241 | Agree | Leonora Cruz | 2402130108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-04 13:20 | stmstation1@avive.life | W1241 | Agree | Leonora Cruz | 2401290958 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-04 13:24 | stmstation2@avive.life | W1241 | Agree | Leonora Cruz | 2401290951 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-04 13:28 | stmstation3@avive.life | W1241 | Agree | Leonora Cruz | 2402130021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.347 | Pass |  | N/A
2024-04-04 13:31 | stmstation1@avive.life | W1241 | Agree | Christina Cruz | 2401290952 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 13:39 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2402130022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 343 | Pass | Pass | 0.362 | Pass |  | n/a
2024-04-04 13:41 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2402130221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-04 13:49 | stmstation3@avive.life | W1241 | Agree | Leonora Cruz | 2402130260 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-04 13:55 | stmstation2@avive.life | W1241 | Agree | Christina Cruz | 2401290973 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.362 | Pass |  | n/a
2024-04-04 14:02 | stmstation1@avive.life | W1241 | Agree | Christina Cruz | 2402130224 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.353 | Pass |  | n/a
2024-04-04 14:03 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2402130250 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-04 14:05 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 14:07 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 14:13 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-04 14:13 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130201 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-04 14:13 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-04 15:00 | stmstation1@avive.life | W1241 | Agree | Christina Cruz | 2402130256 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-04 15:02 | stmstation2@avive.life | W1241 | Agree | Christina Cruz | 2401290829 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-04 15:04 | stmstation5@avive.life | W1241 | Agree | Precious Ungco | 2401290821 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-04 15:05 | stmstation4@avive.life | W1241 | Agree | Precious Ungco | 2401290860 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 15:06 | stmstation1@avive.life | W1241 | Agree | Christina Cruz | 2401290816 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-04 15:08 | stmstation2@avive.life | W1241 | Agree | Christina Cruz | 2401290811 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-04 15:10 | stmstation5@avive.life | W1241 | Agree | Precious Ungco | 2402130278 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-04 15:11 | stmstation3@avive.life | W1241 | Agree | Christina Cruz | 2401290979 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-04 15:11 | stmstation4@avive.life | W1241 | Agree | Precious Ungco | 2402130300 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.354 | Pass |  | n/a
2024-04-04 15:15 | stmstation1@avive.life | W1241 | Agree | Christina Cruz | 2401290825 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-04 15:16 | stmstation5@avive.life | W1241 | Agree | Precious Ungco | 2401290789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-04 15:17 | stmstation2@avive.life | W1241 | Agree | Christina Cruz | 2402130263 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 15:20 | stmstation3@avive.life | W1241 | Agree | Christina Cruz | 2402130252 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-04 15:23 | stmstation5@avive.life | W1241 | Agree | Precious Ungco | 2401290953 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-04 15:25 | stmstation4@avive.life | W1241 | Agree | Precious Ungco | 2401290900 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-04 15:27 | stmstation5@avive.life | W1241 | Agree | Precious Ungco | 2401290833 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-04 17:11 | stmstation3@avive.life | W1241 | Agree | Christina Cruz | 2401290814 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-04 17:12 | stmstation2@avive.life | W1241 | Agree | Christina Cruz | 2402130232 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-04 17:14 | stmstation1@avive.life | W1241 | Agree | Christina Cruz | 2402130228 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-05 06:05 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2401290954 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-05 06:08 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2401290834 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.356 | Pass |  | n/a
2024-04-05 06:08 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2401290835 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.355 | Pass |  | n/a
2024-04-05 06:12 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2401290787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-05 06:16 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2401290842 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-05 06:20 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130255 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 06:21 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.355 | Pass |  | n/a
2024-04-05 06:31 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-05 06:34 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2401290820 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-05 06:34 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130293 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-05 06:41 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2401290822 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-05 06:41 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130230 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-05 06:41 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2401290851 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 06:50 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-05 06:51 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2401290855 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.359 | Pass |  | n/a
2024-04-05 06:51 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2401290826 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-05 06:58 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2401290856 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 06:59 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130253 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-05 06:59 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2401290982 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-05 07:04 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130277 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-05 07:05 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130239 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 07:07 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2401290794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-05 07:11 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130231 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-05 07:11 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2401290785 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-05 07:15 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130236 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-05 07:18 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2401290795 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 07:19 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130296 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-05 07:22 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2401290790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 08:39 | stmstation1@avive.life | W1241 | Agree | Precious Ungco | 2402130235 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-05 08:40 | stmstation2@avive.life | W1241 | Agree | Precious Ungco | 2402130051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-05 08:41 | stmstation3@avive.life | W1241 | Agree | Precious Ungco | 2401290819 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 09:08 | stmstation1@avive.life | W1241 | Agree | Precious Ungco | 2402130247 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-05 09:09 | stmstation2@avive.life | W1241 | Agree | Precious Ungco | 2401290830 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-05 09:09 | stmstation4@avive.life | W1241 | Agree | Precious Ungco | 2402130249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-05 09:12 | stmstation3@avive.life | W1241 | Agree | Precious Ungco | 2401290975 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-05 09:45 | stmstation1@avive.life | W1241 | Agree | Precious Ungco | 2402130038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.339 | Pass |  | n/a
2024-04-05 09:45 | stmstation2@avive.life | W1241 | Agree | Precious Ungco | 2402130041 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.339 | Pass |  | n/a
2024-04-05 09:46 | stmstation3@avive.life | W1241 | Agree | Precious Ungco | 2402130258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 09:46 | stmstation5@avive.life | W1241 | Agree | Precious Ungco | 2402130233 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.339 | Pass |  | n/a
2024-04-05 09:46 | stmstation4@avive.life | W1241 | Agree | Precious Ungco | 2402130006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 09:55 | stmstation5@avive.life | W1241 | Agree | Precious Ungco | 2401290843 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-05 09:56 | stmstation4@avive.life | W1241 | Agree | Precious Ungco | 2401290844 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-05 09:56 | stmstation1@avive.life | W1241 | Agree | Precious Ungco | 2401290853 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-05 09:56 | stmstation2@avive.life | W1241 | Agree | Precious Ungco | 2401290836 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-05 09:57 | stmstation3@avive.life | W1241 | Agree | Precious Ungco | 2401290841 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-05 10:06 | stmstation1@avive.life | W1241 | Agree | Precious Ungco | 2401290838 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 10:06 | stmstation2@avive.life | W1241 | Agree | Precious Ungco | 2401290784 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-05 10:07 | stmstation3@avive.life | W1241 | Agree | Precious Ungco | 2401290870 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.339 | Pass |  | n/a
2024-04-05 10:07 | stmstation5@avive.life | W1241 | Agree | Precious Ungco | 2402130035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-05 10:08 | stmstation4@avive.life | W1241 | Agree | Precious Ungco | 2402130288 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 10:17 | stmstation1@avive.life | W1241 | Agree | Precious Ungco | 2402130251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 10:18 | stmstation2@avive.life | W1241 | Agree | Precious Ungco | 2402130234 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.356 | Pass |  | n/a
2024-04-05 10:18 | stmstation3@avive.life | W1241 | Agree | Precious Ungco | 2402130261 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 10:19 | stmstation5@avive.life | W1241 | Agree | Precious Ungco | 2402130282 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-05 10:19 | stmstation4@avive.life | W1241 | Agree | Precious Ungco | 2402130050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.339 | Pass |  | n/a
2024-04-05 10:36 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2401290846 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-05 10:36 | stmstation5@avive.life | W1241 | Agree | Liza Ordillo | 2401290967 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 10:39 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2401290810 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 10:40 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2401290847 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-05 10:42 | stmstation3@avive.life | W1241 | Agree | Liza Ordillo | 2401290981 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 10:47 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2402130266 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-05 10:48 | stmstation5@avive.life | W1241 | Agree | Liza Ordillo | 2401290827 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.354 | Pass |  | n/a
2024-04-05 10:53 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2401290812 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-05 10:54 | stmstation3@avive.life | W1241 | Agree | Liza Ordillo | 2401290854 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 10:54 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2401290966 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-05 11:01 | stmstation5@avive.life | W1241 | Agree | Liza Ordillo | 2402130087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-05 11:01 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2402130254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 11:45 | stmstation3@avive.life | W1241 | Agree | Liza Ordillo | 2402130257 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 11:46 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2401290962 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-05 11:46 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2402130091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 11:53 | stmstation5@avive.life | W1241 | Agree | Liza Ordillo | 2402130210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-05 11:53 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2401290957 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 11:59 | stmstation3@avive.life | W1241 | Agree | Liza Ordillo | 2401290956 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-05 12:00 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2401290986 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 12:04 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2401291000 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 12:05 | stmstation5@avive.life | W1241 | Agree | Liza Ordillo | 2401290939 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-05 12:10 | stmstation1@avive.life | W1241 | Agree | Liza Ordillo | 2401290852 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-05 12:12 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2402130274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.353 | Pass |  | n/a
2024-04-05 12:12 | stmstation5@avive.life | W1241 | Agree | Liza Ordillo | 2401290839 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.352 | Pass |  | n/a
2024-04-05 12:14 | stmstation2@avive.life | W1241 | Agree | Liza Ordillo | 2401290963 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-05 12:16 | stmstation4@avive.life | W1241 | Agree | Liza Ordillo | 2401290983 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 12:42 | stmstation3@avive.life | W1241 | Agree | Christina Cruz | 2401290849 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-05 12:43 | stmstation2@avive.life | W1241 | Agree | Christina Cruz | 2402130094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 12:44 | stmstation1@avive.life | W1241 | Agree | Christina Cruz | 2402130206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 12:55 | stmstation3@avive.life | W1241 | Agree | Christina Cruz | 2402130103 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 12:57 | stmstation2@avive.life | W1241 | Agree | Christina Cruz | 2402130106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 12:58 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130299 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-05 13:14 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130265 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-05 13:14 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-05 14:13 | stmstation5@avive.life | W1241 | Agree | Christina Cruz | 2402130181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-05 14:13 | stmstation4@avive.life | W1241 | Agree | Christina Cruz | 2401290968 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.353 | Pass |  | n/a
2024-04-05 14:15 | stmstation1@avive.life | W1241 | Agree | Christina Cruz | 2402130215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-05 14:19 | stmstation2@avive.life | W1241 | Agree | Christina Cruz | 2402130173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.353 | Pass |  | n/a
2024-04-05 14:20 | stmstation3@avive.life | W1241 | Agree | Christina Cruz | 2402130088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-05 14:28 | stmstation3@avive.life | W1241 | Agree | Christina Cruz | 2401290869 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-05 14:29 | stmstation2@avive.life | W1241 | Agree | Christina Cruz | 2402130044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.348 | Pass |  | n/a
2024-04-05 14:31 | stmstation5@avive.life | W1241 | Agree | Christina Cruz | 2401290955 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-05 14:31 | stmstation4@avive.life | W1241 | Agree | Christina Cruz | 2401290876 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-05 14:34 | stmstation1@avive.life | W1241 | Agree | Christina Cruz | 2402130180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-05 16:14 | stmstation3@avive.life | W1241 | Agree | Christina Cruz | 2401290972 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-08 06:45 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.34 | Pass |  | n/a
2024-04-08 06:45 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2401290793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-08 06:50 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130267 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-08 06:53 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.334 | Pass |  | n/a
2024-04-08 06:54 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2401290984 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-08 06:58 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2401290978 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-08 07:03 | stmstation4@avive.life | W1241 | Agree | Valeria Manzo | 2402130047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.336 | Pass |  | n/a
2024-04-08 07:03 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2401290877 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.338 | Pass |  | n/a
2024-04-08 07:03 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.341 | Pass |  | n/a
2024-04-08 07:08 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130276 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-08 07:13 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-08 07:13 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-08 07:18 | stmstation4@avive.life | W1241 | Agree | Valeria Manzo | 2401290964 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-08 07:19 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-08 07:23 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.338 | Pass |  | n/a
2024-04-08 07:23 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130182 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-08 07:28 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130268 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.337 | Pass |  | n/a
2024-04-08 07:29 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130269 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.358 | Pass |  | n/a
2024-04-08 07:30 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2401290996 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.353 | Pass |  | n/a
2024-04-08 07:38 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-08 07:38 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2401290997 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.36 | Pass |  | n/a
2024-04-08 07:38 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130240 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-08 07:39 | stmstation4@avive.life | W1241 | Agree | Valeria Manzo | 2402130211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-08 07:47 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-08 07:48 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130298 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-08 07:51 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130275 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-08 07:55 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2401290969 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-08 07:57 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130209 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-08 08:00 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2401290989 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.356 | Pass |  | n/a
2024-04-08 08:05 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.339 | Pass |  | n/a
2024-04-08 08:05 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2401290940 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.342 | Pass |  | n/a
2024-04-08 08:09 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2401290961 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-08 08:12 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130024 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-08 08:13 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2402130213 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-08 08:14 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2401290960 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-08 08:17 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.349 | Pass |  | n/a
2024-04-08 08:19 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2401290988 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.343 | Pass |  | n/a
2024-04-08 08:23 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.339 | Pass |  | n/a
2024-04-08 08:24 | stmstation3@avive.life | W1241 | Agree | Valeria Manzo | 2402130176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-08 08:25 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2401290809 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.412 | Pass |  | n/a
2024-04-08 08:27 | stmstation2@avive.life | W1241 | Agree | Valeria Manzo | 2402130208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-08 08:29 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2401290970 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.347 | Pass |  | n/a
2024-04-08 08:33 | stmstation1@avive.life | W1241 | Agree | Valeria Manzo | 2402130216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-08 08:34 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2401290994 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.355 | Pass |  | n/a
2024-04-08 08:44 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2401290965 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.36 | Pass |  | n/a
2024-04-08 08:45 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2401290971 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-08 09:16 | stmstation3@avive.life | W1241 | Agree | Leonora Cruz | 2401290999 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.351 | Pass |  | n/a
2024-04-08 09:16 | stmstation2@avive.life | W1241 | Agree | Leonora Cruz | 2402130212 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.361 | Pass |  | n/a
2024-04-08 09:19 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2401290786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.36 | Pass |  | n/a
2024-04-08 09:19 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2402130262 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.366 | Pass |  | n/a
2024-04-08 09:31 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2401290788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.361 | Pass |  | n/a
2024-04-08 09:31 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2401290992 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-08 09:34 | stmstation3@avive.life | W1241 | Agree | Leonora Cruz | 2401290848 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.35 | Pass |  | n/a
2024-04-08 09:38 | stmstation2@avive.life | W1241 | Agree | Leonora Cruz | 2401290998 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.377 | Pass |  | n/a
2024-04-08 09:44 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2402130270 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.345 | Pass |  | n/a
2024-04-08 09:48 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2401290991 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.379 | Pass |  | n/a
2024-04-08 09:52 | stmstation5@avive.life | W1241 | Agree | Leonora Cruz | 2401290995 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.344 | Pass |  | n/a
2024-04-08 09:55 | stmstation4@avive.life | W1241 | Agree | Leonora Cruz | 2401290987 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.355 | Pass |  | n/a
2024-04-08 09:59 | stmstation2@avive.life | W1241 | Agree | Leonora Cruz | 2401290791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.346 | Pass |  | n/a
2024-04-08 10:52 | stmstation1@avive.life | W1241 | Agree | Leonora Cruz | 2401290959 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [5 empty] | Can't connect to Cell