## STM Output
STM-03002 Rev P Electronic Data Collection form for Avive Connect Connect Sub-Assembly Release Tests and Inspection Data
Timestamp | Email Address | Work Order Number | General Instructions | Operator's Name | Serial Number | Serial Number Verification / Revision Verification | LCD Inspection | Speaker Test | GPS Satellite sign verification | Wifi Verification | Verify that Device is synced | Cellular icon appears and shows reception | Current (I) Value when wifi and cellular is connected | Verify the values are between 350 mA and 550 mA when wifi and cellular are connected | SIM sign verification | Current (I) value when device switched off (Low Power Mode) | Verify the values are between 0.150 mA and 0.200 mA | Verify the connect turns on. | Specify the Failure and Reason | Notes/Comments
2024-04-23 10:05 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020730 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-23 10:10 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020759 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 10:14 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 10:15 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020758 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-23 10:18 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 10:19 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 10:19 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 10:21 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020592 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-23 10:30 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020477 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 10:31 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020494 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-23 10:32 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020461 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 10:32 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-04-23 10:41 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020467 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 10:54 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020536 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-04-23 10:55 | stmstation2@avive.life | W1253 | Agree | Leonora Cruz | 2404020538 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 11:02 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020498 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-23 11:15 | stmstation4@avive.life | W1253 | Agree | Leonora Cruz | 2404020735 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 11:15 | stmstation6@avive.life | W1253 | Agree | Leonora Cruz | 2404020517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 11:23 | stmstation2@avive.life | W1253 | Agree | Leonora Cruz | 2404020531 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 11:23 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-23 11:31 | stmstation4@avive.life | W1253 | Agree | Leonora Cruz | 2404020693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 11:32 | stmstation6@avive.life | W1253 | Agree | Leonora Cruz | 2404020489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 11:46 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-04-23 11:47 | stmstation2@avive.life | W1253 | Agree | Leonora Cruz | 2404020724 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-23 11:51 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020454 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 12:01 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 12:02 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020437 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 12:05 | stmstation2@avive.life | W1253 | Agree | Leonora Cruz | 2404020772 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 12:14 | stmstation5@avive.life | W1253 | Agree | Leonora Cruz | 2404020695 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 12:15 | stmstation6@avive.life | W1253 | Agree | Leonora Cruz | 2404020486 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 12:20 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020675 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 12:21 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020718 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-23 12:28 | stmstation4@avive.life | W1253 | Agree | Leonora Cruz | 2404020430 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-23 12:34 | stmstation5@avive.life | W1253 | Agree | Leonora Cruz | 2404020556 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-23 12:37 | stmstation6@avive.life | W1253 | Agree | Leonora Cruz | 2404020747 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-23 13:03 | stmstation6@avive.life | W1253 | Agree | Precious Ungco | 2404020466 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-23 13:04 | stmstation5@avive.life | W1253 | Agree | Precious Ungco | 2404020577 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 422 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 13:04 | stmstation4@avive.life | W1253 | Agree | Precious Ungco | 2404020595 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 13:11 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020526 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 13:13 | stmstation6@avive.life | W1253 | Agree | Precious Ungco | 2404020462 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 13:14 | stmstation5@avive.life | W1253 | Agree | Precious Ungco | 2404020470 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 13:14 | stmstation4@avive.life | W1253 | Agree | Precious Ungco | 2404020642 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 13:24 | stmstation6@avive.life | W1253 | Agree | Precious Ungco | 2404020473 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 13:25 | stmstation5@avive.life | W1253 | Agree | Precious Ungco | 2404020634 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 13:26 | stmstation4@avive.life | W1253 | Agree | Precious Ungco | 2404020457 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 13:29 | stmstation6@avive.life | W1253 | Agree | Precious Ungco | 2404020532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-23 13:34 | stmstation4@avive.life | W1253 | Agree | Precious Ungco | 2404020610 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.165 | Pass | Pass |  | n/a
2024-04-23 13:39 | stmstation2@avive.life | W1253 | Agree | Leonora Cruz | 2404020428 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 13:39 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-23 13:49 | stmstation2@avive.life | W1253 | Agree | Leonora Cruz | 2404020771 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 13:49 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020681 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 13:50 | stmstation5@avive.life | W1253 | Agree | Leonora Cruz | 2404020722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 13:53 | stmstation4@avive.life | W1253 | Agree | Leonora Cruz | 2404020694 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.199 | Pass | Pass |  | n/a
2024-04-23 14:10 | stmstation6@avive.life | W1253 | Agree | Liza Ordillo | 2404020495 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 14:11 | stmstation5@avive.life | W1253 | Agree | Liza Ordillo | 2404020522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-23 14:12 | stmstation4@avive.life | W1253 | Agree | Liza Ordillo | 2404020670 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-23 14:19 | stmstation6@avive.life | W1253 | Agree | Liza Ordillo | 2404020491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 14:19 | stmstation5@avive.life | W1253 | Agree | Liza Ordillo | 2404020468 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-23 14:58 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020534 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.173 | Pass | Pass |  | N/A
2024-04-23 15:02 | stmstation2@avive.life | W1253 | Agree | Ananya Jain | 2404020792 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.174 | Pass | Pass |  | N/A
2024-04-23 15:04 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020763 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.175 | Pass | Pass |  | N/A
2024-04-23 15:09 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020777 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 15:15 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020755 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.175 | Pass | Pass |  | N/A
2024-04-23 15:25 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020784 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-23 15:32 | stmstation2@avive.life | W1253 | Agree | Ananya Jain | 2404020780 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 23.5 | Fail |  | High Current during Sleep | N/A
2024-04-23 15:36 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020783 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-23 15:43 | stmstation2@avive.life | W1253 | Agree | Ananya Jain | 2404020624 | Pass | Pass | Pass | Pass | Fail | [8 empty] | Can't connect to WiFi | N/A
2024-04-23 15:50 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020710 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-04-23 15:52 | stmstation2@avive.life | W1253 | Agree | Ananya Jain | 2404020628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | N/A
2024-04-23 15:57 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 15:58 | stmstation3@avive.life | W1253 | Agree | Ananya Jain | 2404020754 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 16:09 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020684 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 16:15 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 16:21 | stmstation2@avive.life | W1253 | Agree | Ananya Jain | 2404020458 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-23 16:26 | stmstation1@avive.life | W1253 | Agree | Ananya Jain | 2404020502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.174 | Pass | Pass |  | N/a
2024-04-23 16:40 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 16:57 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020778 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-23 17:00 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020561 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-23 17:03 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020596 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 17:05 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2404020515 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-23 17:06 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-23 17:07 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020301 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-23 17:12 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020606 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-23 17:15 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 05:51 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020303 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 05:51 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 05:51 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020530 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 05:51 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 05:52 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020519 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 05:53 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020331 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 06:07 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020302 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 06:07 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020587 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 06:09 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020512 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 06:11 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 06:13 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020511 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 06:25 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020471 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 06:25 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020520 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 06:26 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020322 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 06:27 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020507 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-24 06:30 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020518 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 06:35 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020509 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 06:36 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020623 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-24 06:37 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020463 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 06:39 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-24 06:42 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020464 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-24 06:44 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 06:45 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020469 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 06:47 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020500 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 07:31 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020624 | Pass | Pass | Pass | Pass | Fail | [8 empty] | no wifi connection | FAIL
2024-04-24 07:34 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020659 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-24 07:35 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 07:36 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020478 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 07:36 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 07:37 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020761 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 07:49 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020475 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 07:49 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 07:50 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020603 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-24 07:51 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 07:51 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020613 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 07:56 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020631 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-04-24 07:57 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020527 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 07:58 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020621 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-24 07:59 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020529 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 08:08 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020332 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 08:08 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 08:28 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020657 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 08:40 | stmstation3@avive.life | W1253 | Agree | Precious Ungco | 2404020423 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-24 08:40 | stmstation2@avive.life | W1253 | Agree | Precious Ungco | 2404020484 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 08:41 | stmstation1@avive.life | W1253 | Agree | Precious Ungco | 2404020476 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-24 08:42 | stmstation4@avive.life | W1253 | Agree | Precious Ungco | 2404020665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 08:43 | stmstation5@avive.life | W1253 | Agree | Precious Ungco | 2404020697 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 08:44 | stmstation6@avive.life | W1253 | Agree | Precious Ungco | 2404020626 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 08:57 | stmstation4@avive.life | W1253 | Agree | Precious Ungco | 2404020425 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-24 08:58 | stmstation6@avive.life | W1253 | Agree | Precious Ungco | 2404020717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 08:59 | stmstation5@avive.life | W1253 | Agree | Precious Ungco | 2404020692 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 08:59 | stmstation3@avive.life | W1253 | Agree | Precious Ungco | 2404020459 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 424 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 09:00 | stmstation2@avive.life | W1253 | Agree | Precious Ungco | 2404020794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 428 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 09:01 | stmstation1@avive.life | W1253 | Agree | Precious Ungco | 2404020749 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 09:34 | stmstation3@avive.life | W1253 | Agree | Precious Ungco | 2404020756 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-24 09:35 | stmstation2@avive.life | W1253 | Agree | Precious Ungco | 2404020305 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-04-24 09:35 | stmstation1@avive.life | W1253 | Agree | Precious Ungco | 2404020485 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 09:36 | stmstation4@avive.life | W1253 | Agree | Precious Ungco | 2404020768 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 09:36 | stmstation5@avive.life | W1253 | Agree | Precious Ungco | 2404020790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 09:37 | stmstation6@avive.life | W1253 | Agree | Precious Ungco | 2404020713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-04-24 09:46 | stmstation1@avive.life | W1253 | Agree | Precious Ungco | 2404020667 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 09:47 | stmstation2@avive.life | W1253 | Agree | Precious Ungco | 2404020487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 09:47 | stmstation3@avive.life | W1253 | Agree | Precious Ungco | 2404020702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 09:48 | stmstation6@avive.life | W1253 | Agree | Precious Ungco | 2404020728 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-04-24 09:48 | stmstation5@avive.life | W1253 | Agree | Precious Ungco | 2404020685 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 09:49 | stmstation4@avive.life | W1253 | Agree | Precious Ungco | 2404020753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 09:58 | stmstation6@avive.life | W1253 | Agree | Precious Ungco | 2404020504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 09:59 | stmstation5@avive.life | W1253 | Agree | Precious Ungco | 2404020682 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 09:59 | stmstation4@avive.life | W1253 | Agree | Precious Ungco | 2404020751 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 10:01 | stmstation3@avive.life | W1253 | Agree | Precious Ungco | 2404020689 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 10:01 | stmstation2@avive.life | W1253 | Agree | Precious Ungco | 2404020785 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-24 10:01 | stmstation1@avive.life | W1253 | Agree | Precious Ungco | 2404020719 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 10:15 | stmstation6@avive.life | W1253 | Agree | Liza Ordillo | 2404020516 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-04-24 10:15 | stmstation5@avive.life | W1253 | Agree | Liza Ordillo | 2404020648 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 10:15 | stmstation4@avive.life | W1253 | Agree | Liza Ordillo | 2404020793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-24 10:20 | stmstation1@avive.life | W1253 | Agree | Liza Ordillo | 2404020687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 10:24 | stmstation2@avive.life | W1253 | Agree | Liza Ordillo | 2404020543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 10:25 | stmstation3@avive.life | W1253 | Agree | Liza Ordillo | 2404020789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 10:34 | stmstation1@avive.life | W1253 | Agree | Liza Ordillo | 2404020449 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 10:35 | stmstation6@avive.life | W1253 | Agree | Liza Ordillo | 2404020567 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 10:35 | stmstation5@avive.life | W1253 | Agree | Liza Ordillo | 2404020563 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-24 10:35 | stmstation4@avive.life | W1253 | Agree | Liza Ordillo | 2404020497 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 10:39 | stmstation3@avive.life | W1253 | Agree | Liza Ordillo | 2404020564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 10:48 | stmstation6@avive.life | W1253 | Agree | Liza Ordillo | 2404020436 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 10:50 | stmstation5@avive.life | W1253 | Agree | Liza Ordillo | 2404020451 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 10:51 | stmstation4@avive.life | W1253 | Agree | Liza Ordillo | 2404020573 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-24 10:52 | stmstation1@avive.life | W1253 | Agree | Liza Ordillo | 2404020446 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 10:54 | stmstation3@avive.life | W1253 | Agree | Liza Ordillo | 2404020746 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 11:04 | stmstation6@avive.life | W1253 | Agree | Liza Ordillo | 2404020441 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 11:05 | stmstation5@avive.life | W1253 | Agree | Liza Ordillo | 2404020736 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-04-24 11:05 | stmstation4@avive.life | W1253 | Agree | Liza Ordillo | 2404020562 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 11:07 | stmstation1@avive.life | W1253 | Agree | Liza Ordillo | 2404020316 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 11:07 | stmstation2@avive.life | W1253 | Agree | Liza Ordillo | 2404020318 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 11:08 | stmstation3@avive.life | W1253 | Agree | Liza Ordillo | 2404020450 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 11:13 | stmstation6@avive.life | W1253 | Agree | Liza Ordillo | 2404020574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 11:14 | stmstation5@avive.life | W1253 | Agree | Liza Ordillo | 2404020306 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 11:14 | stmstation4@avive.life | W1253 | Agree | Liza Ordillo | 2404020545 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 182 | Pass | Pass |  | n/a
2024-04-24 11:21 | stmstation3@avive.life | W1253 | Agree | Liza Ordillo | 2404020321 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 11:22 | stmstation1@avive.life | W1253 | Agree | Liza Ordillo | 2404020701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-04-24 11:22 | stmstation2@avive.life | W1253 | Agree | Liza Ordillo | 2404020661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 11:24 | stmstation4@avive.life | W1253 | Agree | Liza Ordillo | 2404020304 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 11:24 | stmstation5@avive.life | W1253 | Agree | Liza Ordillo | 2404020324 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 11:41 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020508 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 11:42 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020616 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 11:54 | stmstation6@avive.life | W1253 | Agree | Leonora Cruz | 2404020651 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-04-24 11:54 | stmstation5@avive.life | W1253 | Agree | Leonora Cruz | 2404020608 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 11:55 | stmstation4@avive.life | W1253 | Agree | Leonora Cruz | 2404020765 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 12:03 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020775 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 12:04 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 12:12 | stmstation6@avive.life | W1253 | Agree | Leonora Cruz | 2404020422 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-24 12:13 | stmstation5@avive.life | W1253 | Agree | Leonora Cruz | 2404020652 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 12:19 | stmstation4@avive.life | W1253 | Agree | Leonora Cruz | 2404020560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 23.9 | Fail |  | failed value on shut down | n/a
2024-04-24 12:24 | stmstation2@avive.life | W1253 | Agree | Leonora Cruz | 2404020688 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 12:30 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020617 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 12:31 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020604 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 12:41 | stmstation2@avive.life | W1253 | Agree | Leonora Cruz | 2404020655 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 12:43 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020598 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 12:45 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 12:58 | stmstation5@avive.life | W1253 | Agree | Leonora Cruz | 2404020427 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 12:58 | stmstation6@avive.life | W1253 | Agree | Leonora Cruz | 2404020568 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 13:00 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020696 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 13:00 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020638 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 13:14 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020439 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 13:14 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020440 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 13:16 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020727 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 13:16 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-24 13:17 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020547 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-24 13:35 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020739 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-24 14:04 | stmstation6@avive.life | W1253 | Agree | Leonora Cruz | 2404020742 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 14:04 | stmstation4@avive.life | W1253 | Agree | Leonora Cruz | 2404020612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 14:11 | stmstation2@avive.life | W1253 | Agree | Leonora Cruz | 2404020609 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-24 14:14 | stmstation1@avive.life | W1253 | Agree | Leonora Cruz | 2404020630 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 15:12 | stmstation3@avive.life | W1253 | Agree | Leonora Cruz | 2404020566 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 15:15 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 15:20 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020589 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 15:44 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 15:54 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020752 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 15:56 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020533 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 16:01 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020426 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 16:11 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020762 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 16:12 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020569 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 16:18 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020764 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 16:20 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2404020644 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-24 16:23 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020663 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 16:26 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020546 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-24 16:28 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 16:31 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 16:36 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020544 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 16:44 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 16:57 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2404020455 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-24 17:03 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 17:07 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2404020620 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-24 17:09 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 17:11 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020448 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 17:15 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020619 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 17:16 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020654 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-24 17:31 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020779 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 17:33 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 17:34 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2404020576 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 17:41 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-24 17:43 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020433 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 17:44 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020424 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 17:48 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2404020479 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 17:48 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020311 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 17:49 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-24 17:53 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020766 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 17:56 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020760 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 17:58 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020474 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 18:02 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 24.45 | Fail |  | hi sleep reading | FAIL
2024-04-24 18:04 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020734 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-24 18:08 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020585 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-24 18:09 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020686 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-24 18:14 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020472 | Pass | Fail | [11 empty] | no display | fail
2024-04-24 18:16 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2404020580 | Pass | Fail | [11 empty] | No Display w/ Hi current reading | Fail
2024-04-24 18:22 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020499 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-24 18:29 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020743 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-25 06:55 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020309 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-25 06:55 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020444 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 06:55 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020434 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-25 06:56 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020435 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-25 06:56 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020738 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 06:57 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020453 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 07:20 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020578 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-25 07:21 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020307 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 07:21 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020443 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 07:30 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020442 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-25 07:30 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020571 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-25 07:39 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020605 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-25 07:44 | stmstation5@avive.life | W1253 | Agree | Valeria Manzo | 2404020317 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 07:44 | stmstation4@avive.life | W1253 | Agree | Valeria Manzo | 2404020757 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 07:46 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020622 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 07:47 | stmstation2@avive.life | W1253 | Agree | Valeria Manzo | 2404020548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 07:48 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-25 07:48 | stmstation6@avive.life | W1253 | Agree | Valeria Manzo | 2404020788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-25 08:48 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2404020432 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 08:51 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020542 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 09:10 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020445 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 09:26 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 09:29 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020570 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-25 10:00 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020447 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 10:04 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020590 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-25 10:05 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020460 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 10:06 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2404020748 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 10:12 | stmstation3@avive.life | W1253 | Agree | Liza Ordillo | 2404020639 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-25 10:12 | stmstation1@avive.life | W1253 | Agree | Liza Ordillo | 2404020741 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 10:18 | stmstation2@avive.life | W1253 | Agree | Liza Ordillo | 2404020594 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 10:23 | stmstation5@avive.life | W1253 | Agree | Liza Ordillo | 2404020314 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 10:32 | stmstation6@avive.life | W1253 | Agree | Liza Ordillo | 2404020456 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.433 | Fail |  | sleep mode hi current reading | Fail
2024-04-25 10:35 | stmstation3@avive.life | W1253 | Agree | Liza Ordillo | 2404020438 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-25 10:38 | stmstation5@avive.life | W1253 | Agree | Liza Ordillo | 2404020647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 10:39 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2402130169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-25 10:43 | stmstation4@avive.life | W1253 | Agree | Liza Ordillo | 2404020429 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 10:49 | stmstation1@avive.life | W1253 | Agree | Liza Ordillo | 2404020599 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 10:49 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2402130635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-25 10:50 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2402130613 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-25 10:50 | stmstation2@avive.life | W1253 | Agree | Liza Ordillo | 2404020553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-25 11:11 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2402130629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 423 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 11:16 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2402130545 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 11:19 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2402130612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 424 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-25 11:21 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2402130359 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 11:23 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2404020452 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 11:24 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2402130537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 11:26 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2402130611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 11:29 | stmstation6@avive.life | W1253 | Agree | Christina Cruz | 2402130501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-25 11:30 | stmstation5@avive.life | W1253 | Agree | Christina Cruz | 2404020572 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 11:32 | stmstation4@avive.life | W1253 | Agree | Christina Cruz | 2404020525 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 11:34 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2402130462 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-25 11:40 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2402130074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 11:43 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2402130385 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-25 11:46 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2402130636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-25 11:53 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2402130548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 11:55 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020575 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 12:08 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020580 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 12:08 | stmstation4@avive.life | W1253 | Agree | Ananya Jain | 2404020421 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | No Cell | N/A
2024-04-25 12:10 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2402130553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-25 12:12 | stmstation6@avive.life | W1253 | Agree | Ananya Jain | 2404020565 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 22.5 | Fail |  | High Current during sleep | n/a
2024-04-25 12:14 | stmstation6@avive.life | W1253 | Agree | Ananya Jain | 2404020514 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 44.2 | Fail |  | High Current during sleep | n/a
2024-04-25 12:15 | stmstation5@avive.life | W1253 | Agree | Ananya Jain | 2404020510 | Pass | Pass | Pass | Pass | Fail | [8 empty] | Can't connrct to WiFI | n/a
2024-04-25 12:17 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2404020705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-25 12:18 | stmstation3@avive.life | W1253 | Agree | Christina Cruz | 2404020472 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-25 12:20 | stmstation2@avive.life | W1253 | Agree | Christina Cruz | 2404020506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 12:20 | stmstation6@avive.life | W1253 | Agree | Ananya Jain | 2404020550 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.174 | Pass | Pass |  | N/A
2024-04-25 12:59 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020336 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 13:01 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-04-25 13:04 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020337 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 432 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-04-25 13:09 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 432 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 13:11 | stmstation4@avive.life | W1254 | Agree | Ananya Jain | 2404020209 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-25 13:23 | stmstation1@avive.life | W1253 | Agree | Christina Cruz | 2402130469 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | N/A
2024-04-25 13:24 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 13:24 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020297 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.187 | Pass | Pass |  | n/a
2024-04-25 13:31 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 432 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 14:16 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-25 14:18 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020330 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 14:22 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020365 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 14:25 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020326 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-25 14:28 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 14:32 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020228 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 14:38 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020277 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 14:53 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-25 14:55 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 15:06 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020075 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-04-25 15:07 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020236 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 15:12 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020295 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.193 | Pass | Pass |  | n/a
2024-04-25 15:13 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020287 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 15:19 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-25 15:20 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020325 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 15:22 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020233 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-25 15:27 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-25 15:30 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020411 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 15:36 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020246 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-04-25 15:39 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020404 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 15:41 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020329 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-25 15:44 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-25 15:48 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020281 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-25 15:49 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-04-25 15:54 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020328 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-25 16:01 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 06:13 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 06:14 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020081 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 06:14 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 06:15 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-26 06:15 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 06:20 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 06:27 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020082 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 502 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 06:27 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-26 06:28 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020128 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 06:30 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 06:30 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 06:32 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020061 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 06:36 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-26 06:36 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-26 06:42 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 06:43 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 06:43 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 06:44 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 505 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 06:44 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-26 06:44 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-26 06:57 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020077 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 06:57 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 06:57 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020113 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.192 | Pass | Pass |  | n/a
2024-04-26 06:58 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-26 06:58 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-26 07:17 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-26 07:18 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 07:18 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-26 07:19 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-26 07:19 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 07:20 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 07:29 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 07:29 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 07:30 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-26 07:30 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 07:30 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 442 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 07:34 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 07:40 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-26 07:45 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020041 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-26 07:46 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 07:46 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-26 07:54 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-26 07:56 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.175 | Pass | Fail | tested on all 6 stations, did not wake up with switch | n/a
2024-04-26 07:58 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020050 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | no cell over 1 hour | n/a
2024-04-26 09:02 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020234 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-26 09:03 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 09:05 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 09:06 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 09:07 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 09:10 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-26 10:34 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 10:34 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 10:35 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 10:35 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020070 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 10:35 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 10:36 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 11:15 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 11:17 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 11:24 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-26 11:30 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020065 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-26 11:31 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-26 11:34 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-26 11:36 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-26 11:37 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 11:38 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.177 | Pass | Pass |  | N/A
2024-04-26 11:47 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020227 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-26 11:48 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 11:51 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.177 | Pass | Pass |  | N/A
2024-04-26 11:52 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-26 11:54 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-04-26 11:57 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020232 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-26 11:59 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 11:59 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020018 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 12:03 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020296 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 12:04 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-26 12:06 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 12:09 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020279 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 12:11 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-26 12:11 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020224 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 12:12 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020288 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-26 12:14 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020408 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 12:37 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020343 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 12:38 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020066 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-26 12:38 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020335 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-26 12:44 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020270 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-26 12:48 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 12:49 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 13:00 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-26 13:00 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-26 13:01 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-26 13:18 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 13:21 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020073 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 13:23 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020374 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-26 13:26 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020294 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 509 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-26 13:29 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020172 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 07:04 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 07:04 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020276 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 07:08 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 07:13 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020299 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-04-29 07:14 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-29 07:15 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020265 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-29 07:19 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 07:20 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020174 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 07:29 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 07:31 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-04-29 07:37 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 438 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 07:40 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 07:44 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 07:54 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020319 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.198 | Pass | Pass |  | n/a
2024-04-29 07:57 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020138 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 438 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 08:05 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020158 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.19 | Pass | Pass |  | n/a
2024-04-29 08:06 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020292 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 08:09 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020379 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-04-29 08:11 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020200 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 442 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-29 08:13 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020256 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 08:15 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020293 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-29 08:15 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 185 | Pass | Pass |  | n/a
2024-04-29 08:27 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020023 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 08:31 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020245 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-29 08:33 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020261 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 08:38 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020197 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-29 08:40 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-29 08:46 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-29 08:46 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 08:47 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 438 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-04-29 08:52 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020298 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 08:52 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 08:54 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-29 08:58 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 431 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-29 08:59 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020275 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-29 09:01 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020252 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-29 09:12 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 09:13 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 11:51 | stmstation3@avive.life | W1254 | Agree | Precious Ungco | 2404020238 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-29 11:51 | stmstation2@avive.life | W1254 | Agree | Precious Ungco | 2404020373 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 11:52 | stmstation1@avive.life | W1254 | Agree | Precious Ungco | 2404020376 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-29 11:54 | stmstation6@avive.life | W1254 | Agree | Precious Ungco | 2404020237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 12:06 | stmstation3@avive.life | W1254 | Agree | Precious Ungco | 2404020378 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 12:07 | stmstation2@avive.life | W1254 | Agree | Precious Ungco | 2404020177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 12:08 | stmstation1@avive.life | W1254 | Agree | Precious Ungco | 2404020400 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 12:08 | stmstation4@avive.life | W1254 | Agree | Precious Ungco | 2404020380 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 12:09 | stmstation5@avive.life | W1254 | Agree | Precious Ungco | 2404020268 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 12:09 | stmstation6@avive.life | W1254 | Agree | Precious Ungco | 2404020179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-29 12:19 | stmstation6@avive.life | W1254 | Agree | Precious Ungco | 2404020327 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 12:20 | stmstation5@avive.life | W1254 | Agree | Precious Ungco | 2404020397 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 12:21 | stmstation4@avive.life | W1254 | Agree | Precious Ungco | 2404020333 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 12:21 | stmstation3@avive.life | W1254 | Agree | Precious Ungco | 2404020164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 12:22 | stmstation2@avive.life | W1254 | Agree | Precious Ungco | 2404020382 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 12:22 | stmstation1@avive.life | W1254 | Agree | Precious Ungco | 2404020395 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-29 12:33 | stmstation1@avive.life | W1254 | Agree | Precious Ungco | 2404020248 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 12:33 | stmstation2@avive.life | W1254 | Agree | Precious Ungco | 2404020368 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 12:34 | stmstation3@avive.life | W1254 | Agree | Precious Ungco | 2404020244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 12:35 | stmstation6@avive.life | W1254 | Agree | Precious Ungco | 2404020320 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 12:35 | stmstation5@avive.life | W1254 | Agree | Precious Ungco | 2404020240 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-29 12:36 | stmstation4@avive.life | W1254 | Agree | Precious Ungco | 2404020208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 12:47 | stmstation6@avive.life | W1254 | Agree | Precious Ungco | 2404020357 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 12:48 | stmstation5@avive.life | W1254 | Agree | Precious Ungco | 2404020272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 12:48 | stmstation3@avive.life | W1254 | Agree | Precious Ungco | 2404020285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 12:49 | stmstation4@avive.life | W1254 | Agree | Precious Ungco | 2404020340 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 12:49 | stmstation2@avive.life | W1254 | Agree | Precious Ungco | 2404020271 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 12:50 | stmstation1@avive.life | W1254 | Agree | Precious Ungco | 2404020260 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 13:00 | stmstation1@avive.life | W1254 | Agree | Precious Ungco | 2404020339 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-04-29 13:00 | stmstation2@avive.life | W1254 | Agree | Precious Ungco | 2404020283 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-29 13:01 | stmstation3@avive.life | W1254 | Agree | Precious Ungco | 2404020266 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 13:01 | stmstation4@avive.life | W1254 | Agree | Precious Ungco | 2404020170 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 13:02 | stmstation5@avive.life | W1254 | Agree | Precious Ungco | 2404020257 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 13:03 | stmstation6@avive.life | W1254 | Agree | Precious Ungco | 2404020273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-29 13:12 | stmstation3@avive.life | W1254 | Agree | Precious Ungco | 2404020396 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-04-29 13:12 | stmstation2@avive.life | W1254 | Agree | Precious Ungco | 2404020399 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 13:13 | stmstation1@avive.life | W1254 | Agree | Precious Ungco | 2404020269 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-29 13:13 | stmstation4@avive.life | W1254 | Agree | Precious Ungco | 2404020394 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 13:14 | stmstation5@avive.life | W1254 | Agree | Precious Ungco | 2404020387 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 13:15 | stmstation6@avive.life | W1254 | Agree | Precious Ungco | 2404020142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 13:29 | stmstation4@avive.life | W1254 | Agree | Precious Ungco | 2404020267 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 13:30 | stmstation5@avive.life | W1254 | Agree | Precious Ungco | 2404020173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-04-29 13:31 | stmstation3@avive.life | W1254 | Agree | Precious Ungco | 2404020119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 14:47 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 14:49 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020401 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 14:50 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-29 14:57 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 14:58 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020391 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-04-29 15:00 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020370 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 15:03 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 15:06 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 15:10 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-29 15:14 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 15:15 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 15:18 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020124 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 15:21 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 15:24 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 15:29 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-29 15:30 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020263 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-29 15:32 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 15:35 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020262 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-29 15:35 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020255 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 15:38 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 15:43 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-29 15:45 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-04-29 15:49 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-29 15:56 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-29 15:58 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020346 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-29 16:11 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020406 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 07:25 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 07:25 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020282 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 07:25 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 07:26 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020388 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 07:26 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 07:26 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020377 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-30 07:36 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020393 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-30 07:37 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020375 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-04-30 07:37 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020386 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-30 07:37 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020347 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-30 07:38 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 07:42 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020235 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-30 07:49 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 07:49 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020166 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 07:50 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020159 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 07:50 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020417 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-30 07:50 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020383 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-30 07:53 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020390 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 08:00 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020409 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 08:00 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020250 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-30 08:00 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020230 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 08:00 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020160 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-04-30 08:01 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020362 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 08:01 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020367 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 08:10 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020405 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 09:38 | stmstation4@avive.life | W1254 | Agree | Ananya Jain | 2404020040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-30 09:40 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-30 09:42 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020420 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-04-30 11:04 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020239 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | no cell, fail per ananya | n/a
2024-04-30 11:07 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-30 11:07 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-30 11:07 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 11:09 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020083 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-30 11:09 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-04-30 11:16 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020152 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-30 11:17 | stmstation4@avive.life | W1254 | Agree | Valeria Manzo | 2404020219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 11:18 | stmstation1@avive.life | W1254 | Agree | Valeria Manzo | 2404020165 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-30 11:20 | stmstation6@avive.life | W1254 | Agree | Valeria Manzo | 2404020214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 11:20 | stmstation2@avive.life | W1254 | Agree | Valeria Manzo | 2404020193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-30 11:21 | stmstation3@avive.life | W1254 | Agree | Valeria Manzo | 2404020162 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 11:24 | stmstation5@avive.life | W1254 | Agree | Valeria Manzo | 2404020215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 11:52 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 11:53 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 11:56 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-04-30 11:59 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 12:03 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 12:03 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020167 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-30 12:22 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020402 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-30 12:24 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020243 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 12:28 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020407 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 12:32 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020334 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 12:34 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-30 12:37 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020349 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 12:51 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020103 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-04-30 12:51 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-30 13:01 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020348 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-30 13:01 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 13:03 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020344 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 13:09 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020139 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 14:05 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 433 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 14:06 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 14:06 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020345 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 14:13 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 14:13 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-04-30 14:15 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-30 14:23 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020242 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 14:24 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020412 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-04-30 14:24 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 14:32 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-30 14:33 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020350 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 14:41 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-30 14:41 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020389 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-30 14:45 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 14:56 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-04-30 14:56 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-04-30 14:57 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020019 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 15:05 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-04-30 15:06 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 15:06 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020213 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 15:15 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 15:17 | stmstation5@avive.life | W1254 | Agree | Leonora Cruz | 2404020143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 15:20 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-04-30 15:26 | stmstation2@avive.life | W1254 | Agree | Leonora Cruz | 2404020253 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-04-30 15:27 | stmstation3@avive.life | W1254 | Agree | Leonora Cruz | 2404020216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-04-30 15:28 | stmstation1@avive.life | W1254 | Agree | Leonora Cruz | 2404020130 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-04-30 15:34 | stmstation4@avive.life | W1254 | Agree | Leonora Cruz | 2404020278 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-04-30 15:36 | stmstation6@avive.life | W1254 | Agree | Leonora Cruz | 2404020251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-01 06:22 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020419 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-01 06:25 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020342 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 06:26 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020361 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 06:38 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020247 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 06:40 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 06:42 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020366 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 06:47 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020354 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 06:58 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020416 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 07:04 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020392 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 07:06 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020300 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 07:11 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020353 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-01 07:29 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 07:33 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020359 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 07:36 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020212 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-01 07:38 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-01 07:41 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020201 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 07:44 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020356 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 07:45 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 07:48 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404080275 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 07:52 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-01 07:56 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020360 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 08:00 | stmstation1@avive.life | W1254 | Agree | Christina Cruz | 2404020221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.168 | Pass | Pass |  | n/a
2024-05-01 08:03 | stmstation3@avive.life | W1254 | Agree | Christina Cruz | 2404020169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 08:06 | stmstation2@avive.life | W1254 | Agree | Christina Cruz | 2404020196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 09:23 | stmstation6@avive.life | W1254 | Agree | Christina Cruz | 2404020371 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 09:25 | stmstation5@avive.life | W1254 | Agree | Christina Cruz | 2404020352 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | no cellular | Fail
2024-05-01 09:28 | stmstation4@avive.life | W1254 | Agree | Christina Cruz | 2404020024 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | No cellular icon | FAIL
2024-05-01 09:41 | stmstation1@avive.life | W1254 | Agree | Ananya Jain | 2404020005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 10:10 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 10:10 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080319 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 10:11 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 10:11 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 10:11 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080073 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-01 10:23 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.176 | Pass | Pass |  | N/A
2024-05-01 10:24 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080379 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-05-01 10:24 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080083 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-01 10:25 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 10:25 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 10:35 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080402 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-01 10:35 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-01 10:35 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 10:36 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 10:36 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-01 10:45 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 10:45 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 10:46 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 10:46 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080380 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 10:50 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 10:51 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080019 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-01 10:59 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 10:59 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080333 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-01 11:00 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080334 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-01 11:00 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 11:01 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080308 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-01 11:32 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080384 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 11:33 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 11:33 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080295 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 11:34 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080306 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 11:35 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080174 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.166 | Pass | Pass |  | n/a
2024-05-01 11:37 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080209 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 11:44 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080310 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-01 11:45 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-01 11:46 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 11:47 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 11:48 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080374 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 11:48 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 11:53 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080166 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 11:54 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080375 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-01 11:54 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080212 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-01 11:56 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 11:56 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 11:56 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 12:01 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080245 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 12:01 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-01 12:02 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080241 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 12:09 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080261 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 12:10 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 12:10 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080234 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-01 12:24 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080332 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 12:24 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080320 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 12:25 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080233 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-01 12:26 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-01 12:26 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080288 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 12:27 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080287 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 12:50 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-01 12:57 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 428 | Pass | Pass | 24.207 | Fail |  | tried 7 times still high after sleep | n/a
2024-05-01 12:59 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080061 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 24.088 | Fail |  | tried multiple times but still high | n/a
2024-05-01 13:16 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 13:17 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-01 13:17 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-01 13:18 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080312 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-01 13:19 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-01 13:19 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-01 13:32 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080335 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-01 13:33 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080328 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-01 13:33 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-01 13:33 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-01 13:33 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080330 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-01 13:34 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080336 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 10:07 | stmstation3@avive.life | W1257 | Agree | Christina Cruz | 2404080317 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-02 10:10 | stmstation2@avive.life | W1257 | Agree | Christina Cruz | 2404080325 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-02 10:12 | stmstation1@avive.life | W1257 | Agree | Christina Cruz | 2404080137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-02 10:18 | stmstation2@avive.life | W1257 | Agree | Christina Cruz | 2404080313 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-02 10:28 | stmstation3@avive.life | W1257 | Agree | Christina Cruz | 2404080145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 10:32 | stmstation2@avive.life | W1257 | Agree | Christina Cruz | 2404080324 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-02 10:33 | stmstation1@avive.life | W1257 | Agree | Christina Cruz | 2404080190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-02 10:37 | stmstation3@avive.life | W1257 | Agree | Christina Cruz | 2404080138 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 10:41 | stmstation2@avive.life | W1257 | Agree | Christina Cruz | 2404080039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-02 10:42 | stmstation1@avive.life | W1257 | Agree | Christina Cruz | 2404080044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 10:50 | stmstation2@avive.life | W1257 | Agree | Christina Cruz | 2404080314 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-02 10:51 | stmstation1@avive.life | W1257 | Agree | Christina Cruz | 2404080147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-02 10:53 | stmstation3@avive.life | W1257 | Agree | Christina Cruz | 2404080210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 10:56 | stmstation6@avive.life | W1257 | Agree | Christina Cruz | 2404080326 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-05-02 10:59 | stmstation5@avive.life | W1257 | Agree | Christina Cruz | 2404080264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 11:02 | stmstation2@avive.life | W1257 | Agree | Christina Cruz | 2404080296 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-02 11:04 | stmstation6@avive.life | W1257 | Agree | Christina Cruz | 2404080121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 11:06 | stmstation5@avive.life | W1257 | Agree | Christina Cruz | 2404080207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-02 11:08 | stmstation1@avive.life | W1257 | Agree | Christina Cruz | 2404080056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 11:10 | stmstation3@avive.life | W1257 | Agree | Christina Cruz | 2404080302 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 11:12 | stmstation2@avive.life | W1257 | Agree | Christina Cruz | 2404080043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 180 | Pass | Pass |  | n/a
2024-05-02 11:43 | stmstation6@avive.life | W1257 | Agree | Leonora Cruz | 2404080301 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 11:48 | stmstation4@avive.life | W1257 | Agree | Leonora Cruz | 2404080291 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 12:05 | stmstation1@avive.life | W1257 | Agree | Leonora Cruz | 2404080205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-05-02 12:07 | stmstation3@avive.life | W1257 | Agree | Leonora Cruz | 2404080040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 12:13 | stmstation2@avive.life | W1257 | Agree | Leonora Cruz | 2404080218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 12:15 | stmstation3@avive.life | W1257 | Agree | Leonora Cruz | 2404080290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-02 12:15 | stmstation1@avive.life | W1257 | Agree | Leonora Cruz | 2404080204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 12:23 | stmstation4@avive.life | W1257 | Agree | Leonora Cruz | 2404080341 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-02 12:23 | stmstation6@avive.life | W1257 | Agree | Leonora Cruz | 2404080037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 13:17 | stmstation6@avive.life | W1257 | Agree | Leonora Cruz | 2404080042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.192 | Pass | Pass |  | n/a
2024-05-02 13:22 | stmstation3@avive.life | W1257 | Agree | Leonora Cruz | 2404080045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 13:27 | stmstation1@avive.life | W1257 | Agree | Leonora Cruz | 2404080292 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-02 13:31 | stmstation6@avive.life | W1257 | Agree | Leonora Cruz | 2404080398 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 13:32 | stmstation5@avive.life | W1257 | Agree | Leonora Cruz | 2404080351 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-02 13:40 | stmstation6@avive.life | W1257 | Agree | Leonora Cruz | 2404080338 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-02 13:41 | stmstation5@avive.life | W1257 | Agree | Leonora Cruz | 2404080214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-02 13:44 | stmstation1@avive.life | W1257 | Agree | Leonora Cruz | 2404080034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-02 13:46 | stmstation3@avive.life | W1257 | Agree | Leonora Cruz | 2404080309 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-02 13:55 | stmstation2@avive.life | W1257 | Agree | Leonora Cruz | 2404080048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 14:14 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 14:14 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-02 14:15 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080152 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 14:15 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080305 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 14:16 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 14:25 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 14:25 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080247 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-02 14:26 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-02 14:27 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-02 14:27 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-05-02 14:28 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 14:37 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-02 14:38 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-02 14:39 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080182 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-02 14:39 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-02 14:40 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 14:40 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-02 14:50 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-02 14:51 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-02 14:51 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-02 14:52 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080230 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-02 14:52 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-02 14:52 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080162 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 14:57 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-02 14:58 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080200 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 14:59 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080197 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 14:59 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 14:59 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080252 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 14:59 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080024 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 15:03 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 15:06 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 15:06 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080103 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 15:06 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 15:07 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-02 15:07 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 15:10 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-02 15:13 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 15:14 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 15:14 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 15:15 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 15:15 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 15:19 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-02 15:19 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080113 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-02 15:22 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-02 15:23 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080281 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-02 15:23 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-02 15:24 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-02 15:26 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080276 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-02 15:27 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-02 15:29 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080283 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-02 15:43 | stmstation6@avive.life | W1257 | Agree | Ananya Jain | 2404080397 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 08:31 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 08:31 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-03 08:32 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 08:40 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 08:41 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080279 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-05-03 08:51 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 08:51 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 08:51 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 08:51 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 08:52 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 08:52 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 09:01 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-03 09:01 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 09:01 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 09:02 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 09:02 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 09:04 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-03 09:14 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 09:14 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 09:15 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 09:17 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 09:21 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080286 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-03 09:24 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080130 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 09:25 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 09:26 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080315 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 09:26 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 09:33 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080018 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-03 09:35 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 09:36 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-03 09:36 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 09:37 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-03 09:43 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080331 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 09:43 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 09:50 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 24.28 | Fail |  | to high when on sleep | n/a
2024-05-03 09:53 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080349 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 10:00 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080294 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 24.99 | Fail |  | to high when on sleep | n/a
2024-05-03 10:14 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080350 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 10:15 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-03 10:15 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 10:16 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 10:36 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-03 10:37 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080023 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-03 10:37 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080159 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 10:38 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 442 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 10:38 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 10:39 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080277 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-03 10:48 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-03 10:48 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080392 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 10:49 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080329 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 10:49 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080070 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 10:50 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 10:51 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 11:00 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 11:00 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-03 11:01 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 11:01 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080063 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-03 11:02 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-03 11:02 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 11:12 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080401 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-03 11:13 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080399 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 11:13 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080144 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 11:14 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080065 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 11:15 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-03 11:15 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 11:23 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 11:24 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080161 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 11:24 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080228 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 11:25 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 11:26 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080077 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-03 11:26 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080158 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 11:45 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 11:45 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080128 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 11:46 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080282 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 11:46 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 11:47 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080278 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 11:48 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080227 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 11:59 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 12:00 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080250 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 12:00 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080393 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 12:00 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-03 12:03 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-03 12:05 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404080143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 12:11 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080303 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 12:11 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-03 12:11 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 12:14 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080299 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 12:14 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080311 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-03 12:17 | stmstation3@avive.life | W1257 | Agree | Valeria Manzo | 2404020418 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 12:21 | stmstation5@avive.life | W1257 | Agree | Valeria Manzo | 2404080353 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 12:23 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 12:24 | stmstation4@avive.life | W1257 | Agree | Valeria Manzo | 2404080339 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 12:25 | stmstation1@avive.life | W1257 | Agree | Valeria Manzo | 2404080124 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 12:25 | stmstation2@avive.life | W1257 | Agree | Valeria Manzo | 2404080298 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 12:29 | stmstation6@avive.life | W1257 | Agree | Valeria Manzo | 2404080307 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 13:07 | stmstation2@avive.life | W1257 | Agree | Leonora Cruz | 2404080007 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 13:08 | stmstation3@avive.life | W1257 | Agree | Leonora Cruz | 2404080184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 13:14 | stmstation6@avive.life | W1257 | Agree | Leonora Cruz | 2404080181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 13:15 | stmstation5@avive.life | W1257 | Agree | Leonora Cruz | 2404080273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 13:15 | stmstation4@avive.life | W1257 | Agree | Leonora Cruz | 2404080154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 13:17 | stmstation1@avive.life | W1257 | Agree | Leonora Cruz | 2404080030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-03 13:26 | stmstation3@avive.life | W1257 | Agree | Leonora Cruz | 2404080055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 13:26 | stmstation2@avive.life | W1257 | Agree | Leonora Cruz | 2404080229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 13:27 | stmstation1@avive.life | W1257 | Agree | Leonora Cruz | 2404080318 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 13:33 | stmstation6@avive.life | W1257 | Agree | Leonora Cruz | 2404080125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-03 13:35 | stmstation5@avive.life | W1257 | Agree | Leonora Cruz | 2404080075 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 13:45 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 25.36 | Fail |  | high current during sleep mode | n/a
2024-05-03 13:56 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404020369 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 14:00 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-03 14:00 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 14:01 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 14:02 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 14:03 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080300 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 14:03 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 14:25 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 14:26 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 14:27 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080263 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 14:27 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080293 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 14:28 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404080316 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-03 14:28 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-03 14:39 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404020413 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 14:39 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404020185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 14:40 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404020308 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 14:41 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404080053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 14:41 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404080157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.187 | Pass | Pass |  | n/a
2024-05-03 14:42 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404080289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-03 14:56 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404020315 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-03 14:57 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404020258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 14:58 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404020264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-03 14:58 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404080304 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-03 14:59 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404020403 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 432 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 14:59 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404020182 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-03 15:13 | stmstation6@avive.life | W1257 | Agree | Precious Ungco | 2404020226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-03 15:14 | stmstation5@avive.life | W1257 | Agree | Precious Ungco | 2404020259 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-03 15:14 | stmstation4@avive.life | W1257 | Agree | Precious Ungco | 2404020310 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-03 15:14 | stmstation3@avive.life | W1257 | Agree | Precious Ungco | 2404020186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 15:14 | stmstation2@avive.life | W1257 | Agree | Precious Ungco | 2404020202 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-03 15:15 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404020203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-03 15:26 | stmstation1@avive.life | W1257 | Agree | Precious Ungco | 2404080386 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-06 11:55 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080743 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 11:55 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080456 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 11:57 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080744 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-06 11:57 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080711 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-06 11:57 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-06 12:00 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080573 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-06 12:01 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080562 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 12:03 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080563 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-05-06 12:08 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080751 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-06 12:08 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080511 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-06 12:08 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080752 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-06 12:09 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080741 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-06 12:09 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080738 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 12:10 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080746 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-05-06 12:18 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080587 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 12:18 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-06 12:19 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 12:21 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080575 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-06 12:22 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 12:27 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-06 12:27 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080526 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-06 12:28 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080571 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-06 12:31 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080660 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-06 12:31 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-06 12:32 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080750 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 12:47 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080469 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-06 12:56 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080470 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-06 12:57 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-06 13:05 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-06 13:05 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 13:06 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-06 13:12 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080480 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-06 13:13 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080494 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-06 13:20 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080459 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-06 13:21 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080498 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-06 13:21 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080568 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 514 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 13:29 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080565 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 13:29 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-06 13:38 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-06 13:39 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-06 13:40 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080465 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-06 13:41 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080569 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-05-06 13:42 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080472 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-06 13:54 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080468 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-06 13:54 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080570 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.192 | Pass | Pass |  | n/a
2024-05-06 14:01 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080593 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-06 14:02 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 505 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-06 14:02 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080733 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 14:24 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080871 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-06 14:24 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080652 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-06 14:25 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080657 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-06 14:26 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080649 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-06 14:26 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080748 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 14:27 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080742 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 506 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-06 14:37 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080650 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-06 14:37 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080872 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-06 14:38 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080654 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-06 14:38 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-06 14:39 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080856 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-06 14:39 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-06 14:49 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080736 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-06 14:49 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080567 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 437 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-06 14:50 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-06 14:50 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-06 14:51 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080653 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-06 14:52 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080747 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-06 15:01 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-06 15:01 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080561 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-06 15:01 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080514 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.186 | Pass | Pass |  | n/a
2024-05-06 15:01 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080642 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-06 15:01 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-06 15:05 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080869 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-06 15:12 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080419 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-07 09:56 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080450 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-07 09:57 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080454 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-07 09:57 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080473 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-07 09:58 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080737 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-05-07 09:59 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-07 09:59 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080745 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 510 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 10:09 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080462 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-07 10:09 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-07 10:10 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-07 10:11 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080614 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 519 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-07 10:13 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080495 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-07 10:13 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 10:28 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080516 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-07 10:29 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080592 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-07 10:30 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080590 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-07 10:30 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-07 10:31 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080644 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-07 10:31 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080581 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 10:44 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-07 10:45 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080485 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-07 10:45 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080474 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.168 | Pass | Pass |  | n/a
2024-05-07 10:46 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080580 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-07 10:47 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080497 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-07 10:50 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-07 11:01 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080453 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 11:02 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080457 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-07 11:02 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080464 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-07 11:03 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080471 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-07 11:03 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080418 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-07 11:08 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080420 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 24.74 | Fail |  | high current | n/a
2024-05-07 11:28 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080416 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-07 11:32 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-07 11:32 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080446 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-07 11:43 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080460 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 11:44 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080585 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.177 | Pass | Pass |  | N/A
2024-05-07 11:44 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080566 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-05-07 12:33 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-07 12:35 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080441 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-07 12:38 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080443 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-07 12:42 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080432 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-07 12:43 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080423 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 12:44 | stmstation4@avive.life | W1258 | Agree | Leonora Cruz | 2404080797 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-07 12:50 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080433 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-07 12:53 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 12:53 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080762 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-07 12:59 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-05-07 12:59 | stmstation4@avive.life | W1258 | Agree | Leonora Cruz | 2404080583 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-07 12:59 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-07 13:12 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080429 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-07 13:12 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080559 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 13:13 | stmstation4@avive.life | W1258 | Agree | Leonora Cruz | 2404080557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-07 13:14 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080466 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 13:14 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 432 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-07 13:15 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080431 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-07 13:40 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080421 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-07 13:40 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-07 13:40 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080603 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-07 13:42 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080417 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-07 13:43 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080425 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-07 13:43 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080442 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-07 13:49 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080509 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-07 13:49 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080424 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-07 13:50 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080437 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-07 13:51 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080589 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 13:51 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-07 13:51 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080819 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-07 13:59 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080838 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-07 13:59 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080602 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.187 | Pass | Pass |  | n/a
2024-05-07 13:59 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080576 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-07 14:01 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080829 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-07 14:02 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080435 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-07 14:02 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080830 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-05-07 14:12 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080448 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-07 14:12 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080542 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-07 14:12 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080434 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 14:15 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-05-07 14:15 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080577 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 14:19 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080445 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-07 14:20 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080601 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 14:20 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080538 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-05-07 14:21 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080617 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-07 14:22 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 14:22 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080444 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-07 14:25 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-07 14:42 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080440 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 14:44 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080415 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-07 14:45 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080449 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.169 | Pass | Pass |  | n/a
2024-05-07 14:47 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080467 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-07 14:47 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080519 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-05-07 14:51 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080428 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-07 14:53 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080521 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-07 14:55 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080422 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-07 14:59 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080638 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 23.423 | Fail |  | Hi reading during sleep | Fail
2024-05-07 15:00 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080556 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 22.456 | Fail |  | hi reading | Fail
2024-05-08 08:16 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080609 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 08:16 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080595 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-08 08:16 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-08 08:17 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080778 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-08 08:18 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080861 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-08 08:19 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080458 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 08:27 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080853 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 08:28 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080634 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-08 08:29 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080572 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-05-08 08:29 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 08:29 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080768 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 08:30 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080785 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 08:38 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080544 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 08:38 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-08 08:39 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080550 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-08 08:39 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080798 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-08 08:42 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080792 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 08:45 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080810 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 23.444 | Fail |  | value to high | n/a
2024-05-08 09:12 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080795 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-08 09:12 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080808 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-08 09:12 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080766 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-08 09:13 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-08 09:13 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-08 09:13 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 09:21 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080620 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 09:21 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080499 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 09:22 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-08 09:22 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080767 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 09:23 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080630 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 09:23 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080552 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-08 09:29 | stmstation4@avive.life | W1258 | Agree | Valeria Manzo | 2404080728 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-08 09:45 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080833 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-08 09:46 | stmstation4@avive.life | W1258 | Agree | Leonora Cruz | 2404080843 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 09:47 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080827 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.187 | Pass | Pass |  | n/a
2024-05-08 09:50 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080191 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 09:51 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080518 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-08 09:54 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080520 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-08 10:09 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-08 10:10 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 10:10 | stmstation4@avive.life | W1258 | Agree | Leonora Cruz | 2404080647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 10:13 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080632 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 10:14 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080659 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 10:16 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080631 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 10:26 | stmstation4@avive.life | W1258 | Agree | Leonora Cruz | 2404080828 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 10:28 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080826 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-05-08 10:28 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080824 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-08 10:39 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080655 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-08 10:40 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080646 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-08 10:40 | stmstation4@avive.life | W1258 | Agree | Leonora Cruz | 2404080866 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 10:47 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080825 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-08 10:49 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 10:49 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080536 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-08 10:56 | stmstation4@avive.life | W1258 | Agree | Leonora Cruz | 2404080187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-08 10:59 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-08 11:00 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 11:05 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080624 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 11:12 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 11:12 | stmstation4@avive.life | W1258 | Agree | Leonora Cruz | 2404080651 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 11:13 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080596 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-08 11:21 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 11:21 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080623 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 13:45 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080772 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-08 13:49 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-08 13:52 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2404080427 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 513 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-08 14:00 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080447 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 14:03 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-08 14:05 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080534 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 14:06 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2404080636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 14:08 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-08 14:12 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080551 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 14:14 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080438 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-08 14:16 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080844 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 14:24 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080618 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 14:26 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080841 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-08 14:27 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2404080839 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-08 14:28 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080508 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 14:30 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080823 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 14:39 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080531 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 14:41 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 14:42 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2404080193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 14:43 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-08 14:44 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 14:46 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080610 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-08 14:47 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 14:49 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080545 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-08 14:56 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080515 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-08 14:57 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080774 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-08 15:02 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2404080840 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-08 15:03 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-08 15:03 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080842 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-08 15:12 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080530 | Pass | Fail | [11 empty] | no display | Fail
2024-05-09 07:12 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080732 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-09 07:13 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-09 07:13 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080484 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 07:14 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080775 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 07:14 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 07:15 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080784 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 07:25 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080525 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 07:26 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-09 07:27 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080771 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 07:28 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080822 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 07:28 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080779 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-09 07:39 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 07:40 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080622 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 07:41 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080730 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 07:41 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080640 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 07:42 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080879 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 07:42 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 07:53 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080868 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 433 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 07:54 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080604 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 07:55 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080867 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 07:55 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-09 07:56 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 07:56 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080765 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 08:07 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080679 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 08:08 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080773 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 08:08 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080637 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-09 08:09 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080878 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 08:09 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 08:10 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080527 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 08:25 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080535 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 08:25 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080533 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 08:26 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080549 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 08:26 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080430 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-09 08:27 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080486 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 08:28 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 08:39 | stmstation4@avive.life | W1258 | Agree | Precious Ungco | 2404080870 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 08:39 | stmstation5@avive.life | W1258 | Agree | Precious Ungco | 2404080776 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 08:40 | stmstation6@avive.life | W1258 | Agree | Precious Ungco | 2404080706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 08:40 | stmstation3@avive.life | W1258 | Agree | Precious Ungco | 2404080855 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 08:40 | stmstation2@avive.life | W1258 | Agree | Precious Ungco | 2404080875 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 08:41 | stmstation1@avive.life | W1258 | Agree | Precious Ungco | 2404080619 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-09 09:21 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080887 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.187 | Pass | Pass |  | n/a
2024-05-09 09:25 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080477 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-09 09:25 | stmstation5@avive.life | W1258 | Agree | Valeria Manzo | 2404080764 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-09 09:25 | stmstation1@avive.life | W1258 | Agree | Valeria Manzo | 2404080888 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 09:28 | stmstation2@avive.life | W1258 | Agree | Valeria Manzo | 2404080761 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 09:29 | stmstation3@avive.life | W1258 | Agree | Valeria Manzo | 2404080512 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 09:38 | stmstation6@avive.life | W1258 | Agree | Valeria Manzo | 2404080461 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 09:57 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080615 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 10:03 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080892 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 10:09 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080835 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 431 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 10:12 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080479 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-09 10:13 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080836 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 10:20 | stmstation6@avive.life | W1258 | Agree | Leonora Cruz | 2404080621 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 10:21 | stmstation4@avive.life | W1258 | Agree | Leonora Cruz | 2404080734 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 10:22 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080599 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 10:31 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080598 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 10:31 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080639 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 10:45 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080680 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 10:46 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 10:46 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080865 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 10:57 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080883 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 10:58 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 11:00 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080886 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 11:07 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080854 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 11:08 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080769 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-09 11:10 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080725 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 11:21 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080864 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 11:22 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080881 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 11:25 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080877 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 11:36 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080507 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 22.67 | Fail |  | not within range | n/a
2024-05-09 11:44 | stmstation2@avive.life | W1258 | Agree | Leonora Cruz | 2404080605 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 11:44 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 11:47 | stmstation3@avive.life | W1258 | Agree | Leonora Cruz | 2404080626 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 11:57 | stmstation1@avive.life | W1258 | Agree | Leonora Cruz | 2404080885 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 12:02 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080510 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 24.06 | Fail |  | value reading high | n/a
2024-05-09 12:05 | stmstation5@avive.life | W1258 | Agree | Leonora Cruz | 2404080862 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | no cell/sim | n/a
2024-05-09 13:40 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080594 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 13:45 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080777 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 13:46 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-09 13:49 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080607 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 13:52 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080749 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 13:54 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2404080882 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 13:56 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080801 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 13:59 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080884 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 14:00 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080727 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 22.75 | Fail |  | hi reading | FAIL
2024-05-09 14:02 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080546 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 14:05 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 14:09 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080891 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 14:13 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 14:16 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080890 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 14:18 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080555 | Pass | Fail | [11 empty] | not working | n/a
2024-05-09 14:20 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2404080475 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 14:22 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-05-09 14:27 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080710 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 14:28 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080889 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 15:11 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-09 15:13 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 15:14 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080648 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 15:18 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 15:21 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080763 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 15:24 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080608 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-09 15:27 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080352 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 15:32 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080863 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 15:34 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080426 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 15:35 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-09 15:42 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 15:44 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080482 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 15:46 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2404080723 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-09 15:48 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080876 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 15:50 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080880 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-09 15:52 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2402130496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-09 16:00 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2402130588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-09 16:03 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2402130586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-09 16:04 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2402130634 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 16:05 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2402130560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 16:07 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2402130633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-09 16:15 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-10 07:06 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2402130541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-10 07:10 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080478 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-10 07:11 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080476 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-10 07:14 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2402130397 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-10 07:20 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2402130582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-10 07:21 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080436 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-10 07:24 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080832 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-10 07:26 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2402130389 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-10 07:28 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2402130570 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-10 07:34 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2402130555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-10 07:36 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2402130395 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.194 | Pass | Pass |  | n/a
2024-05-10 07:38 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080873 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-10 08:24 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-10 08:28 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2404080678 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-10 08:29 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-10 08:39 | stmstation2@avive.life | W1258 | Agree | Christina Cruz | 2404080451 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-10 08:39 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-10 08:43 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2402130408 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-10 08:43 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2404080735 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-10 08:47 | stmstation4@avive.life | W1258 | Agree | Christina Cruz | 2404080834 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-10 08:48 | stmstation5@avive.life | W1258 | Agree | Christina Cruz | 2402130150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-10 08:53 | stmstation3@avive.life | W1258 | Agree | Christina Cruz | 2404080837 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-10 08:56 | stmstation1@avive.life | W1258 | Agree | Christina Cruz | 2404080831 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-10 08:57 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080463 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-10 13:46 | stmstation6@avive.life | W1258 | Agree | Christina Cruz | 2404080800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Fail |  |  |  | No Cell No SIM | N.A
2024-05-10 13:51 | stmstation5@avive.life | W1258 | Agree | Ananya Jain | 2404080530 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.176 | Pass | Pass |  | Rework
2024-05-13 12:56 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.186 | Pass | Pass |  | n/a
2024-05-13 12:57 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404080661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 12:58 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 361 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 13:07 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-13 13:13 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404120004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-13 13:14 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404120002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-13 13:14 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404120059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-13 13:20 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.175 | Pass | Pass |  | N/A
2024-05-13 13:22 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-13 13:23 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-13 13:29 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404120181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-13 13:29 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404120127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-13 13:39 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404080358 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 13:39 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 13:40 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-13 13:50 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404120013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-13 13:51 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404120170 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-13 13:51 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404120123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-13 13:59 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404080690 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-13 13:59 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404080898 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-13 14:00 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-13 14:06 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404120202 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 14:07 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404120132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-13 14:07 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404120065 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-13 14:15 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-13 14:16 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-13 14:25 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404120031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-13 14:25 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404120073 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-13 14:26 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404120039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 14:30 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404020750 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-13 14:31 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404020579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-13 14:32 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080394 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-13 14:49 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080395 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 14:50 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404080400 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 423 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-13 14:51 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080396 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-05-13 14:53 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404120113 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-13 14:54 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404120014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-13 14:55 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404120088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-13 15:05 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404120120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 15:06 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404120174 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-13 15:06 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404120173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-13 15:07 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404120130 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 15:08 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404120128 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 15:08 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404120145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 442 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-13 15:19 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404080267 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 15:20 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404080163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-13 15:21 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404080165 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-13 15:22 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404120172 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 15:23 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404080170 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 423 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-13 15:24 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404120149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-13 15:36 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404080270 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-13 15:36 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404080272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-13 15:37 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404080254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-13 15:38 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080266 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-13 15:38 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404120169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.191 | Pass | Pass |  | n/a
2024-05-13 15:39 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080268 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-13 16:01 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080255 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-13 16:02 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404080347 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-13 16:02 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404120125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-13 16:03 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404120122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 432 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-13 16:04 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404120131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-13 16:05 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404120180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 07:02 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404020767 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 432 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-14 07:06 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 427 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 07:06 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404120165 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-14 07:10 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404020774 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 07:13 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404120090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 07:16 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404120102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 07:23 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404020600 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 07:24 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404120101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 07:25 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404120093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 07:32 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 07:33 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404120037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 07:34 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404120114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-14 07:41 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404020715 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-14 07:44 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404020583 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 07:46 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120124 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 07:51 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 07:52 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404120171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-14 07:52 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404120178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-14 07:57 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404020745 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-14 08:03 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404020723 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | NO Cell | FAIL
2024-05-14 08:30 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080388 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 08:31 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080371 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-14 08:31 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404080370 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-14 08:33 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404080348 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 08:33 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404080387 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 08:34 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404080269 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-14 08:43 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404080376 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-14 08:44 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404080378 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 08:44 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404080224 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 08:45 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 08:45 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404080246 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 08:46 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2402130618 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 08:57 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080382 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-14 08:57 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404080381 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 08:58 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080372 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-14 08:58 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404080368 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-14 08:59 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404080354 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 09:01 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404080373 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 09:34 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404120010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 09:35 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404020650 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 09:37 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404120177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-14 09:38 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.21 | Fail |  | value out of range | n/a
2024-05-14 09:44 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.172 | Pass | Fail | not turning on after power switch off | n/a
2024-05-14 10:09 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2402130627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 10:10 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2402130444 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 10:11 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2402130491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-14 10:17 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404020737 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-14 10:17 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404020664 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-14 10:18 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404020677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 10:23 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 10:26 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120144 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 10:26 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 10:27 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404020640 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 10:27 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404020668 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 10:28 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404020707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-14 10:38 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 10:39 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-14 10:40 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 10:47 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404120025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-14 10:48 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404080701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-14 10:49 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404120075 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-14 10:55 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 10:56 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404080817 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 10:57 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 11:05 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404120034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-14 11:06 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404120058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-14 11:06 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404120063 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 11:13 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-14 11:13 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 11:14 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 11:25 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404120027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-14 11:26 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404120182 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 11:31 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 11:32 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404080818 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 11:33 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 11:36 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404120028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 11:39 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404120067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-14 11:40 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404120135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 11:50 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 15:34 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-14 15:35 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404120009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 15:36 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404120150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-05-14 15:38 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404120082 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 15:38 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404080345 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 15:39 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404080806 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | N/A
2024-05-14 15:49 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404080667 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 15:50 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404080783 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 15:51 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-14 15:52 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120167 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-14 15:53 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404080691 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-14 15:58 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-14 16:02 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404120152 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-14 16:03 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404120080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-14 16:11 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404080704 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 23.75 | Fail |  | hi reading | FAIL
2024-05-15 07:54 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404080757 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 07:54 | stmstation5@avive.life | W1262 | Agree | Valeria Manzo | 2404020554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 07:55 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404020558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 07:56 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404020797 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-15 07:58 | stmstation1@avive.life | W1262 | Agree | Valeria Manzo | 2404020773 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 08:01 | stmstation4@avive.life | W1262 | Agree | Valeria Manzo | 2404020641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 08:06 | stmstation5@avive.life | W1262 | Agree | Valeria Manzo | 2404120091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-15 08:06 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404120068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 08:07 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404120005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 08:08 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404020721 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 08:13 | stmstation4@avive.life | W1262 | Agree | Valeria Manzo | 2404120020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-15 08:26 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404020798 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-15 08:26 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404020725 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 08:27 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404020680 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-15 08:28 | stmstation5@avive.life | W1262 | Agree | Valeria Manzo | 2404020714 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 08:28 | stmstation4@avive.life | W1262 | Agree | Valeria Manzo | 2404020557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-15 08:37 | stmstation4@avive.life | W1262 | Agree | Valeria Manzo | 2404020559 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 08:37 | stmstation5@avive.life | W1262 | Agree | Valeria Manzo | 2404120136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-15 08:38 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404020770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-15 08:38 | stmstation1@avive.life | W1262 | Agree | Valeria Manzo | 2404020700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-05-15 08:38 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404080804 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-15 08:39 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404020795 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-15 08:46 | stmstation1@avive.life | W1262 | Agree | Valeria Manzo | 2404120074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 08:46 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404020800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-15 08:46 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404080407 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 08:47 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404020671 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 08:49 | stmstation5@avive.life | W1262 | Agree | Valeria Manzo | 2404020769 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 08:53 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404020732 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.305 | Fail |  | current value to high | n/a
2024-05-15 08:55 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404080894 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 22.64 | Fail |  | value to high | n/a
2024-05-15 08:57 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404020799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 22.55 | Fail |  | current value to high | n/a
2024-05-15 09:33 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404080256 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 09:33 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404080666 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 09:34 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404080665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-15 09:35 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080716 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 09:35 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404080265 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 09:36 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080759 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 432 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-15 09:48 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404120018 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-15 09:48 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404020796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-15 09:49 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404020627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-15 09:50 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080406 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 09:50 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404120143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-15 09:51 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080720 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 10:00 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080405 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-15 10:01 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404120137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 10:01 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404120085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 10:02 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404080702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 10:02 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404080403 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 10:03 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404080675 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-15 10:14 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404120066 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 10:15 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404120056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 10:15 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404120011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 422 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 10:16 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080755 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 422 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-15 10:16 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404120133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 10:17 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404120064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 10:29 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404120071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 10:30 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404080756 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-15 10:30 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404080896 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-15 10:31 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080893 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 10:31 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404080782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-15 10:32 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404120053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 16:39 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404080169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 16:40 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404080377 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 16:41 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404080164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 16:52 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404080662 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-15 16:55 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404080168 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 16:56 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404080167 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-15 16:57 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404080674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-15 16:59 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404080389 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-15 17:01 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404080670 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 17:11 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404120168 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 17:13 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404080668 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-15 17:16 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404080171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 17:21 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404080895 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-15 17:23 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404120134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-15 17:26 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404120162 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-15 17:28 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404080803 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-15 17:30 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404080724 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-15 17:38 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404080712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 17:39 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404080369 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-15 17:41 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404080857 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-15 17:45 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404120006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-15 17:45 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404080900 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 17:45 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-15 17:47 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404080897 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 17:48 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404120019 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 17:49 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-15 17:58 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-15 18:05 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404080758 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 07:20 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404120077 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 07:21 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404120176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 07:21 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404120084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 07:22 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404120146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 07:23 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404120201 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-16 07:23 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404120204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-16 07:31 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080681 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 07:32 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404080781 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 07:32 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080715 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 429 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 07:33 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404080851 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 07:34 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404080807 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 07:34 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404080809 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 423 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 07:46 | stmstation6@avive.life | W1262 | Agree | Precious Ungco | 2404080820 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 07:46 | stmstation3@avive.life | W1262 | Agree | Precious Ungco | 2404120094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 07:47 | stmstation2@avive.life | W1262 | Agree | Precious Ungco | 2404080812 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 07:47 | stmstation1@avive.life | W1262 | Agree | Precious Ungco | 2404120083 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 07:51 | stmstation4@avive.life | W1262 | Agree | Precious Ungco | 2404080719 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 428 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 07:52 | stmstation5@avive.life | W1262 | Agree | Precious Ungco | 2404080683 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 22.27 | Fail |  | high current value | n/a
2024-05-16 07:59 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404080413 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 08:00 | stmstation1@avive.life | W1262 | Agree | Valeria Manzo | 2404080858 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-16 08:00 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404080717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 08:27 | stmstation4@avive.life | W1262 | Agree | Valeria Manzo | 2404080816 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 08:27 | stmstation5@avive.life | W1262 | Agree | Valeria Manzo | 2404080322 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 08:28 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404080860 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 08:37 | stmstation1@avive.life | W1262 | Agree | Valeria Manzo | 2404080899 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 08:38 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404020524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 08:38 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404080859 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 08:39 | stmstation5@avive.life | W1262 | Agree | Valeria Manzo | 2404080821 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 08:39 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404080852 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-16 08:42 | stmstation4@avive.life | W1262 | Agree | Valeria Manzo | 2404080688 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 08:50 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404080346 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-16 08:50 | stmstation1@avive.life | W1262 | Agree | Valeria Manzo | 2404080344 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 08:50 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404080323 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 08:52 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404080391 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 08:53 | stmstation4@avive.life | W1262 | Agree | Valeria Manzo | 2404080343 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 08:54 | stmstation5@avive.life | W1262 | Agree | Valeria Manzo | 2404020643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 09:00 | stmstation4@avive.life | W1262 | Agree | Valeria Manzo | 2404080689 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 09:01 | stmstation5@avive.life | W1262 | Agree | Valeria Manzo | 2404120076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 09:01 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404080687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-16 09:03 | stmstation1@avive.life | W1262 | Agree | Valeria Manzo | 2404020669 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 09:03 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404020666 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 09:04 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404020555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 09:10 | stmstation4@avive.life | W1262 | Agree | Valeria Manzo | 2404080847 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 09:10 | stmstation5@avive.life | W1262 | Agree | Valeria Manzo | 2404080356 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 09:11 | stmstation6@avive.life | W1262 | Agree | Valeria Manzo | 2404020674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-16 09:11 | stmstation1@avive.life | W1262 | Agree | Valeria Manzo | 2404020521 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-05-16 09:14 | stmstation2@avive.life | W1262 | Agree | Valeria Manzo | 2404020535 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 09:16 | stmstation3@avive.life | W1262 | Agree | Valeria Manzo | 2404080849 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-16 10:13 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404080385 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 10:13 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404080753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-16 10:15 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404080408 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 10:21 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404080703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 10:25 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404120070 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 22.11 | Fail |  | Value out of range after low power mode | n/a
2024-05-16 10:29 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404020678 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 10:35 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404020691 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-16 10:36 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-16 10:36 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120139 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 10:44 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404020708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 10:45 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404080780 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 10:45 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404120121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-16 10:56 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404080815 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 10:57 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404080760 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 11:04 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404080685 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 11:05 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404080682 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-16 11:06 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404020656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 11:12 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404080814 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-16 11:13 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.186 | Pass | Pass |  | n/a
2024-05-16 11:24 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404080850 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 11:25 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 11:26 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404080664 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 11:29 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404020699 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-16 11:39 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404020683 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 11:39 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404020645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-16 11:40 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404120095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 11:48 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404080714 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 11:50 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404080271 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 11:50 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 12:01 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404120258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 12:02 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404080686 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 12:04 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404080673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 12:05 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404120253 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-16 14:58 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 15:01 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404080845 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 15:04 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404080713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-16 15:06 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404080813 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 15:08 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404080684 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 15:11 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404080848 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 15:13 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-16 15:21 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404020541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 15:29 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404120161 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 15:30 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120257 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 15:34 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-16 15:37 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404080367 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 15:39 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404080722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 15:41 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404020552 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 15:45 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404020551 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 15:47 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-16 15:51 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404080404 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-16 15:54 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404020703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 15:57 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-16 16:00 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404120166 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-16 16:02 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-16 16:06 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404020646 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-16 16:06 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404120158 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 16:08 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404080410 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-16 16:32 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120079 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | no cell | FAIL
2024-05-16 16:49 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404020720 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 07:01 | stmstation1@avive.life | W1262 | Agree | Leonora Cruz | 2404020602 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-05-17 07:04 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404080672 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 07:04 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404120159 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 07:58 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404120210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 07:59 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404020134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-17 08:01 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404020381 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 08:03 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404120262 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-17 08:04 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404120256 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 08:06 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-17 08:14 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404020188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 08:15 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404020372 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-17 08:18 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404020184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 08:21 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404080383 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 08:22 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404020313 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 08:26 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404120017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 09:02 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404020183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 09:04 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404120229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-17 09:06 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 09:08 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404120110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-17 09:10 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120231 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-17 09:16 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404120016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-17 09:21 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404020410 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-17 09:23 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404020414 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 09:24 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404020187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-17 09:27 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404120089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 09:29 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404020363 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 09:31 | stmstation4@avive.life | W1262 | Agree | Christina Cruz | 2404120263 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-17 09:39 | stmstation6@avive.life | W1262 | Agree | Christina Cruz | 2404020690 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-17 09:41 | stmstation5@avive.life | W1262 | Agree | Christina Cruz | 2404120081 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 09:43 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404020706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-17 09:47 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404020615 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 09:59 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404020711 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-17 10:05 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120160 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-17 10:07 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404120021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-17 10:14 | stmstation3@avive.life | W1262 | Agree | Christina Cruz | 2404120184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-17 10:15 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120259 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-17 10:16 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-17 10:16 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404080697 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-17 10:17 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404120024 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 10:18 | stmstation2@avive.life | W1262 | Agree | Christina Cruz | 2404120206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-17 10:26 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120260 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-17 10:26 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404120157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-17 10:27 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404020698 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 10:31 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404020601 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-17 10:36 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404080411 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 10:36 | stmstation6@avive.life | W1262 | Agree | Leonora Cruz | 2404120155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-17 10:37 | stmstation1@avive.life | W1262 | Agree | Christina Cruz | 2404020358 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 10:40 | stmstation4@avive.life | W1262 | Agree | Leonora Cruz | 2404020355 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-17 10:45 | stmstation2@avive.life | W1262 | Agree | Leonora Cruz | 2404120142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-17 10:51 | stmstation5@avive.life | W1262 | Agree | Leonora Cruz | 2404120029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.167 | Pass | Fail | won't turn back on after power switch mode | n/a
2024-05-17 10:55 | stmstation3@avive.life | W1262 | Agree | Leonora Cruz | 2404020716 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-20 08:31 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-20 08:31 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120601 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 08:32 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-20 08:32 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120593 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 08:32 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120704 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-20 08:33 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120589 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-20 08:57 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120644 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-20 08:58 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120594 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-20 08:58 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 08:59 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120583 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-20 08:59 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-20 09:11 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120570 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-20 09:11 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120587 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 09:13 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120604 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 09:13 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120576 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 09:13 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120590 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-20 09:15 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120598 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 09:27 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120556 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-20 09:28 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120606 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-20 09:28 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120478 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-20 09:29 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 09:33 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120660 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 09:34 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 09:37 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120575 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 23.83 | Fail |  | value to high when on sleep | n/a
2024-05-20 09:42 | stmstation4@avive.life | W1262 | Agree | Ananya Jain | 2404120055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 22.85 | Fail |  | High Current During Sleep | N/A
2024-05-20 09:50 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120695 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 24.16 | Fail |  | value to high when on sleep | n/a
2024-05-20 10:02 | stmstation4@avive.life | W1262 | Agree | Ananya Jain | 2404080409 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 430 | Pass | Pass | 0.175 | Pass | Pass |  | N/A
2024-05-20 10:52 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120666 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-20 10:53 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120650 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 11:02 | stmstation4@avive.life | W1266 | Agree | Leonora Cruz | 2404120658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-20 11:03 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-20 11:08 | stmstation5@avive.life | W1266 | Agree | Leonora Cruz | 2404120653 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 11:10 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120745 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 11:11 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120742 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 11:13 | stmstation4@avive.life | W1262 | Agree | Ananya Jain | 2404120103 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | No Cell | N/A
2024-05-20 11:16 | stmstation5@avive.life | W1266 | Agree | Leonora Cruz | 2404120685 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-20 11:24 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120686 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 11:25 | stmstation5@avive.life | W1266 | Agree | Leonora Cruz | 2404120661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 11:26 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120662 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-20 11:27 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120670 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 11:40 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120680 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 11:41 | stmstation4@avive.life | W1266 | Agree | Leonora Cruz | 2404120681 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 11:47 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120663 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 11:47 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 11:48 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120675 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-20 11:57 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-20 11:58 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120668 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-20 12:02 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120744 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.151 | Fail |  | Value after power down out of range | n/a
2024-05-20 12:07 | stmstation4@avive.life | W1266 | Agree | Leonora Cruz | 2404120654 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 12:10 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120595 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-20 12:10 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120669 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-20 14:54 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 427 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 14:58 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120679 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 427 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 15:04 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 15:08 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120739 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-20 15:11 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120615 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-20 15:12 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120741 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-20 15:21 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120699 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-20 15:23 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120497 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 23.84 | Fail |  | hi reading | fail
2024-05-20 15:31 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120698 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-20 15:34 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120585 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-20 15:35 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-20 15:40 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120746 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-20 15:44 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120607 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-20 15:46 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120683 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-20 17:11 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120682 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-21 08:09 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.168 | Pass | Pass |  | n/a
2024-05-21 08:09 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120657 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 08:09 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120405 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 08:11 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-21 08:11 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120410 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 08:12 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120566 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 08:23 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120387 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-21 08:23 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120357 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 08:23 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120378 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 08:25 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120485 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 08:25 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120347 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-21 08:26 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120344 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 08:32 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120360 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-21 08:34 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120384 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-21 08:35 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 08:35 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120392 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 08:36 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120401 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 08:36 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120561 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 08:43 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120397 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 08:44 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120465 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 08:45 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120391 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-21 08:45 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120358 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-21 08:47 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120359 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 08:48 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120385 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-21 08:53 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120362 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 08:53 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120409 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 08:53 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120398 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 08:54 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120386 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-21 08:59 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120363 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 22.2 | Fail |  | value to high when on sleep | n/a
2024-05-21 09:28 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120667 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | no cell but has sim | n/a
2024-05-21 09:30 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120396 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | no cell no sim, 30+ min | n/a
2024-05-21 10:34 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-21 10:35 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120512 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-21 10:36 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 10:38 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120577 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 10:40 | stmstation5@avive.life | W1266 | Agree | Leonora Cruz | 2404120492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 10:40 | stmstation4@avive.life | W1266 | Agree | Leonora Cruz | 2404120636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 10:54 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120610 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 10:54 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120317 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 11:00 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 11:01 | stmstation5@avive.life | W1266 | Agree | Leonora Cruz | 2404120468 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-21 11:01 | stmstation4@avive.life | W1266 | Agree | Leonora Cruz | 2404120560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 11:12 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120632 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-21 11:13 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120571 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-21 11:15 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120631 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 317 | Fail |  |  |  |  | value not in range | n/a
2024-05-21 11:18 | stmstation5@avive.life | W1266 | Agree | Leonora Cruz | 2404120608 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 11:19 | stmstation4@avive.life | W1266 | Agree | Leonora Cruz | 2404120634 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-21 11:24 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120652 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 11:26 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120581 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 11:31 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120464 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 11:32 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120449 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 11:34 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120565 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 11:47 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120547 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 11:47 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-21 11:48 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120424 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-21 11:57 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-21 11:58 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120696 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 353 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 11:59 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120649 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 12:00 | stmstation5@avive.life | W1266 | Agree | Leonora Cruz | 2404120446 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 12:01 | stmstation4@avive.life | W1266 | Agree | Leonora Cruz | 2404120545 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 12:01 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120549 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 12:29 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120432 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 12:29 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120538 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.186 | Pass | Pass |  | n/a
2024-05-21 12:29 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-21 12:30 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120498 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 12:30 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120429 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 12:30 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120435 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 12:43 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120235 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 12:44 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120361 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 12:45 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120232 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 12:45 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120617 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 12:49 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120438 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 12:51 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120383 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 13:01 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120616 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 13:01 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120371 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 13:02 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120365 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-21 13:03 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120406 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 13:03 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120500 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 13:25 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 13:25 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120389 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 13:26 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120366 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-21 13:26 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120390 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-21 13:26 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-21 13:27 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120364 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 13:40 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120395 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 13:41 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120436 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 13:45 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120499 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 13:46 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-21 13:46 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120346 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 13:46 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120394 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 13:54 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 17:19 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120530 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 422 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 17:20 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120520 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-21 17:23 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 17:26 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-21 17:29 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120509 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 428 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 17:34 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120519 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 17:35 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120411 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-21 17:36 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120536 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-21 17:38 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-21 17:38 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120418 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-21 17:45 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120416 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-22 10:31 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120534 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-22 10:31 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 10:32 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120379 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 10:33 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-22 10:34 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120380 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 10:50 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 10:51 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-22 10:51 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120486 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 10:52 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120382 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-05-22 10:53 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 10:55 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 11:05 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120621 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-22 11:06 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120618 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 21.1 | Fail |  | high when on sleep | n/a
2024-05-22 11:07 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 23.49 | Fail |  | high when on sleep | n/a
2024-05-22 11:09 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120421 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 11:09 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120599 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-22 11:10 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120516 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 11:15 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120622 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 11:17 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120624 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-22 11:18 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120550 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 11:19 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120457 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 11:21 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120458 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 11:28 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120533 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 11:46 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 11:46 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120343 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 11:47 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120514 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 11:48 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120470 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 11:48 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 11:51 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120426 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 12:01 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120423 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 12:01 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-22 12:02 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-22 12:02 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120535 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 12:04 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120427 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 25.22 | Fail |  | high current on sleep mode | n/a
2024-05-22 12:06 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120525 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 12:21 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120434 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 12:22 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120341 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 12:25 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120412 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 12:30 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.428 | Fail |  | high reading on sleep mode | n/a
2024-05-22 12:32 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120376 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 12:32 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120531 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 12:36 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 12:40 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120440 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 12:41 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120527 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 12:43 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120377 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 12:45 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120460 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-22 12:48 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120342 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-22 16:34 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120482 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 16:36 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-22 16:44 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 16:46 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120484 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 16:48 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 16:50 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120452 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-22 16:51 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120526 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 16:54 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120451 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 16:57 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120417 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-22 16:59 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120619 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.175 | Pass | Pass |  | N/A
2024-05-22 17:02 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120419 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 17:05 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 17:07 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120453 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 17:08 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120552 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 17:11 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 17:15 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120422 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 17:16 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120638 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 17:18 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120623 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 17:19 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 17:21 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120551 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 17:32 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 17:34 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120447 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 17:37 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120439 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 17:41 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120596 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 17:42 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 17:45 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120463 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-22 17:46 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120459 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 427 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 17:48 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120559 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 17:51 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120414 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 17:53 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120580 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 17:54 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-22 17:57 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120572 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-22 17:59 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120467 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-22 18:00 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120542 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-22 18:02 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120672 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-22 18:07 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120626 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-22 18:09 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120620 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-22 18:11 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120448 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 07:18 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 23.53 | Fail |  | value out range HIgh Reading during sleep | n/a
2024-05-23 07:21 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120648 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 07:24 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120640 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 07:25 | stmstation5@avive.life | W1266 | Agree | Leonora Cruz | 2404120614 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-23 07:25 | stmstation4@avive.life | W1266 | Agree | Leonora Cruz | 2404120677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 07:37 | stmstation5@avive.life | W1266 | Agree | Leonora Cruz | 2404120511 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 422 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-23 07:37 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120646 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-23 07:38 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-23 07:43 | stmstation4@avive.life | W1266 | Agree | Leonora Cruz | 2404120563 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 07:47 | stmstation3@avive.life | W1266 | Agree | Leonora Cruz | 2404120613 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-23 07:48 | stmstation2@avive.life | W1266 | Agree | Leonora Cruz | 2404120678 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 08:23 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120462 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 318 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 08:37 | stmstation1@avive.life | W1266 | Agree | Leonora Cruz | 2404120656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 08:43 | stmstation4@avive.life | W1266 | Agree | Leonora Cruz | 2404120515 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 08:44 | stmstation5@avive.life | W1266 | Agree | Leonora Cruz | 2404120651 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 08:44 | stmstation6@avive.life | W1266 | Agree | Leonora Cruz | 2404120628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-23 09:01 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120495 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 09:03 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120507 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 09:05 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120444 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 09:15 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 09:15 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120475 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 09:16 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120544 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 09:19 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 09:19 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 09:26 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 09:31 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120480 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-23 09:32 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 09:32 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120472 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-23 09:37 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 09:37 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120296 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-23 09:37 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120469 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-23 09:47 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 09:47 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 09:48 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120445 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 09:51 | stmstation6@avive.life | W1266 | Agree | Liza Ordillo | 2404120476 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-23 09:51 | stmstation5@avive.life | W1266 | Agree | Liza Ordillo | 2404120295 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-23 09:51 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120477 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 09:59 | stmstation1@avive.life | W1266 | Agree | Liza Ordillo | 2404120609 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 09:59 | stmstation2@avive.life | W1266 | Agree | Liza Ordillo | 2404120603 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 09:59 | stmstation3@avive.life | W1266 | Agree | Liza Ordillo | 2404120442 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-23 10:00 | stmstation4@avive.life | W1266 | Agree | Liza Ordillo | 2404120592 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-23 10:08 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120393 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-23 10:09 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120508 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 10:09 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120461 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 10:10 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120466 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-23 10:11 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120306 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-23 10:11 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120443 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 10:19 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120479 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 10:19 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120312 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 10:21 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120311 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 10:25 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120471 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 10:25 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 10:25 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120305 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 10:28 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120297 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 10:29 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120473 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 10:29 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120299 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 10:34 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-23 10:34 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120336 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 10:38 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120474 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-23 10:39 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120569 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 10:40 | stmstation2@avive.life | W1266 | Agree | Valeria Manzo | 2404120415 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 10:41 | stmstation3@avive.life | W1266 | Agree | Valeria Manzo | 2404120567 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 10:41 | stmstation6@avive.life | W1266 | Agree | Valeria Manzo | 2404120291 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-23 10:46 | stmstation5@avive.life | W1266 | Agree | Valeria Manzo | 2404120286 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 10:47 | stmstation4@avive.life | W1266 | Agree | Valeria Manzo | 2404120314 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 10:47 | stmstation1@avive.life | W1266 | Agree | Valeria Manzo | 2404120287 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-23 11:31 | stmstation1@avive.life | W1266 | Agree | Precious Ungco | 2404120529 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 422 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 11:32 | stmstation2@avive.life | W1266 | Agree | Precious Ungco | 2404120400 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 11:32 | stmstation3@avive.life | W1266 | Agree | Precious Ungco | 2404120388 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 12:17 | stmstation1@avive.life | W1266 | Agree | Precious Ungco | 2404120340 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-23 12:18 | stmstation2@avive.life | W1266 | Agree | Precious Ungco | 2404120578 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 424 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 12:18 | stmstation3@avive.life | W1266 | Agree | Precious Ungco | 2404120293 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 12:20 | stmstation6@avive.life | W1266 | Agree | Precious Ungco | 2404120285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 430 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-23 12:20 | stmstation5@avive.life | W1266 | Agree | Precious Ungco | 2404120399 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 12:21 | stmstation4@avive.life | W1266 | Agree | Precious Ungco | 2404120546 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 433 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 12:32 | stmstation6@avive.life | W1266 | Agree | Precious Ungco | 2404120310 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 12:33 | stmstation5@avive.life | W1266 | Agree | Precious Ungco | 2404120430 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-23 12:33 | stmstation4@avive.life | W1266 | Agree | Precious Ungco | 2404120490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-23 12:34 | stmstation3@avive.life | W1266 | Agree | Precious Ungco | 2404120334 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 433 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 12:35 | stmstation2@avive.life | W1266 | Agree | Precious Ungco | 2404120333 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 427 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 12:35 | stmstation1@avive.life | W1266 | Agree | Precious Ungco | 2404120437 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-23 13:27 | stmstation1@avive.life | W1266 | Agree | Precious Ungco | 2404120407 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 13:28 | stmstation2@avive.life | W1266 | Agree | Precious Ungco | 2404120404 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 13:29 | stmstation3@avive.life | W1266 | Agree | Precious Ungco | 2404120518 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 13:29 | stmstation6@avive.life | W1266 | Agree | Precious Ungco | 2404120403 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 13:30 | stmstation5@avive.life | W1266 | Agree | Precious Ungco | 2404120450 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 13:31 | stmstation4@avive.life | W1266 | Agree | Precious Ungco | 2404120456 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 13:50 | stmstation6@avive.life | W1266 | Agree | Precious Ungco | 2404120431 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 13:50 | stmstation5@avive.life | W1266 | Agree | Precious Ungco | 2404120433 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 13:50 | stmstation4@avive.life | W1266 | Agree | Precious Ungco | 2404120425 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-23 13:50 | stmstation3@avive.life | W1266 | Agree | Precious Ungco | 2404120665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 13:51 | stmstation2@avive.life | W1266 | Agree | Precious Ungco | 2404120420 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 13:51 | stmstation1@avive.life | W1266 | Agree | Precious Ungco | 2404120402 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 16:49 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120671 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 16:50 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120288 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 16:51 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 17:11 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120338 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 17:11 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120372 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-23 17:12 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120573 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 17:14 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120233 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-23 17:16 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120298 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-23 17:17 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120239 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-24 11:30 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-24 11:34 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-24 11:37 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120320 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-24 11:39 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120271 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-24 11:42 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120324 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-24 11:45 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120230 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-24 11:47 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120268 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-24 11:49 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120356 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-24 11:50 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120265 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-24 11:56 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-24 11:59 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.185 | Pass | Pass |  | N/A
2024-05-24 12:00 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120269 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-24 12:01 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120227 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-24 12:02 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120318 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-24 12:05 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-24 12:08 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120270 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-24 12:10 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-24 12:12 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-24 12:16 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120337 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-24 12:18 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120369 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-24 12:19 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120370 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-24 12:22 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120375 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-24 12:25 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120373 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-24 12:33 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120315 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-24 12:34 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120339 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-24 12:35 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120332 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-24 12:37 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120308 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-24 12:38 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120374 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-24 13:25 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120307 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-24 13:27 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120602 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-24 13:27 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-24 13:43 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120345 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-24 13:44 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120335 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-24 13:49 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120294 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-24 13:52 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120313 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-24 13:54 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-24 13:57 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404120367 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-24 13:58 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120292 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-24 13:59 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120630 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-24 14:02 | stmstation6@avive.life | W1266 | Agree | Christina Cruz | 2404120224 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.171 | Pass | Pass |  | n/a
2024-05-24 14:09 | stmstation5@avive.life | W1266 | Agree | Christina Cruz | 2404120323 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 22.35 | Fail |  | Hi reading | Hi reading during sleep
2024-05-24 14:11 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404120289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-24 14:12 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404120267 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | no cell | FAIL NO CELL
2024-05-24 14:32 | stmstation2@avive.life | W1266 | Agree | Christina Cruz | 2404300289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-24 14:33 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404300217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-24 14:33 | stmstation1@avive.life | W1266 | Agree | Christina Cruz | 2404300222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-24 14:50 | stmstation3@avive.life | W1266 | Agree | Christina Cruz | 2404300147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-24 15:08 | stmstation4@avive.life | W1266 | Agree | Christina Cruz | 2404120276 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 22.34 | Fail |  | Hi reading | FAIL
2024-05-28 10:29 | stmstation1@avive.life | W1266 | Agree | Ananya jain | 2404120524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 436 | Pass | Pass | 0.176 | Pass | Pass |  | Checked again on STM and passed
2024-05-28 11:06 | stmstation6@avive.life | W1272 | Agree | Liza Ordillo | 2404120278 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-28 11:06 | stmstation4@avive.life | W1272 | Agree | Liza Ordillo | 2404120319 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 11:12 | stmstation5@avive.life | W1272 | Agree | Liza Ordillo | 2404120353 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 11:14 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404120191 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-28 11:15 | stmstation6@avive.life | W1272 | Agree | Liza Ordillo | 2404120207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 11:16 | stmstation4@avive.life | W1272 | Agree | Liza Ordillo | 2404120213 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 11:16 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 11:20 | stmstation5@avive.life | W1272 | Agree | Liza Ordillo | 2404120327 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-28 11:20 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404120119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 11:22 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120354 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 11:24 | stmstation6@avive.life | W1272 | Agree | Liza Ordillo | 2404120322 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 11:25 | stmstation4@avive.life | W1272 | Agree | Liza Ordillo | 2404120041 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-28 11:26 | stmstation5@avive.life | W1272 | Agree | Liza Ordillo | 2404120045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 11:26 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404120330 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-28 11:30 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120326 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-28 11:40 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404120049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 11:40 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120250 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 11:41 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404120349 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-28 11:47 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404120238 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-28 11:52 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404120104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 11:53 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404120106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 11:56 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 11:59 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404120351 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 12:04 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 12:05 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404120099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-28 12:08 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404120328 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-28 12:09 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404120043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-28 12:10 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404120219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 23.15 | Fail |  | values out of range | n/a
2024-05-28 12:12 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404120266 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-28 12:22 | stmstation6@avive.life | W1272 | Agree | Liza Ordillo | 2404120216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-28 12:22 | stmstation5@avive.life | W1272 | Agree | Liza Ordillo | 2404120301 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 12:22 | stmstation4@avive.life | W1272 | Agree | Liza Ordillo | 2404120247 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-28 12:26 | stmstation2@avive.life | W1272 | Agree | Liza Ordillo | 2404120302 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-28 12:26 | stmstation3@avive.life | W1272 | Agree | Liza Ordillo | 2404120284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-28 12:26 | stmstation1@avive.life | W1272 | Agree | Liza Ordillo | 2404120248 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 12:34 | stmstation6@avive.life | W1272 | Agree | Liza Ordillo | 2404120355 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 12:34 | stmstation5@avive.life | W1272 | Agree | Liza Ordillo | 2404120321 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 12:35 | stmstation4@avive.life | W1272 | Agree | Liza Ordillo | 2404120316 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 12:40 | stmstation2@avive.life | W1272 | Agree | Liza Ordillo | 2404120348 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 12:43 | stmstation3@avive.life | W1272 | Agree | Liza Ordillo | 2404120277 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-28 12:45 | stmstation1@avive.life | W1272 | Agree | Liza Ordillo | 2404120325 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 12:46 | stmstation4@avive.life | W1272 | Agree | Liza Ordillo | 2404120304 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-28 12:47 | stmstation5@avive.life | W1272 | Agree | Liza Ordillo | 2404120283 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 12:55 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 22.45 | Fail |  | value out of range | n/a
2024-05-28 12:56 | stmstation4@avive.life | W1272 | Agree | Liza Ordillo | 2404120303 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 12:58 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404120209 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 12:59 | stmstation5@avive.life | W1272 | Agree | Liza Ordillo | 2404120281 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 21.94 | Fail |  | high current 21.94 | n/a
2024-05-28 13:01 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404120274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 280 | Fail |  |  |  |  | value out of range | n/a
2024-05-28 13:06 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404120061 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-28 13:07 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-28 16:26 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404120251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 438 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-28 16:29 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404120252 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-28 16:37 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404120193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 16:37 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404120195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 422 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-28 16:38 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404120194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-28 16:39 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404120057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-28 16:39 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404120196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 16:40 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404120197 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-28 16:51 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404120047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-28 16:52 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404120118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 16:52 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404120051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 16:57 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404120236 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-28 16:58 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404120042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-28 16:58 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404120046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-28 17:11 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404120035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 17:12 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404080700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 17:12 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404080696 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 17:13 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404080699 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-28 17:13 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404080695 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-28 17:14 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404120050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 17:23 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404120199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-28 17:24 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404120190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 17:24 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404120036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-28 17:25 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404120205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-28 17:25 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404120054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-28 17:25 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404120044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-29 09:41 | stmstation6@avive.life | W1272 | Agree | Ananya Jain | 2404120274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.176 | Pass | Pass |  | N/A
2024-05-29 09:56 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404120242 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-29 09:57 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404120212 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-29 09:58 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404080694 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-29 09:59 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-05-29 10:01 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404120032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-29 10:17 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404120331 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-29 10:18 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404120198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 10:18 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404120040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-29 10:19 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-29 10:20 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300263 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-29 10:31 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404120111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 10:35 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404120023 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-29 10:35 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404080698 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 10:36 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404080692 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 10:37 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404120255 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-29 10:37 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404080693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-29 12:02 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404120245 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 12:03 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 12:03 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300152 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-29 12:06 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404120208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-29 12:07 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300202 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-29 12:08 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404120241 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 12:16 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 12:17 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300174 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 12:27 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300172 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 12:27 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404120240 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 12:29 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404120105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-29 12:38 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300282 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-29 12:38 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-29 12:39 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404120100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 12:40 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120234 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 13:11 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 359 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 14:06 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300333 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 14:07 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404300148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 14:08 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300318 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 14:08 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 14:09 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300209 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 14:10 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404120771 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-29 14:24 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300170 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 14:25 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300269 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 14:26 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300301 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 14:26 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300231 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 14:27 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404300299 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 424 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 14:28 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300298 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 14:47 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300317 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-29 14:48 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 438 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-29 14:49 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300168 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-29 14:49 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 433 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 14:50 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404120800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 14:50 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 15:01 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300197 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-29 15:01 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404120246 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-29 15:02 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-29 15:02 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 15:03 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-29 15:04 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300247 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 15:14 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404120793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 15:14 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300233 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-29 15:15 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404300125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 15:15 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300232 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 15:16 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300259 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-29 15:16 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 07:11 | stmstation2@avive.life | W1272 | Agree | Christina Cruz | 2404300306 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 07:13 | stmstation3@avive.life | W1272 | Agree | Christina Cruz | 2404120792 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 07:24 | stmstation1@avive.life | W1272 | Agree | Christina Cruz | 2404300319 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-30 07:31 | stmstation2@avive.life | W1272 | Agree | Christina Cruz | 2404300249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 07:32 | stmstation3@avive.life | W1272 | Agree | Christina Cruz | 2404300130 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 07:34 | stmstation6@avive.life | W1272 | Agree | Christina Cruz | 2404300134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 07:35 | stmstation5@avive.life | W1272 | Agree | Christina Cruz | 2404300321 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 07:36 | stmstation4@avive.life | W1272 | Agree | Christina Cruz | 2404300337 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-30 07:38 | stmstation1@avive.life | W1272 | Agree | Christina Cruz | 2404300313 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 08:27 | stmstation3@avive.life | W1272 | Agree | Christina Cruz | 2404300142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 08:31 | stmstation1@avive.life | W1272 | Agree | Christina Cruz | 2404300265 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 08:45 | stmstation2@avive.life | W1272 | Agree | Christina Cruz | 2404300249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 08:48 | stmstation5@avive.life | W1272 | Agree | Christina Cruz | 2404300119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-30 08:49 | stmstation4@avive.life | W1272 | Agree | Christina Cruz | 2404120243 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 08:50 | stmstation6@avive.life | W1272 | Agree | Christina Cruz | 2404300132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-30 08:52 | stmstation1@avive.life | W1272 | Agree | Christina Cruz | 2404300126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 08:54 | stmstation2@avive.life | W1272 | Agree | Christina Cruz | 2404300136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-30 08:56 | stmstation3@avive.life | W1272 | Agree | Christina Cruz | 2404300116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 10:26 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404300290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 10:26 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 10:27 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300124 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-30 10:28 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404120214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-30 10:29 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 10:30 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 10:39 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 10:41 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-30 10:41 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300311 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 10:50 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404300131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 10:50 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300288 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 10:51 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300238 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 11:09 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300325 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 11:09 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404120798 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 11:12 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300336 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 11:12 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300294 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 11:13 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404300109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 11:17 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300303 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-30 11:25 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300296 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-30 11:28 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 11:28 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300082 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-30 11:29 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300236 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 11:31 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 11:34 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300234 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-30 11:45 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300256 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-30 11:47 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300138 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-30 11:51 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300295 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 11:58 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300309 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 12:04 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404300304 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-30 13:10 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300061 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 13:10 | stmstation2@avive.life | W1272 | Agree | Valeria Manzo | 2404300149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-30 13:10 | stmstation3@avive.life | W1272 | Agree | Valeria Manzo | 2404300072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 13:11 | stmstation4@avive.life | W1272 | Agree | Valeria Manzo | 2404300166 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 13:11 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404300066 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 13:12 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404300150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-30 13:22 | stmstation4@avive.life | W1272 | Agree | Valeria Manzo | 2404300001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 13:23 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404300010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 13:23 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404300011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 13:24 | stmstation2@avive.life | W1272 | Agree | Valeria Manzo | 2404300081 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 13:24 | stmstation3@avive.life | W1272 | Agree | Valeria Manzo | 2404300145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.19 | Pass | Pass |  | n/a
2024-05-30 13:24 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-30 13:34 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404300006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 13:35 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300018 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-30 13:35 | stmstation2@avive.life | W1272 | Agree | Valeria Manzo | 2404300004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 13:36 | stmstation3@avive.life | W1272 | Agree | Valeria Manzo | 2404300029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 13:36 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404300063 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 14:32 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300302 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 14:33 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404300297 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 14:33 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300024 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 428 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-30 14:34 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 14:35 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 14:35 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300261 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 14:49 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300339 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 14:49 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300322 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 14:50 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 14:50 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300334 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 14:51 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404300335 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 14:51 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300331 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 15:07 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 15:08 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 15:09 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 15:10 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300340 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.181 | Pass | Pass |  | Precious Ungco
2024-05-30 15:11 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404300320 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 15:11 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300310 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 424 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 15:19 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300262 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 15:20 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-30 15:21 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 15:21 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404300338 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 15:22 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-30 15:22 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-30 15:31 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-30 15:32 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300113 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 15:33 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-30 15:34 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-30 15:34 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404300127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-30 15:35 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 15:45 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300213 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-30 15:45 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-30 15:46 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-30 15:46 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300083 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-30 15:47 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404300023 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-30 15:48 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-30 16:04 | stmstation6@avive.life | W1272 | Agree | Precious Ungco | 2404300225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-30 16:04 | stmstation5@avive.life | W1272 | Agree | Precious Ungco | 2404120797 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-30 16:04 | stmstation4@avive.life | W1272 | Agree | Precious Ungco | 2404300112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-30 16:04 | stmstation1@avive.life | W1272 | Agree | Precious Ungco | 2404300287 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-30 16:05 | stmstation2@avive.life | W1272 | Agree | Precious Ungco | 2404300110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-30 16:05 | stmstation3@avive.life | W1272 | Agree | Precious Ungco | 2404300108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-31 06:15 | stmstation4@avive.life | W1272 | Agree | Valeria Manzo | 2404300026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-31 06:15 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404120795 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-31 06:15 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404120189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 06:16 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-31 06:16 | stmstation2@avive.life | W1272 | Agree | Valeria Manzo | 2404300101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 06:16 | stmstation3@avive.life | W1272 | Agree | Valeria Manzo | 2404300033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-31 06:43 | stmstation4@avive.life | W1272 | Agree | Valeria Manzo | 2404300065 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 06:43 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404080669 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-31 06:43 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404300184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-31 06:44 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-31 06:44 | stmstation2@avive.life | W1272 | Agree | Valeria Manzo | 2404300115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-31 06:44 | stmstation3@avive.life | W1272 | Agree | Valeria Manzo | 2404120791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-31 06:54 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404300095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 23.44 | Fail |  | to high when on sleep mode | n/a
2024-05-31 06:55 | stmstation4@avive.life | W1272 | Agree | Valeria Manzo | 2404300012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-05-31 06:56 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404300091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 06:56 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-31 08:02 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404300300 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 10:02 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 22.82 | Fail |  | value out of range | n/a
2024-05-31 10:07 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300308 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-31 10:09 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300324 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-31 10:10 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404300315 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 10:11 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-31 10:12 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404120568 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 10:18 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 10:26 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-31 10:26 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-31 10:27 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300316 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 10:30 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 10:32 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-31 10:32 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404300274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-31 10:47 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300307 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-31 10:48 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300312 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 10:52 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 10:52 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 10:55 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404300270 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-31 11:05 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 11:05 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 11:06 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404300067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 11:07 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-31 11:07 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300041 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 11:20 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 23.45 | Fail |  | value out of range | n/a
2024-05-31 11:21 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 11:23 | stmstation4@avive.life | W1272 | Agree | Leonora Cruz | 2404300089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 22.65 | Fail |  | value out of range | n/a
2024-05-31 11:23 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-31 11:24 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-05-31 12:23 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300103 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 12:24 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-31 12:27 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300229 | Pass | Fail | [11 empty] | screen won't turn on | n/a
2024-05-31 12:32 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 22.55 | Fail |  | value out of range | n/a
2024-05-31 12:37 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300019 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-05-31 12:39 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300314 | Pass | Pass | Pass | Pass | Pass | Fail | [7 empty] | not syncing | n/a
2024-05-31 12:43 | stmstation2@avive.life | W1272 | Agree | Leonora Cruz | 2404300266 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-31 12:44 | stmstation3@avive.life | W1272 | Agree | Leonora Cruz | 2404300048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.186 | Pass | Pass |  | n/a
2024-05-31 12:46 | stmstation1@avive.life | W1272 | Agree | Leonora Cruz | 2404300210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-31 12:47 | stmstation6@avive.life | W1272 | Agree | Leonora Cruz | 2404300043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-31 12:49 | stmstation5@avive.life | W1272 | Agree | Leonora Cruz | 2404300129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-31 13:29 | stmstation4@avive.life | W1272 | Agree | Valeria Manzo | 2404300235 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-31 13:30 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404120796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-31 13:31 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404300002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 13:31 | stmstation2@avive.life | W1272 | Agree | Valeria Manzo | 2404300039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-31 13:32 | stmstation3@avive.life | W1272 | Agree | Valeria Manzo | 2404300071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-05-31 13:32 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-31 13:42 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404300332 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-31 13:43 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404300284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-31 13:43 | stmstation4@avive.life | W1272 | Agree | Valeria Manzo | 2404300323 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-31 13:45 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-31 13:46 | stmstation2@avive.life | W1272 | Agree | Valeria Manzo | 2404300020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-05-31 13:47 | stmstation3@avive.life | W1272 | Agree | Valeria Manzo | 2404300037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-31 13:54 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404300036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-05-31 13:55 | stmstation2@avive.life | W1272 | Agree | Valeria Manzo | 2404300046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 13:56 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300283 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 13:56 | stmstation4@avive.life | W1272 | Agree | Valeria Manzo | 2404300038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-31 13:57 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404300025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 13:58 | stmstation3@avive.life | W1272 | Agree | Valeria Manzo | 2404300171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 14:05 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404300044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 14:06 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404300084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-31 14:07 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-05-31 14:09 | stmstation3@avive.life | W1272 | Agree | Valeria Manzo | 2404300248 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-31 14:10 | stmstation2@avive.life | W1272 | Agree | Valeria Manzo | 2404300075 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-05-31 14:16 | stmstation4@avive.life | W1272 | Agree | Valeria Manzo | 2404300092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-05-31 14:17 | stmstation3@avive.life | W1272 | Agree | Valeria Manzo | 2404300085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-05-31 14:17 | stmstation1@avive.life | W1272 | Agree | Valeria Manzo | 2404300086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-05-31 14:18 | stmstation5@avive.life | W1272 | Agree | Valeria Manzo | 2404300047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-05-31 14:18 | stmstation6@avive.life | W1272 | Agree | Valeria Manzo | 2404300268 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-05-31 14:20 | stmstation2@avive.life | W1272 | Agree | Valeria Manzo | 2404300271 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-03 09:41 | stmstation4@avive.life | W1272 | Agree | Ananya Jain | 2404300229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.179 | Pass | Pass |  | N/A
2024-06-04 10:33 | stmstation6@avive.life | W1276 | Agree | Leonora Cruz | 2404300477 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-04 10:34 | stmstation5@avive.life | W1276 | Agree | Leonora Cruz | 2404300672 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-04 10:34 | stmstation4@avive.life | W1276 | Agree | Leonora Cruz | 2404300669 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-04 10:37 | stmstation1@avive.life | W1276 | Agree | Leonora Cruz | 2404300592 | Pass | Fail | [11 empty] | screen issue | n/a
2024-06-04 10:45 | stmstation2@avive.life | W1276 | Agree | Leonora Cruz | 2404300416 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-04 10:51 | stmstation1@avive.life | W1276 | Agree | Leonora Cruz | 2404300164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 318 | Fail |  |  |  |  | n/a | n/a
2024-06-05 10:22 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 10:23 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300714 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 10:23 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300279 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-05 10:27 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300568 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 10:28 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300571 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 10:38 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 10:38 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300639 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 10:38 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300372 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 10:39 | stmstation2@avive.life | W1276 | Agree | Liza Ordillo | 2404300418 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 10:39 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 10:43 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300466 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 10:57 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300697 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-05 10:57 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 10:58 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300792 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 11:00 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300620 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 11:01 | stmstation2@avive.life | W1276 | Agree | Liza Ordillo | 2404300622 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 11:01 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300440 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 11:10 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300689 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 11:11 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300453 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 11:11 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300797 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-05 11:12 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 11:13 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300385 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 11:13 | stmstation2@avive.life | W1276 | Agree | Liza Ordillo | 2404300476 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 11:43 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300450 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 11:45 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300435 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 11:47 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 11:47 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300423 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 11:49 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300447 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 11:56 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300380 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 11:56 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 11:57 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300424 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 11:57 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300243 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 11:59 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 12:01 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 12:10 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300590 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 24.54 | Fail |  | current to high when asleep | n/a
2024-06-05 12:12 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 12:13 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300602 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 12:14 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 12:14 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 12:17 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300464 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 12:25 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300608 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 12:25 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300734 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 12:26 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300432 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 12:26 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-05 12:27 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300698 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 12:27 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300694 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-05 12:41 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 12:42 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300696 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 12:42 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300610 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 12:43 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 12:45 | stmstation2@avive.life | W1276 | Agree | Liza Ordillo | 2404300648 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 12:45 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300508 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 12:47 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300776 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 13:03 | stmstation4@avive.life | W1276 | Agree | Precious Ungco | 2404300401 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 13:04 | stmstation5@avive.life | W1276 | Agree | Precious Ungco | 2404300690 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 13:04 | stmstation6@avive.life | W1276 | Agree | Precious Ungco | 2404300361 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-05 13:05 | stmstation3@avive.life | W1276 | Agree | Precious Ungco | 2404300790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 13:06 | stmstation2@avive.life | W1276 | Agree | Precious Ungco | 2404300428 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 13:06 | stmstation1@avive.life | W1276 | Agree | Precious Ungco | 2404300691 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 13:36 | stmstation3@avive.life | W1276 | Agree | Precious Ungco | 2404300546 | Pass | Pass | Fail | [10 empty] | audio doesnt work | n/a
2024-06-05 13:39 | stmstation4@avive.life | W1276 | Agree | Precious Ungco | 2404300757 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-05 13:39 | stmstation5@avive.life | W1276 | Agree | Precious Ungco | 2404300404 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 13:40 | stmstation6@avive.life | W1276 | Agree | Precious Ungco | 2404300496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 13:41 | stmstation1@avive.life | W1276 | Agree | Precious Ungco | 2404300397 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 13:41 | stmstation2@avive.life | W1276 | Agree | Precious Ungco | 2404300357 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 13:42 | stmstation3@avive.life | W1276 | Agree | Precious Ungco | 2404300499 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-05 13:47 | stmstation3@avive.life | W1276 | Agree | Precious Ungco | 2404300185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 429 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 13:49 | stmstation6@avive.life | W1276 | Agree | Precious Ungco | 2404300474 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 13:50 | stmstation5@avive.life | W1276 | Agree | Precious Ungco | 2404300539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 13:50 | stmstation4@avive.life | W1276 | Agree | Precious Ungco | 2404300506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-05 14:00 | stmstation4@avive.life | W1276 | Agree | Precious Ungco | 2404300618 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 14:01 | stmstation5@avive.life | W1276 | Agree | Precious Ungco | 2404300422 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 14:01 | stmstation6@avive.life | W1276 | Agree | Precious Ungco | 2404300379 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 14:02 | stmstation3@avive.life | W1276 | Agree | Precious Ungco | 2404300409 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 14:02 | stmstation2@avive.life | W1276 | Agree | Precious Ungco | 2404300374 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 14:03 | stmstation1@avive.life | W1276 | Agree | Precious Ungco | 2404300619 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 14:20 | stmstation3@avive.life | W1276 | Agree | Precious Ungco | 2404300278 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 14:20 | stmstation2@avive.life | W1276 | Agree | Precious Ungco | 2404300389 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 422 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 14:21 | stmstation1@avive.life | W1276 | Agree | Precious Ungco | 2404300410 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-05 14:22 | stmstation4@avive.life | W1276 | Agree | Precious Ungco | 2404300354 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 14:22 | stmstation5@avive.life | W1276 | Agree | Precious Ungco | 2404300408 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-05 14:23 | stmstation6@avive.life | W1276 | Agree | Precious Ungco | 2404300617 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 23.365 | Fail |  | high current | n/a
2024-06-05 14:29 | stmstation6@avive.life | W1276 | Agree | Precious Ungco | 2404300406 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-05 15:04 | stmstation1@avive.life | W1276 | Agree | Precious Ungco | 2404300475 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-05 15:05 | stmstation2@avive.life | W1276 | Agree | Precious Ungco | 2404300638 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 15:05 | stmstation3@avive.life | W1276 | Agree | Precious Ungco | 2404300467 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-05 15:07 | stmstation6@avive.life | W1276 | Agree | Precious Ungco | 2404300402 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-05 15:07 | stmstation5@avive.life | W1276 | Agree | Precious Ungco | 2404300400 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 15:08 | stmstation4@avive.life | W1276 | Agree | Precious Ungco | 2404300472 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 15:18 | stmstation1@avive.life | W1276 | Agree | Precious Ungco | 2404300712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 15:19 | stmstation2@avive.life | W1276 | Agree | Precious Ungco | 2404300718 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-05 15:20 | stmstation6@avive.life | W1276 | Agree | Precious Ungco | 2404300604 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 433 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-05 15:20 | stmstation5@avive.life | W1276 | Agree | Precious Ungco | 2404300629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-05 15:21 | stmstation4@avive.life | W1276 | Agree | Precious Ungco | 2404300687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 15:23 | stmstation3@avive.life | W1276 | Agree | Precious Ungco | 2404300641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 422 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-05 15:36 | stmstation3@avive.life | W1276 | Agree | Precious Ungco | 2404300431 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-05 15:36 | stmstation2@avive.life | W1276 | Agree | Precious Ungco | 2404300759 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-05 15:36 | stmstation1@avive.life | W1276 | Agree | Precious Ungco | 2404300651 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-05 15:36 | stmstation4@avive.life | W1276 | Agree | Precious Ungco | 2404300649 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-05 15:36 | stmstation5@avive.life | W1276 | Agree | Precious Ungco | 2404300471 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-05 15:37 | stmstation6@avive.life | W1276 | Agree | Precious Ungco | 2404300434 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-06 05:41 | stmstation5@avive.life | W1276 | Agree | Leonora Cruz | 2404300504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 318 | Fail |  |  |  |  | value low | n/a
2024-06-06 05:49 | stmstation3@avive.life | W1276 | Agree | Leonora Cruz | 2404300492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 05:50 | stmstation2@avive.life | W1276 | Agree | Leonora Cruz | 2404300395 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-06 05:51 | stmstation1@avive.life | W1276 | Agree | Leonora Cruz | 2404300396 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 05:53 | stmstation6@avive.life | W1276 | Agree | Leonora Cruz | 2404300578 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 05:53 | stmstation4@avive.life | W1276 | Agree | Leonora Cruz | 2404300595 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 06:06 | stmstation5@avive.life | W1276 | Agree | Leonora Cruz | 2404300503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.297 | Fail |  | value high | n/a
2024-06-06 06:09 | stmstation4@avive.life | W1276 | Agree | Leonora Cruz | 2404300159 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 06:13 | stmstation6@avive.life | W1276 | Agree | Leonora Cruz | 2404300528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 06:28 | stmstation1@avive.life | W1276 | Agree | Leonora Cruz | 2404300384 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-06 06:30 | stmstation2@avive.life | W1276 | Agree | Leonora Cruz | 2404300161 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-06 06:38 | stmstation6@avive.life | W1276 | Agree | Leonora Cruz | 2404300491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 06:38 | stmstation4@avive.life | W1276 | Agree | Leonora Cruz | 2404300452 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 06:42 | stmstation2@avive.life | W1276 | Agree | Leonora Cruz | 2404300647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 22.5 | Fail |  | value out of range | n/a
2024-06-06 06:58 | stmstation1@avive.life | W1276 | Agree | Leonora Cruz | 2404300399 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 07:00 | stmstation3@avive.life | W1276 | Agree | Leonora Cruz | 2404300489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 07:01 | stmstation4@avive.life | W1276 | Agree | Leonora Cruz | 2404300439 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 07:05 | stmstation2@avive.life | W1276 | Agree | Leonora Cruz | 2404300653 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 07:05 | stmstation1@avive.life | W1276 | Agree | Leonora Cruz | 2404300494 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 07:09 | stmstation5@avive.life | W1276 | Agree | Leonora Cruz | 2404300657 | Pass | Pass | Pass | Pass | Pass | Fail | [7 empty] | not syncing | n/a
2024-06-06 07:12 | stmstation6@avive.life | W1276 | Agree | Leonora Cruz | 2404300652 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 07:17 | stmstation4@avive.life | W1276 | Agree | Leonora Cruz | 2404300392 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 07:19 | stmstation5@avive.life | W1276 | Agree | Leonora Cruz | 2404300737 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-06 07:24 | stmstation6@avive.life | W1276 | Agree | Leonora Cruz | 2404300787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 07:25 | stmstation4@avive.life | W1276 | Agree | Leonora Cruz | 2404300719 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-06 07:31 | stmstation1@avive.life | W1276 | Agree | Leonora Cruz | 2404300541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 07:32 | stmstation2@avive.life | W1276 | Agree | Leonora Cruz | 2404300542 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 07:34 | stmstation6@avive.life | W1276 | Agree | Leonora Cruz | 2404300588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 07:35 | stmstation5@avive.life | W1276 | Agree | Leonora Cruz | 2404300755 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 07:39 | stmstation4@avive.life | W1276 | Agree | Leonora Cruz | 2404300738 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 07:41 | stmstation2@avive.life | W1276 | Agree | Leonora Cruz | 2404300599 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.234 | Fail |  | value out of range | n/a
2024-06-06 07:58 | stmstation1@avive.life | W1276 | Agree | Leonora Cruz | 2404300564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 08:01 | stmstation6@avive.life | W1276 | Agree | Leonora Cruz | 2404300594 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 08:01 | stmstation5@avive.life | W1276 | Agree | Leonora Cruz | 2404300530 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-06 08:02 | stmstation4@avive.life | W1276 | Agree | Leonora Cruz | 2404300411 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 08:04 | stmstation3@avive.life | W1276 | Agree | Leonora Cruz | 2404300160 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 08:14 | stmstation2@avive.life | W1276 | Agree | Leonora Cruz | 2404300522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 08:24 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300529 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 08:24 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300773 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 08:25 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 08:25 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300383 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-06-06 08:25 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 08:26 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300660 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 08:37 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 08:37 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300495 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 08:38 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300441 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 08:39 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300538 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-06 08:40 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300699 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 08:47 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 22.4 | Fail |  | to high when asleep, tested 4 times | n/a
2024-06-06 08:53 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300765 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 08:54 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300670 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 08:55 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300407 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 08:55 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300415 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 08:56 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300444 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 09:38 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 09:39 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 09:39 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 09:40 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300327 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 09:40 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300275 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 09:42 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300445 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 09:55 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300777 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 09:55 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300417 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 09:56 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300644 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 09:57 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300593 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 22.52 | Fail |  | value to high when on sleep mode | n/a
2024-06-06 09:58 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 10:27 | stmstation2@avive.life | W1276 | Agree | Liza Ordillo | 2404300795 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 10:32 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300774 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 10:33 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300585 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 10:33 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300562 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 10:36 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300565 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 10:37 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300533 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 10:47 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300398 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-06 10:48 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300587 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 10:49 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 10:51 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300566 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 10:51 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300581 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 10:53 | stmstation2@avive.life | W1276 | Agree | Liza Ordillo | 2404300561 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 11:01 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300545 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 23.19 | Fail |  | high current on sleep mode | n/s
2024-06-06 11:02 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 11:03 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 11:05 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 11:06 | stmstation2@avive.life | W1276 | Agree | Liza Ordillo | 2404300405 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 11:07 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300778 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 11:16 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300162 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 11:17 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 11:18 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 11:21 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300394 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 11:22 | stmstation2@avive.life | W1276 | Agree | Liza Ordillo | 2404300577 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 11:23 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300393 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-06-06 11:35 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300583 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 11:37 | stmstation2@avive.life | W1276 | Agree | Liza Ordillo | 2404300793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 11:38 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 11:40 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300783 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 11:41 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 24.2 | Fail |  | current to slow when on high | n/a
2024-06-06 11:41 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 12:39 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300451 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 12:41 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300442 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 12:42 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300358 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 12:44 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300621 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 12:45 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300741 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 12:45 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300630 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 12:54 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300640 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 12:55 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300346 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 12:56 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300498 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 12:57 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300527 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 12:58 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300536 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-06 12:59 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300758 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 13:08 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300382 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-06 13:08 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 13:09 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300756 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 13:10 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300381 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 13:10 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300182 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 13:15 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300497 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 13:21 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 13:21 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300362 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 13:23 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 13:24 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 13:24 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 13:30 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300414 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 13:33 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300376 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 13:33 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300601 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.19 | Pass | Pass |  | n/a
2024-06-06 13:34 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300624 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 13:34 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 22.23 | Fail |  | current to high when asleep | n/a
2024-06-06 13:35 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300330 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-06 13:42 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300711 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 13:44 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 13:44 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300720 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 13:45 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300715 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 23.64 | Fail |  | to high when on sleep mode | n/a
2024-06-06 14:53 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 14:54 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 15:04 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-06 15:05 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 15:06 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300326 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-06 15:16 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 15:19 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300342 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 15:27 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 15:31 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300253 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-06 15:32 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300239 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 15:39 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300003 | Pass | Fail | [11 empty] | NO DISPLAY | FAIL
2024-06-06 15:44 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 22.45 | Fail |  | hi current during sleep | FAIL
2024-06-06 15:47 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 15:48 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300212 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-06 15:52 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 15:53 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 15:54 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300240 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 15:55 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 16:04 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 16:08 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300007 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-06 16:10 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 16:12 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300781 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-06 16:13 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 16:15 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-06 16:16 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-06 16:26 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300328 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 16:48 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300688 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-06 16:58 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-06 17:01 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300576 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-06 17:05 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300473 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-06 17:07 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300575 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-07 06:18 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300569 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-07 07:09 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300637 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-07 07:11 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-07 07:18 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300573 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-07 07:38 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-07 10:00 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300572 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-07 11:01 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-07 11:05 | stmstation2@avive.life | W1276 | Agree | Liza Ordillo | 2404300586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 382 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-07 11:12 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300603 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 373 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-07 12:43 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-07 12:43 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300761 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-07 12:47 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300767 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-07 12:50 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300570 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-07 12:54 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300606 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-07 12:58 | stmstation6@avive.life | W1276 | Agree | Liza Ordillo | 2404300709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-07 12:59 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300764 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-07 13:02 | stmstation3@avive.life | W1276 | Agree | Liza Ordillo | 2404300188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-07 13:02 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300668 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-07 13:08 | stmstation1@avive.life | W1276 | Agree | Liza Ordillo | 2404300596 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 23.51 | Fail |  | high reading on sleep mode | n/a
2024-06-07 13:13 | stmstation4@avive.life | W1276 | Agree | Liza Ordillo | 2404300369 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-07 13:14 | stmstation5@avive.life | W1276 | Agree | Liza Ordillo | 2404300359 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-10 07:49 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300580 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 07:51 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:02 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 08:03 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300348 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 08:07 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300772 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 08:13 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300421 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 08:16 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300454 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:17 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300436 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:22 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300341 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:27 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300438 | Pass | Fail | [11 empty] | TOUCHSCREEN NOT RESPONDING | Fail LCD
2024-06-10 08:35 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300349 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 08:37 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300433 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:38 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:40 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300659 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:42 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300535 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:46 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300642 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:48 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300779 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 08:52 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300420 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:54 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300671 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 08:56 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 08:59 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300427 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 09:05 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300355 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 09:06 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300353 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 09:07 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300716 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 09:19 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 09:21 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300419 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 09:24 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300785 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 09:30 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300329 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 09:32 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300443 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 09:34 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 09:43 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300631 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 09:45 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300448 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 09:46 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300634 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 09:52 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300632 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 09:53 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300403 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 09:54 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-10 10:00 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300371 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 10:02 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300351 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 362 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-10 10:28 | stmstation6@avive.life | W1276 | Agree | Christina Cruz | 2404300365 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 10:28 | stmstation5@avive.life | W1276 | Agree | Christina Cruz | 2404300356 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 10:29 | stmstation4@avive.life | W1276 | Agree | Christina Cruz | 2404300154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-10 10:34 | stmstation2@avive.life | W1276 | Agree | Christina Cruz | 2404300613 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 10:36 | stmstation3@avive.life | W1276 | Agree | Christina Cruz | 2404300368 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 10:37 | stmstation1@avive.life | W1276 | Agree | Christina Cruz | 2404300470 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 10:53 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300769 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 10:53 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300377 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-10 10:54 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300437 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 10:54 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300277 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 10:55 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300364 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 11:00 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300242 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-06-10 11:07 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300352 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 11:07 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300378 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 11:08 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300650 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 11:08 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300733 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 416 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 11:08 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300663 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 11:08 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300704 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 11:20 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 23.18 | Fail |  | to high when asleep | n/a
2024-06-10 11:21 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300662 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 11:22 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300732 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 11:23 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300775 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 11:23 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300607 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 11:24 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300646 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-10 11:37 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 11:37 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 11:38 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300413 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-10 11:38 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300531 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 11:39 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-10 11:43 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300526 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 11:50 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 11:51 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300654 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 11:51 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-10 11:51 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300391 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 11:53 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 11:54 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300426 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 12:04 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300412 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 12:04 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300429 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 12:06 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300390 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 12:06 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 12:07 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300605 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 12:07 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-10 12:18 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300246 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-10 12:18 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300363 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 12:18 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300469 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-10 12:19 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300276 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 12:19 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 12:20 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300509 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 12:30 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-10 12:30 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300446 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 12:31 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300241 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 12:31 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300343 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 12:32 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300360 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 12:41 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 12:41 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300366 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 12:42 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300655 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 12:42 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300370 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 12:42 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300387 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 12:45 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300768 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 12:58 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300762 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 12:59 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300600 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 13:00 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300347 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-10 13:00 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300344 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 13:01 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300345 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 13:03 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 22.65 | Fail |  | to high when asleep | n/a
2024-06-10 13:13 | stmstation1@avive.life | W1276 | Agree | Valeria Manzo | 2404300388 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 13:14 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300666 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-10 13:14 | stmstation3@avive.life | W1276 | Agree | Valeria Manzo | 2404300674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 13:15 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-10 13:15 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300373 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-10 13:16 | stmstation2@avive.life | W1276 | Agree | Valeria Manzo | 2404300367 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 13:30 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300547 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-10 13:39 | stmstation4@avive.life | W1276 | Agree | Valeria Manzo | 2404300598 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | no cell or sim but it does sync, opened and moved sim but still nothing | n/a
2024-06-10 13:42 | stmstation5@avive.life | W1276 | Agree | Valeria Manzo | 2404300798 | Pass | Fail | [11 empty] | no touch, needs new display | n/a
2024-06-10 14:23 | stmstation6@avive.life | W1276 | Agree | Valeria Manzo | 2404300090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-10 15:26 | stmstation4@avive.life | W1276 | Agree | Ananya Jain | 2404300438 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.182 | Pass | Pass |  | Rework on ribbon cable
2024-06-10 15:53 | stmstation5@avive.life | W1276 | Agree | Ananya Jain | 2404300003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.175 | Pass | Pass |  | Rework Ribbon Cable
2024-06-10 16:02 | stmstation2@avive.life | W1276 | Agree | Ananya Jain | 2404300504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.179 | Pass | Pass |  | Retested
2024-06-10 16:18 | stmstation5@avive.life | W1276 | Agree | Ananya Jain | 2404300592 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.178 | Pass | Pass |  | Reworked Screen
2024-06-10 16:19 | stmstation6@avive.life | W1276 | Agree | Ananya Jain | 2404300657 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.181 | Pass | Pass |  | Retested
2024-06-10 16:22 | stmstation4@avive.life | W1276 | Agree | Ananya Jain | 2404300546 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.18 | Pass | Pass |  | Retested
2024-06-11 05:22 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080160 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-11 05:23 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080529 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-11 05:25 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-11 05:28 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 05:29 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080209 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 05:31 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080296 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-06-11 05:40 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2404300770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 05:41 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080161 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-11 05:44 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2404300771 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 05:46 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2404300763 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 05:48 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080237 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.19 | Pass | Pass |  | n/a
2024-06-11 05:55 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080304 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 24.11 | Fail |  | high reading | n/a
2024-06-11 05:59 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-11 06:02 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080239 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-11 06:02 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-11 06:04 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 06:07 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 06:12 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-11 06:13 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080600 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 06:17 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 06:18 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 437 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-11 06:19 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 06:25 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080166 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 437 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 06:28 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080536 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-11 06:33 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 06:39 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080165 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 428 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 06:40 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080081 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 06:44 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 437 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 06:46 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080599 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 06:50 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-11 06:53 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080162 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 06:55 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-11 06:56 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080172 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 438 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-11 06:59 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-11 07:06 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.17 | Pass | Pass |  | n/a
2024-06-11 07:06 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 07:07 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-06-11 07:10 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 07:11 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080238 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 07:12 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080047 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 442 | Pass | Pass | 0.19 | Pass | Pass |  | n/a
2024-06-11 07:22 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 24.17 | Fail |  | high reading | n/a
2024-06-11 07:24 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080128 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-11 07:26 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 07:29 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-11 07:30 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-11 07:37 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080130 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-11 07:38 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080511 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 07:38 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-11 07:41 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-11 07:41 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 07:45 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 23.86 | Fail |  | high reading | n/a
2024-06-11 07:47 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 24.53 | Fail |  | high reading | n/a
2024-06-11 07:51 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 23.62 | Fail |  | high reading | n/a
2024-06-11 09:10 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080228 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-11 09:11 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 09:13 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 09:14 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080608 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-11 09:15 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080312 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 431 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 09:25 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 09:28 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080167 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 09:33 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.19 | Pass | Pass |  | n/a
2024-06-11 09:33 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 09:34 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 10:05 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 10:06 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 10:08 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 10:20 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080315 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.137 | Fail |  | out of range | n /a
2024-06-11 10:21 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-11 10:22 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080041 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 10:25 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 10:27 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 10:28 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-11 10:34 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080152 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 10:35 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080144 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 10:36 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-11 10:38 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080297 | Pass | Pass | Pass | Pass | Pass | Fail | [7 empty] | not syncing | n/a
2024-06-11 11:07 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080200 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-11 12:10 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.193 | Pass | Pass |  | n/a
2024-06-11 12:12 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-11 12:16 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080313 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-11 12:24 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080170 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-11 12:24 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 12:26 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080526 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.187 | Pass | Pass |  | n/a
2024-06-11 12:33 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080114 | Pass | Fail | [11 empty] | screen not turning on | n/a
2024-06-11 12:35 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080311 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 507 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 12:36 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-06-11 13:19 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2405080299 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-11 13:21 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2405080306 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 13:21 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2404300730 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 13:23 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2405080302 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 13:24 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2404300521 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 13:26 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2404300463 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-11 13:31 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2405080532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 13:45 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300514 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-11 13:49 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2405080032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 13:51 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2405080618 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 13:52 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2404300515 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 15:56 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2405080180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-11 15:57 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2405080153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 16:08 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2405080137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-11 16:13 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2405080138 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 16:15 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2405080607 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-11 16:16 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2404300754 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 16:19 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2405080597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-11 16:23 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2404300511 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-11 16:24 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2405080530 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-11 16:26 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2405080195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-11 16:29 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2405080298 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 16:30 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2405080300 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-11 16:32 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2405080303 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-11 16:46 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2405080187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-06-11 16:48 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2404300753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-11 16:48 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2405080135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.186 | Pass | Pass |  | n/a
2024-06-12 05:25 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 438 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 05:27 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-12 05:28 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2404300550 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 05:32 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-12 05:33 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 05:34 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2404300556 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-12 05:47 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 23.92 | Fail |  | high reading | n/a
2024-06-12 05:52 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080158 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 436 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-12 05:54 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080606 | Pass | Fail | [11 empty] | line on the screen | n/a
2024-06-12 05:57 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-12 05:59 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 06:01 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080212 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 06:03 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 06:05 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 06:15 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 06:16 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080527 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 06:17 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 06:23 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-12 06:24 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080533 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-12 06:25 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 06:30 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300510 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 06:31 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080307 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-12 06:32 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080534 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 06:39 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 06:40 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 06:44 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 06:44 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2404300760 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 06:45 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 06:46 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080525 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 06:50 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 06:52 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 07:02 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 07:04 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 07:06 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 07:08 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 07:09 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 07:16 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080233 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 07:18 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2404300518 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-12 07:19 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080202 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 07:25 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-12 07:26 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-12 07:28 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080227 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 07:37 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-12 07:39 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 07:41 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 07:45 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 07:49 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080234 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 07:50 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 07:52 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-12 07:53 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080610 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 07:59 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 08:05 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 08:07 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 08:10 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 539 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-12 08:10 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 08:11 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080531 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-12 08:12 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 08:19 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080236 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 08:20 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-12 08:21 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080235 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 10:08 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2404300483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 10:09 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2405080018 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-12 10:10 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2405080017 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 10:14 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 10:16 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2404300745 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 10:16 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2405080191 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 10:18 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2405080178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-12 10:22 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2405080091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 10:24 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2405080099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 10:35 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2404300735 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-12 10:36 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080007 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 520 | Pass | Pass | 0.179 | Pass | Pass |  | N/A
2024-06-12 10:42 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080016 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 10:43 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 10:57 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 500 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-12 10:58 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-12 10:59 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080595 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.19 | Pass | Pass |  | n/a
2024-06-12 11:00 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 11:12 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2404300722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 11:15 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2404300519 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-12 11:16 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080201 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 11:20 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 11:24 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080023 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 503 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-12 11:28 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080609 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 11:29 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080061 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 11:31 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-12 11:34 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2404300751 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 11:37 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2404300456 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-12 11:38 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080310 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 12:29 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-12 12:37 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 502 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-12 12:41 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2404300750 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 12:42 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080309 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 12:49 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-12 12:49 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-12 12:50 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 13:02 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080535 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 13:02 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 13:05 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080213 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 13:06 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 433 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-12 13:16 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080507 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-12 13:17 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 13:18 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2404300746 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-12 13:27 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-06-12 13:27 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080169 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-12 13:28 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-12 13:35 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2404300723 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-06-12 13:36 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080013 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-12 13:39 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080594 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 13:41 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2404300728 | Pass | Pass | Pass | Fail | [9 empty] | no GPS | n/a
2024-06-12 17:07 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2405080242 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 17:11 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2405080231 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-12 17:12 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2405080064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-12 17:13 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2405080075 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-12 17:15 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300683 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 05:29 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2404300480 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-13 05:30 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-13 05:32 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2404300481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-13 05:38 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2404300684 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-13 05:39 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080593 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-13 05:40 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-13 05:48 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-13 05:49 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2404300485 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-13 05:50 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080139 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-13 05:56 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-13 05:56 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 06:05 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 21.66 | Fail |  | high reading on sleep mode | n/a
2024-06-13 06:11 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080083 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 23.39 | Fail |  | high reading | n/a
2024-06-13 06:13 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-13 06:14 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-13 06:18 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-13 06:20 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080545 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-06-13 06:26 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2404300679 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-13 06:28 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300525 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-13 06:30 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2404300502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 431 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-13 06:31 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080613 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-13 06:33 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2404300487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-13 06:47 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2404300500 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 23.61 | Fail |  | high reading | n/a
2024-06-13 06:49 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300725 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 06:50 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2404300486 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-13 06:51 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-13 06:51 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2404300559 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-13 06:52 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080592 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-13 07:00 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-13 07:00 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2404300682 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-13 09:45 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2405080204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 09:47 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2405080024 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 09:50 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2405080544 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-13 09:52 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2405080065 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 09:58 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300685 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-13 10:06 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-06-13 10:09 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2405080230 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-13 10:11 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2405080074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 10:18 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2405080054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 22.543 | Fail |  | Hi reading during sleep | HI READING low/hi during sleep
2024-06-13 10:21 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2405080208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 10:23 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2405080021 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 10:25 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2405080078 | Pass | Fail | [11 empty] | Hi reading at start | FAIL @ start
2024-06-13 10:27 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2404300482 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 10:32 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2405080589 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 431 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-13 10:36 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2404300558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-13 11:07 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2404300549 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.176 | Pass | Pass |  | N/A
2024-06-13 11:07 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080082 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-13 11:10 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-13 11:10 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080550 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-13 11:11 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2404300557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-13 11:17 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.175 | Pass | Pass |  | N/A
2024-06-13 11:24 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.179 | Pass | Fail | Won't turn on after power switch | n/a
2024-06-13 11:28 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2404300512 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 505 | Pass | Pass | 0.18 | Pass | Pass |  | N/A
2024-06-13 11:29 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-13 11:30 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-13 11:40 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 500 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-13 11:41 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-13 11:46 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080175 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-13 11:47 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2404300748 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-13 12:48 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080177 | Pass | Pass | Pass | Fail | [9 empty] | no GPS | n/a
2024-06-13 12:51 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2404300484 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-13 12:52 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080077 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-06-13 12:54 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080508 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-13 12:54 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-13 12:55 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2405080190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-13 13:02 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2405080216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-13 13:04 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 323 | Fail |  |  |  |  | out of range | n/a
2024-06-13 13:14 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-13 13:23 | stmstation5@avive.life | W1280 | Agree | Leonora Cruz | 2405080126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-13 13:24 | stmstation6@avive.life | W1280 | Agree | Leonora Cruz | 2404300742 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 513 | Pass | Pass | 0.192 | Pass | Pass |  | n/a
2024-06-13 13:25 | stmstation4@avive.life | W1280 | Agree | Leonora Cruz | 2405080026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.192 | Pass | Pass |  | n/a
2024-06-13 13:26 | stmstation1@avive.life | W1280 | Agree | Leonora Cruz | 2405080084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-13 13:28 | stmstation2@avive.life | W1280 | Agree | Leonora Cruz | 2405080103 | Pass | Pass | Pass | Fail | [9 empty] | NO GPS | n/a
2024-06-13 13:30 | stmstation3@avive.life | W1280 | Agree | Leonora Cruz | 2404300677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 505 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-14 05:26 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 05:27 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2404300747 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 05:28 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 05:35 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2404300721 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-14 05:36 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2404300726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-14 05:38 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080521 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-14 05:44 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-14 05:44 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2404300780 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-14 05:44 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 05:51 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2404300686 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 05:53 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080020 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 05:53 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080224 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-14 06:05 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080070 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 06:09 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 06:10 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 06:21 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2404300680 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-14 06:22 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080073 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 06:27 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-14 06:33 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-14 06:33 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080244 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-14 06:39 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 24.12 | Fail |  | high reading | n/a
2024-06-14 06:42 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-14 06:43 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2404300681 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 06:45 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080241 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 06:52 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-14 06:52 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 06:55 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2404300460 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 06:58 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080509 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 06:59 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2404300458 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 06:59 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080240 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-06-14 07:13 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080019 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 07:14 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2404300710 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 424 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 07:16 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2404300478 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 07:19 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080022 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 07:21 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2404300457 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-14 07:22 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-14 07:30 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080510 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 07:30 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 07:34 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300501 | Pass | Fail | [11 empty] | No Display Hi current at start | FAIL DISPLAY
2024-06-14 07:35 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2404300727 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 07:36 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 08:15 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2404300552 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-14 08:18 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-14 08:21 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080066 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 08:24 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2404300517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 08:24 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080546 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 08:27 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 08:30 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080079 | Pass | Fail | [11 empty] | no display, high current | n/a
2024-06-14 09:02 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2404300724 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-14 09:03 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2404300455 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 09:05 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2404300678 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 09:09 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080015 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-14 09:10 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 09:17 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080232 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 09:21 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2404300554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 23.7 | Fail |  | high current | n/a
2024-06-14 09:26 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-14 09:27 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-14 09:31 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080014 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 09:32 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-14 09:35 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2404300563 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-14 09:45 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080243 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 09:47 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2405080034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-14 09:52 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 437 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-14 09:55 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-14 09:56 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 09:58 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2404300551 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-14 10:01 | stmstation2@avive.life | W1280 | Agree | Liza Ordillo | 2405080028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-14 10:02 | stmstation1@avive.life | W1280 | Agree | Liza Ordillo | 2405080205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 433 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 10:07 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080549 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-14 10:08 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300461 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-14 10:10 | stmstation3@avive.life | W1280 | Agree | Liza Ordillo | 2404300462 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 13:27 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2404300076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-14 13:38 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2404300203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-14 13:39 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2404300743 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-14 13:47 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2404300055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-14 13:48 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2405080140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 13:49 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2404300219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 13:51 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2404300054 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 13:52 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300291 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 442 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 14:17 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2405080121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 14:19 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2404300070 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 14:22 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2404300097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 14:23 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2404300286 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-14 14:33 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300250 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-14 14:35 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2404300056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 14:37 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2404300051 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-14 14:40 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2404300093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-14 14:45 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2404300204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-14 14:47 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-14 14:49 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2404300293 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-14 14:51 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2404300292 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 14:54 | stmstation6@avive.life | W1280 | Agree | Christina Cruz | 2404300189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-14 14:56 | stmstation5@avive.life | W1280 | Agree | Christina Cruz | 2404300749 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-14 14:56 | stmstation4@avive.life | W1280 | Agree | Christina Cruz | 2405080619 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 14:57 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-14 14:59 | stmstation3@avive.life | W1280 | Agree | Christina Cruz | 2404300167 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-14 14:59 | stmstation2@avive.life | W1280 | Agree | Christina Cruz | 2404300520 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-17 07:27 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300165 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-17 07:42 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2405080088 | Pass | Pass | Fail | [10 empty] | Manuf card removed before testing | FAIL
2024-06-17 07:53 | stmstation1@avive.life | W1280 | Agree | Christina Cruz | 2404300057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-17 09:27 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.181 | Pass | Pass |  | Retested
2024-06-17 09:29 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.18 | Pass | Pass |  | Retested
2024-06-17 09:35 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080088 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.178 | Pass | Pass |  | Retested
2024-06-17 09:37 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.182 | Pass | Pass |  | Retested
2024-06-17 09:40 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2404300501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.179 | Pass | Pass |  | Retested
2024-06-17 09:46 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080606 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.176 | Pass | Pass |  | Retested
2024-06-17 09:46 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080297 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | Retested
2024-06-17 09:53 | stmstation6@avive.life | W1280 | Agree | Liza Ordillo | 2405080193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 427 | Pass | Pass | 0.179 | Pass | Pass |  | Retested
2024-06-17 09:54 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2404300500 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.181 | Pass | Pass |  | Retested
2024-06-17 10:01 | stmstation4@avive.life | W1280 | Agree | Liza Ordillo | 2405080114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | Retested
2024-06-17 10:03 | stmstation5@avive.life | W1280 | Agree | Liza Ordillo | 2405080078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.181 | Pass | Pass |  | Retested
2024-06-17 14:54 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080252 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-17 14:55 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080403 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-17 14:57 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080570 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-17 14:58 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080383 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-17 14:59 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080440 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-17 15:00 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-17 15:05 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080410 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-17 15:05 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080346 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-17 15:08 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080349 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-17 15:10 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080457 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-17 15:20 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080322 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-17 15:20 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080278 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-17 15:20 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080562 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-17 15:21 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080272 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-17 15:21 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080771 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-17 15:21 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080294 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-17 15:29 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080569 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.192 | Pass | Pass |  | n/a
2024-06-17 15:31 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080437 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-17 15:33 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080245 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-17 15:33 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080385 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-17 15:34 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080264 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-17 15:36 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080268 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-17 15:36 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080378 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-17 15:43 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080273 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-17 15:44 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080266 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-17 15:45 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080431 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-17 15:47 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080380 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-17 15:50 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080585 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-17 15:54 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080465 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-17 15:54 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080721 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-17 15:55 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080563 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-17 15:56 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080678 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-17 15:58 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-17 15:58 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080777 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-17 16:01 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080602 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-17 16:01 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080615 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-17 16:07 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080568 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-18 09:03 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080286 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 09:03 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080500 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-18 09:04 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080247 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-18 09:04 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080340 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 09:05 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080774 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-18 09:13 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-18 09:13 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080450 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-18 09:14 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 09:15 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080277 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-18 09:16 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080341 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-18 09:17 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080477 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-18 09:32 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080348 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 09:32 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080498 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-18 09:33 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080395 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-18 09:34 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080289 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-18 09:34 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080766 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-18 09:34 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080463 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 10:13 | stmstation3@avive.life | W1282 | Agree | Liza Ordillo | 2405080394 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 10:16 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080392 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-18 10:19 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080432 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 442 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-18 10:24 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080436 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 10:25 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080282 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-18 10:26 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080359 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-18 10:35 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080344 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 10:36 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-18 10:37 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080462 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 10:39 | stmstation3@avive.life | W1282 | Agree | Liza Ordillo | 2405080524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 10:42 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-18 10:43 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080353 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 23.87 | Fail |  | high current on sleep mode | n/a
2024-06-18 10:50 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080458 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-18 10:52 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080281 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 10:59 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080362 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-18 11:01 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.179 | Pass | Fail | Weird noise, reading high- 00.765 a | N/A
2024-06-18 11:01 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080387 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 11:03 | stmstation3@avive.life | W1282 | Agree | Liza Ordillo | 2405080480 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 11:07 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080279 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-18 11:16 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080459 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-18 11:18 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080455 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 11:18 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080347 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 11:21 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080283 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-18 11:22 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080475 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-18 11:23 | stmstation3@avive.life | W1282 | Agree | Liza Ordillo | 2405080276 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-18 11:29 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080460 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 11:30 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080382 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 11:30 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080390 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 11:32 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080274 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 11:33 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080388 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 12:33 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080614 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-18 12:33 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080246 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-18 12:34 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080773 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 12:34 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080456 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-18 12:34 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080308 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-18 12:39 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-18 12:43 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080350 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-18 12:44 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080336 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-18 12:46 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-18 12:48 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080257 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-18 12:48 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-18 12:48 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080601 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-18 12:55 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-20 08:58 | stmstation1@avive.life | W1282 | Agree | Christina Cruz | 2405080275 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-20 09:07 | stmstation2@avive.life | W1282 | Agree | Christina Cruz | 2405080292 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-20 09:08 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080255 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-20 09:14 | stmstation5@avive.life | W1282 | Agree | Christina Cruz | 2405080314 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-20 09:20 | stmstation4@avive.life | W1282 | Agree | Christina Cruz | 2405080250 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-20 09:28 | stmstation2@avive.life | W1282 | Agree | Christina Cruz | 2405080262 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-20 09:30 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080398 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-20 09:33 | stmstation6@avive.life | W1282 | Agree | Christina Cruz | 2405080357 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-20 09:34 | stmstation5@avive.life | W1282 | Agree | Christina Cruz | 2405080249 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-20 09:40 | stmstation4@avive.life | W1282 | Agree | Christina Cruz | 2405080269 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-20 09:41 | stmstation1@avive.life | W1282 | Agree | Christina Cruz | 2405080626 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-20 09:42 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080446 | Pass | Pass | Fail | [10 empty] | No Audio | FAIL
2024-06-20 09:47 | stmstation2@avive.life | W1282 | Agree | Christina Cruz | 2405080470 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-20 11:12 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080655 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-20 11:13 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080623 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-20 11:14 | stmstation3@avive.life | W1282 | Agree | Liza Ordillo | 2405080433 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-20 11:15 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080372 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-20 11:16 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080284 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-20 11:17 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080345 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-20 11:24 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080384 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-20 11:24 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080386 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-20 11:27 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080424 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-20 11:27 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080754 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-20 11:29 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080373 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-20 11:31 | stmstation3@avive.life | W1282 | Agree | Liza Ordillo | 2405080792 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 23.47 | Fail |  | high current on sleep mode tried 2x | n/a
2024-06-20 11:40 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080798 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-20 11:40 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080370 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-20 11:41 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080461 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-20 13:40 | stmstation3@avive.life | W1282 | Agree | Leonora Cruz | 2405080755 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-20 13:44 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080476 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-20 13:51 | stmstation2@avive.life | W1282 | Agree | Leonora Cruz | 2405080722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 505 | Pass | Pass | 0.185 | Pass | Pass |  | N/A
2024-06-20 13:51 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080391 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-20 13:54 | stmstation6@avive.life | W1282 | Agree | Leonora Cruz | 2405080654 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-20 13:54 | stmstation5@avive.life | W1282 | Agree | Leonora Cruz | 2405080660 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 500 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-20 13:55 | stmstation4@avive.life | W1282 | Agree | Leonora Cruz | 2405080756 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-20 13:59 | stmstation3@avive.life | W1282 | Agree | Leonora Cruz | 2405080576 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 437 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-20 15:03 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080367 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-20 15:04 | stmstation2@avive.life | W1282 | Agree | Leonora Cruz | 2405080417 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-20 15:06 | stmstation3@avive.life | W1282 | Agree | Leonora Cruz | 2405080418 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-21 07:40 | stmstation1@avive.life | W1282 | Agree | Christina Cruz | 2405080451 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-06-21 08:04 | stmstation2@avive.life | W1282 | Agree | Christina Cruz | 2405080416 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-21 08:08 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-21 08:09 | stmstation4@avive.life | W1282 | Agree | Christina Cruz | 2405080442 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 08:09 | stmstation6@avive.life | W1282 | Agree | Christina Cruz | 2405080554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-21 08:11 | stmstation5@avive.life | W1282 | Agree | Christina Cruz | 2405080426 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 08:27 | stmstation5@avive.life | W1282 | Agree | Christina Cruz | 2405080339 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-21 08:28 | stmstation1@avive.life | W1282 | Agree | Christina Cruz | 2405080552 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-21 08:29 | stmstation2@avive.life | W1282 | Agree | Christina Cruz | 2405080389 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-21 08:30 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080441 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-21 08:32 | stmstation4@avive.life | W1282 | Agree | Christina Cruz | 2405080321 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 08:33 | stmstation6@avive.life | W1282 | Agree | Christina Cruz | 2405080578 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 08:51 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080444 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-21 08:53 | stmstation1@avive.life | W1282 | Agree | Christina Cruz | 2405080523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-21 08:54 | stmstation2@avive.life | W1282 | Agree | Christina Cruz | 2405080469 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-21 09:10 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080453 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-21 09:12 | stmstation6@avive.life | W1282 | Agree | Christina Cruz | 2405080466 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-21 09:14 | stmstation5@avive.life | W1282 | Agree | Christina Cruz | 2405080513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 09:18 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080288 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 09:28 | stmstation5@avive.life | W1282 | Agree | Christina Cruz | 2405080479 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-21 09:30 | stmstation6@avive.life | W1282 | Agree | Christina Cruz | 2405080582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-21 09:34 | stmstation2@avive.life | W1282 | Agree | Christina Cruz | 2405080423 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-21 09:34 | stmstation1@avive.life | W1282 | Agree | Christina Cruz | 2405080320 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-21 09:37 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080331 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 09:38 | stmstation4@avive.life | W1282 | Agree | Christina Cruz | 2405080499 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-21 09:41 | stmstation5@avive.life | W1282 | Agree | Christina Cruz | 2405080428 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 09:45 | stmstation6@avive.life | W1282 | Agree | Christina Cruz | 2405080497 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.191 | Pass | Pass |  | n/a
2024-06-21 09:46 | stmstation4@avive.life | W1282 | Agree | Christina Cruz | 2405080580 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 10:03 | stmstation2@avive.life | W1282 | Agree | Christina Cruz | 2405080715 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-21 10:13 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080360 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 23.45 | Fail |  | Hi current reading during sleep | FAIL
2024-06-21 11:16 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080263 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-21 11:18 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080515 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-21 11:18 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-21 11:19 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080414 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-21 11:21 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080518 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-21 11:21 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080327 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-21 11:32 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080741 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-21 11:33 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080368 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-21 11:36 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080573 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-21 11:36 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080464 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-21 11:37 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-06-21 12:54 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080421 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-21 12:54 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080323 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 12:54 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080581 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 12:58 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 508 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 12:58 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080342 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-21 12:58 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080438 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 13:05 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080358 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-21 13:07 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080317 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-21 13:07 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080723 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-21 13:09 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-21 13:10 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080408 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-21 13:11 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080420 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 06:25 | stmstation4@avive.life | W1282 | Agree | Precious Ungco | 2405080364 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 06:26 | stmstation5@avive.life | W1282 | Agree | Precious Ungco | 2405080542 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 06:27 | stmstation6@avive.life | W1282 | Agree | Precious Ungco | 2405080413 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 06:27 | stmstation1@avive.life | W1282 | Agree | Precious Ungco | 2405080586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 06:28 | stmstation2@avive.life | W1282 | Agree | Precious Ungco | 2405080489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-24 06:29 | stmstation3@avive.life | W1282 | Agree | Precious Ungco | 2405080512 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 06:53 | stmstation3@avive.life | W1282 | Agree | Precious Ungco | 2405080485 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 06:54 | stmstation2@avive.life | W1282 | Agree | Precious Ungco | 2405080295 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 06:54 | stmstation1@avive.life | W1282 | Agree | Precious Ungco | 2405080575 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-24 06:55 | stmstation4@avive.life | W1282 | Agree | Precious Ungco | 2405080584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 06:56 | stmstation5@avive.life | W1282 | Agree | Precious Ungco | 2405080483 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 06:56 | stmstation6@avive.life | W1282 | Agree | Precious Ungco | 2405080376 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 07:12 | stmstation6@avive.life | W1282 | Agree | Precious Ungco | 2405080325 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 07:12 | stmstation5@avive.life | W1282 | Agree | Precious Ungco | 2405080335 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 23.27 | Fail |  | high current | n/a
2024-06-24 07:14 | stmstation4@avive.life | W1282 | Agree | Precious Ungco | 2405080567 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 07:14 | stmstation3@avive.life | W1282 | Agree | Precious Ungco | 2405080491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 07:15 | stmstation2@avive.life | W1282 | Agree | Precious Ungco | 2405080564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 07:16 | stmstation1@avive.life | W1282 | Agree | Precious Ungco | 2405080587 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 07:38 | stmstation3@avive.life | W1282 | Agree | Precious Ungco | 2405080641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 07:39 | stmstation2@avive.life | W1282 | Agree | Precious Ungco | 2405080256 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-24 07:41 | stmstation1@avive.life | W1282 | Agree | Precious Ungco | 2405080768 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-24 07:42 | stmstation4@avive.life | W1282 | Agree | Precious Ungco | 2405080657 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 07:43 | stmstation5@avive.life | W1282 | Agree | Precious Ungco | 2405080369 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-06-24 07:44 | stmstation6@avive.life | W1282 | Agree | Precious Ungco | 2405080769 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 07:58 | stmstation6@avive.life | W1282 | Agree | Precious Ungco | 2405080775 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 07:59 | stmstation5@avive.life | W1282 | Agree | Precious Ungco | 2405080251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 07:59 | stmstation4@avive.life | W1282 | Agree | Precious Ungco | 2405080650 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 08:00 | stmstation3@avive.life | W1282 | Agree | Precious Ungco | 2405080454 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.194 | Pass | Pass |  | n/a
2024-06-24 08:00 | stmstation2@avive.life | W1282 | Agree | Precious Ungco | 2405080648 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-24 08:02 | stmstation1@avive.life | W1282 | Agree | Precious Ungco | 2405080571 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-24 08:23 | stmstation4@avive.life | W1282 | Agree | Precious Ungco | 2405080645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 08:24 | stmstation5@avive.life | W1282 | Agree | Precious Ungco | 2405080287 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 08:24 | stmstation6@avive.life | W1282 | Agree | Precious Ungco | 2405080538 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 08:25 | stmstation3@avive.life | W1282 | Agree | Precious Ungco | 2405080254 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 08:26 | stmstation2@avive.life | W1282 | Agree | Precious Ungco | 2405080642 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 08:26 | stmstation1@avive.life | W1282 | Agree | Precious Ungco | 2405080375 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 08:41 | stmstation3@avive.life | W1282 | Agree | Precious Ungco | 2405080577 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 08:41 | stmstation2@avive.life | W1282 | Agree | Precious Ungco | 2405080644 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 08:41 | stmstation1@avive.life | W1282 | Agree | Precious Ungco | 2405080561 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 08:42 | stmstation4@avive.life | W1282 | Agree | Precious Ungco | 2405080248 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 426 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 08:42 | stmstation5@avive.life | W1282 | Agree | Precious Ungco | 2405080539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 08:42 | stmstation6@avive.life | W1282 | Agree | Precious Ungco | 2405080290 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-24 08:58 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 24.14 | Fail |  | out of range | n/a
2024-06-24 09:03 | stmstation2@avive.life | W1282 | Agree | Leonora Cruz | 2405080667 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 525 | Pass | Pass | 24.82 | Fail |  | out of range & no GPS signal | n/a
2024-06-24 09:06 | stmstation3@avive.life | W1282 | Agree | Leonora Cruz | 2405080689 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-24 09:44 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080664 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 09:46 | stmstation3@avive.life | W1282 | Agree | Leonora Cruz | 2405080324 | Pass | Pass | Pass | Pass | Pass | Fail | [7 empty] | not syncing | n/a
2024-06-24 09:55 | stmstation2@avive.life | W1282 | Agree | Leonora Cruz | 2405080496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 09:58 | stmstation3@avive.life | W1282 | Agree | Leonora Cruz | 2405080625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 23.39 | Fail |  | out of range | n/a
2024-06-24 10:00 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 10:01 | stmstation6@avive.life | W1282 | Agree | Leonora Cruz | 2405080693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-24 10:02 | stmstation5@avive.life | W1282 | Agree | Leonora Cruz | 2405080680 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 10:02 | stmstation4@avive.life | W1282 | Agree | Leonora Cruz | 2405080688 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-24 10:10 | stmstation3@avive.life | W1282 | Agree | Leonora Cruz | 2405080495 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-24 10:11 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080757 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 10:18 | stmstation6@avive.life | W1282 | Agree | Leonora Cruz | 2405080352 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 10:18 | stmstation5@avive.life | W1282 | Agree | Leonora Cruz | 2405080445 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 10:19 | stmstation4@avive.life | W1282 | Agree | Leonora Cruz | 2405080622 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-24 10:25 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080603 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-24 10:31 | stmstation6@avive.life | W1282 | Agree | Leonora Cruz | 2405080724 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 423 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 10:31 | stmstation5@avive.life | W1282 | Agree | Leonora Cruz | 2405080659 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 10:32 | stmstation4@avive.life | W1282 | Agree | Leonora Cruz | 2405080351 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 10:37 | stmstation2@avive.life | W1282 | Agree | Leonora Cruz | 2405080604 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 10:39 | stmstation3@avive.life | W1282 | Agree | Leonora Cruz | 2405080653 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 506 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-24 10:44 | stmstation4@avive.life | W1282 | Agree | Leonora Cruz | 2405080574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 505 | Pass | Pass | 24.54 | Fail |  | out of range | n/a
2024-06-24 10:47 | stmstation6@avive.life | W1282 | Agree | Leonora Cruz | 2405080393 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-24 10:54 | stmstation4@avive.life | W1282 | Agree | Leonora Cruz | 2405080516 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 500 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 10:55 | stmstation5@avive.life | W1282 | Agree | Leonora Cruz | 2405080488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-24 10:56 | stmstation6@avive.life | W1282 | Agree | Leonora Cruz | 2405080363 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 11:07 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080365 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 11:11 | stmstation3@avive.life | W1282 | Agree | Leonora Cruz | 2405080790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-24 11:19 | stmstation4@avive.life | W1282 | Agree | Leonora Cruz | 2405080566 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-24 11:19 | stmstation5@avive.life | W1282 | Agree | Leonora Cruz | 2405080355 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 11:20 | stmstation6@avive.life | W1282 | Agree | Leonora Cruz | 2405080677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 11:25 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080411 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 11:26 | stmstation2@avive.life | W1282 | Agree | Leonora Cruz | 2405080517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 11:30 | stmstation1@avive.life | W1282 | Agree | Leonora Cruz | 2405080572 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 11:30 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080616 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 11:31 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080448 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-24 11:33 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080430 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 11:34 | stmstation3@avive.life | W1282 | Agree | Leonora Cruz | 2405080522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 11:36 | stmstation2@avive.life | W1282 | Agree | Leonora Cruz | 2405080519 | Pass | Pass | Pass | Fail | [9 empty] | no GPS & sim | n/a
2024-06-24 11:43 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080419 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-24 11:44 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080439 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 11:46 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080494 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 11:55 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080330 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 11:58 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080338 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 12:00 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080332 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 442 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 12:01 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080429 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 12:03 | stmstation3@avive.life | W1282 | Agree | Liza Ordillo | 2405080473 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 22.46 | Fail |  | high current when asleep, tried on different fixtures still high | n/a
2024-06-24 12:48 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080651 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 12:49 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-24 12:50 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 12:53 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080759 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-24 12:54 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080253 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-24 12:54 | stmstation3@avive.life | W1282 | Agree | Liza Ordillo | 2405080647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-24 13:24 | stmstation3@avive.life | W1282 | Agree | Liza Ordillo | 2405080267 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 13:24 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080559 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 13:25 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 13:26 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-24 13:26 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080758 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 13:27 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080366 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 13:40 | stmstation5@avive.life | W1282 | Agree | Liza Ordillo | 2405080540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 13:41 | stmstation6@avive.life | W1282 | Agree | Liza Ordillo | 2405080397 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.176 | Pass | Pass |  | N/A
2024-06-24 13:41 | stmstation2@avive.life | W1282 | Agree | Liza Ordillo | 2405080443 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 13:41 | stmstation3@avive.life | W1282 | Agree | Liza Ordillo | 2405080326 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-24 13:42 | stmstation1@avive.life | W1282 | Agree | Liza Ordillo | 2405080471 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-24 13:44 | stmstation4@avive.life | W1282 | Agree | Liza Ordillo | 2405080328 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-24 13:52 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080401 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 13:52 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080486 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 13:55 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080422 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 13:57 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080329 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 14:01 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080415 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-06-24 14:02 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080409 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 14:02 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080259 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 14:08 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080767 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 14:09 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080374 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 14:10 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080371 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-24 14:11 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080405 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 14:15 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080402 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 500 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 14:16 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080694 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-24 14:17 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080435 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-24 14:17 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080354 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-24 14:20 | stmstation2@avive.life | W1282 | Agree | Valeria Manzo | 2405080406 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.186 | Pass | Pass |  | n/a
2024-06-24 14:24 | stmstation5@avive.life | W1282 | Agree | Valeria Manzo | 2405080434 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 14:26 | stmstation6@avive.life | W1282 | Agree | Valeria Manzo | 2405080343 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-06-24 14:26 | stmstation3@avive.life | W1282 | Agree | Valeria Manzo | 2405080400 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 14:28 | stmstation4@avive.life | W1282 | Agree | Valeria Manzo | 2405080493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 24.61 | Fail |  | high current , tried 3 other fixtures still high | n/a
2024-06-24 14:30 | stmstation1@avive.life | W1282 | Agree | Valeria Manzo | 2405080656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-24 14:40 | stmstation1@avive.life | W1282 | Agree | Christina Cruz | 2405080649 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 14:42 | stmstation2@avive.life | W1282 | Agree | Christina Cruz | 2405080776 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 14:46 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080427 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 14:51 | stmstation1@avive.life | W1282 | Agree | Christina Cruz | 2405080543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-24 14:54 | stmstation2@avive.life | W1282 | Agree | Christina Cruz | 2405080490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-24 14:58 | stmstation3@avive.life | W1282 | Agree | Christina Cruz | 2405080583 | Pass | Fail | [11 empty] | NO DISPLAY | R/R display. FAIL
2024-06-24 15:09 | stmstation1@avive.life | W1282 | Agree | Christina Cruz | 2405080487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 23.18 | Fail |  | Hi current during sleep | FAIL
2024-06-24 16:51 | stmstation6@avive.life | W1282 | Agree | Ananya Jain | 2405080324 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.182 | Pass | Pass |  | Retested
2024-06-24 17:06 | stmstation6@avive.life | W1282 | Agree | Ananya Jain | 2405080583 | Pass | Pass | Pass | Pass | Pass | Pass | Fail | [6 empty] | No Cell | n/A
2024-06-25 07:16 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150600 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 07:20 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405080700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-06-25 07:27 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150104 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 07:30 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405080412 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 07:32 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150493 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 428 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-25 07:44 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150566 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-25 07:45 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150570 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 423 | Pass | Pass | 0.176 | Pass | Pass |  | n /a
2024-06-25 07:48 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080763 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-25 07:49 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150638 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 07:49 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150103 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 07:55 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150559 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-25 08:01 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150488 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 08:02 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 08:03 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150567 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 08:06 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150637 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 08:41 | stmstation5@avive.life | W1288 | Agree | Valeria Manzo | 2405150536 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 08:41 | stmstation6@avive.life | W1288 | Agree | Valeria Manzo | 2405150544 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 08:41 | stmstation4@avive.life | W1288 | Agree | Valeria Manzo | 2405150538 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 08:42 | stmstation3@avive.life | W1288 | Agree | Valeria Manzo | 2405150641 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 08:45 | stmstation2@avive.life | W1288 | Agree | Valeria Manzo | 2405150543 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-25 08:46 | stmstation1@avive.life | W1288 | Agree | Valeria Manzo | 2405150124 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 08:54 | stmstation3@avive.life | W1288 | Agree | Valeria Manzo | 2405150492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-25 08:55 | stmstation2@avive.life | W1288 | Agree | Valeria Manzo | 2405150537 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-25 08:55 | stmstation1@avive.life | W1288 | Agree | Valeria Manzo | 2405150563 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-25 08:57 | stmstation6@avive.life | W1288 | Agree | Valeria Manzo | 2405150541 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 08:57 | stmstation5@avive.life | W1288 | Agree | Valeria Manzo | 2405150530 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 08:57 | stmstation4@avive.life | W1288 | Agree | Valeria Manzo | 2405150546 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 09:05 | stmstation2@avive.life | W1288 | Agree | Valeria Manzo | 2405150545 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 09:06 | stmstation6@avive.life | W1288 | Agree | Valeria Manzo | 2405150132 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-25 09:07 | stmstation4@avive.life | W1288 | Agree | Valeria Manzo | 2405150491 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 09:08 | stmstation3@avive.life | W1288 | Agree | Valeria Manzo | 2405150527 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 09:09 | stmstation1@avive.life | W1288 | Agree | Valeria Manzo | 2405150126 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-06-25 09:12 | stmstation2@avive.life | W1288 | Agree | Valeria Manzo | 2405150534 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 09:17 | stmstation6@avive.life | W1288 | Agree | Valeria Manzo | 2405150495 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-25 09:18 | stmstation3@avive.life | W1288 | Agree | Valeria Manzo | 2405150571 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 09:18 | stmstation1@avive.life | W1288 | Agree | Valeria Manzo | 2405150526 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-25 09:19 | stmstation2@avive.life | W1288 | Agree | Valeria Manzo | 2405150148 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 09:20 | stmstation4@avive.life | W1288 | Agree | Valeria Manzo | 2405150525 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 09:25 | stmstation6@avive.life | W1288 | Agree | Valeria Manzo | 2405150149 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 423 | Pass | Pass | 0.172 | Pass | Pass |  | n/a
2024-06-25 09:28 | stmstation5@avive.life | W1288 | Agree | Valeria Manzo | 2405080305 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.179 | Pass | Pass |  | n/A
2024-06-25 09:28 | stmstation4@avive.life | W1288 | Agree | Valeria Manzo | 2405150572 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-25 09:29 | stmstation3@avive.life | W1288 | Agree | Valeria Manzo | 2405080520 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 09:29 | stmstation2@avive.life | W1288 | Agree | valeria Manzo | 2405150150 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 09:29 | stmstation1@avive.life | W1288 | Agree | Valeria Manzo | 2405150151 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 10:08 | stmstation3@avive.life | W1288 | Agree | Valeria Manzo | 2405150540 | Pass | Pass | Pass | Pass | Pass | Fail | [7 empty] | not syncing over 1 hour, tried other wifi networks and still doesnt sync..also made sure sim was ok still doesnt sync also doesnt show sim | n/a
2024-06-25 10:13 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405150121 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-25 10:18 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405150490 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 10:19 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405150542 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 10:20 | stmstation1@avive.life | W1288 | Agree | Liza Ordillo | 2404300459 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 10:22 | stmstation2@avive.life | W1288 | Agree | Liza Ordillo | 2405150539 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-25 10:29 | stmstation3@avive.life | W1288 | Agree | Liza Ordillo | 2405150129 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 10:36 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405150131 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 10:37 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405150118 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 10:38 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405150579 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 10:39 | stmstation1@avive.life | W1288 | Agree | Liza Ordillo | 2405150553 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-25 10:40 | stmstation2@avive.life | W1288 | Agree | Liza Ordillo | 2405080301 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 10:41 | stmstation3@avive.life | W1288 | Agree | Liza Ordillo | 2405150547 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 10:51 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405150504 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 10:52 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405080639 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 10:53 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405080634 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 10:55 | stmstation1@avive.life | W1288 | Agree | Liza Ordillo | 2405150112 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 10:57 | stmstation2@avive.life | W1288 | Agree | Liza Ordillo | 2405080631 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 10:58 | stmstation3@avive.life | W1288 | Agree | Liza Ordillo | 2405150613 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 11:01 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405150611 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 11:01 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405150548 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 11:02 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405080781 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 427 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-25 11:06 | stmstation2@avive.life | W1288 | Agree | Liza Ordillo | 2405150583 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-25 11:07 | stmstation3@avive.life | W1288 | Agree | Liza Ordillo | 2405080785 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 11:08 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405150582 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 11:08 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405150592 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 11:09 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405150096 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 11:14 | stmstation3@avive.life | W1288 | Agree | Liza Ordillo | 2405150595 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-25 11:16 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405150092 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-25 11:17 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405150616 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 430 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 11:18 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405080632 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 11:20 | stmstation3@avive.life | W1288 | Agree | Liza Ordillo | 2405080629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 11:24 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405150122 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 11:25 | stmstation6@avive.life | W1288 | Agree | Precious Ungco | 2405150564 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 11:25 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405150111 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-25 11:28 | stmstation3@avive.life | W1288 | Agree | Precious Ungco | 2405150496 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 11:33 | stmstation4@avive.life | W1288 | Agree | Precious Ungco | 2405080630 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 432 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 11:34 | stmstation5@avive.life | W1288 | Agree | Precious Ungco | 2405150489 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 11:34 | stmstation6@avive.life | W1288 | Agree | Precious Ungco | 2405150107 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 427 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 11:38 | stmstation1@avive.life | W1288 | Agree | Liza Ordillo | 2405150093 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 11:39 | stmstation3@avive.life | W1288 | Agree | Precious Ungco | 2405150557 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 11:43 | stmstation4@avive.life | W1288 | Agree | Precious Ungco | 2405150494 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 11:43 | stmstation5@avive.life | W1288 | Agree | Precious Ungco | 2405150574 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-25 11:44 | stmstation6@avive.life | W1288 | Agree | Precious Ungco | 2405150130 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 425 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 11:45 | stmstation3@avive.life | W1288 | Agree | Precious Ungco | 2405080636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 11:47 | stmstation1@avive.life | W1288 | Agree | Precious Ungco | 2405150108 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.178 | Pass | Pass |  | Aviv3mFg3m@il5
2024-06-25 11:47 | stmstation2@avive.life | W1288 | Agree | Liza Ordillo | 2405080633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Fail |  |  |  | no SIM | n/a
2024-06-25 11:50 | stmstation3@avive.life | W1288 | Agree | Precious Ungco | 2405150134 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 11:53 | stmstation4@avive.life | W1288 | Agree | Precious Ungco | 2405080753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 11:54 | stmstation5@avive.life | W1288 | Agree | Precious Ungco | 2405150601 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 11:54 | stmstation6@avive.life | W1288 | Agree | Precious Ungco | 2405150577 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 11:57 | stmstation3@avive.life | W1288 | Agree | Precious Ungco | 2405150619 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 11:57 | stmstation2@avive.life | W1288 | Agree | Precious Ungco | 2405150123 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 429 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-25 11:58 | stmstation1@avive.life | W1288 | Agree | Precious Ungco | 2405150558 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.173 | Pass | Pass |  | Aviv3mFg3m@il5
2024-06-25 12:09 | stmstation4@avive.life | W1288 | Agree | Precious Ungco | 2405150608 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 12:09 | stmstation5@avive.life | W1288 | Agree | Precious Ungco | 2405150618 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 12:24 | stmstation3@avive.life | W1288 | Agree | Precious Ungco | 2405150210 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 12:32 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2404300245 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-25 12:35 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150631 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 12:38 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150100 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-25 12:39 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405080635 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 12:40 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150005 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 12:40 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150073 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-25 15:12 | stmstation4@avive.life | W1282 | Agree | Ananya Jain | 2405080280 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.18 | Pass | Pass |  | Retested
2024-06-25 17:26 | stmstation6@avive.life | W1288 | Agree | Christina Cruz | 2405150561 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 412 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 17:27 | stmstation5@avive.life | W1288 | Agree | Christina Cruz | 2405080732 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 417 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 17:28 | stmstation4@avive.life | W1288 | Agree | Christina Cruz | 2405080748 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 17:30 | stmstation1@avive.life | W1288 | Agree | Christina Cruz | 2405150560 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-25 17:32 | stmstation3@avive.life | W1288 | Agree | Christina Cruz | 2405080736 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-25 17:33 | stmstation2@avive.life | W1288 | Agree | Christina Cruz | 2405080749 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 429 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 17:40 | stmstation2@avive.life | W1288 | Agree | Christina Cruz | 2405080743 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 17:42 | stmstation3@avive.life | W1288 | Agree | Christina Cruz | 2405080745 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 423 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-25 17:44 | stmstation6@avive.life | W1288 | Agree | Christina Cruz | 2405150003 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-25 17:45 | stmstation5@avive.life | W1288 | Agree | Christina Cruz | 2405150636 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-25 17:46 | stmstation4@avive.life | W1288 | Agree | Christina Cruz | 2405080699 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 423 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-25 17:47 | stmstation1@avive.life | W1288 | Agree | Christina Cruz | 2405080701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 423 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-25 17:54 | stmstation6@avive.life | W1288 | Agree | Christina Cruz | 2405150127 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.179 | Pass | Pass |  | N/A
2024-06-25 17:56 | stmstation4@avive.life | W1288 | Agree | Christina Cruz | 2405080762 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 17:57 | stmstation5@avive.life | W1288 | Agree | Christina Cruz | 2405150640 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 17:59 | stmstation1@avive.life | W1288 | Agree | Christina Cruz | 2405150584 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 18:00 | stmstation2@avive.life | W1288 | Agree | Christina Cruz | 2405150639 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-25 18:01 | stmstation3@avive.life | W1288 | Agree | Christina Cruz | 2405150125 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.179 | Pass | Pass |  | Christina Cruz
2024-06-25 18:08 | stmstation6@avive.life | W1288 | Agree | Christina Cruz | 2405080747 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 18:09 | stmstation5@avive.life | W1288 | Agree | Christina Cruz | 2405080737 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-25 18:13 | stmstation3@avive.life | W1288 | Agree | Christina Cruz | 2405080750 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 18:15 | stmstation2@avive.life | W1288 | Agree | Christina Cruz | 2405150549 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 18:16 | stmstation1@avive.life | W1288 | Agree | Christina Cruz | 2405150565 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-25 18:17 | stmstation6@avive.life | W1288 | Agree | Christina Cruz | 2405080728 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 18:19 | stmstation4@avive.life | W1288 | Agree | Christina Cruz | 2405080746 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-25 18:22 | stmstation2@avive.life | W1288 | Agree | Christina Cruz | 2405080744 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-25 18:23 | stmstation3@avive.life | W1288 | Agree | Christina Cruz | 2405080751 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-26 06:47 | stmstation3@avive.life | W1288 | Agree | Christina Cruz | 2404300350 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-26 06:54 | stmstation3@avive.life | W1288 | Agree | Christina Cruz | 2405080778 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-26 07:08 | stmstation6@avive.life | W1288 | Agree | Precious Ungco | 2405150099 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-26 07:09 | stmstation5@avive.life | W1288 | Agree | Precious Ungco | 2405080271 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-26 07:10 | stmstation2@avive.life | W1288 | Agree | Precious Ungco | 2405150610 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-26 07:14 | stmstation3@avive.life | W1288 | Agree | Precious Ungco | 2405150109 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-26 07:17 | stmstation1@avive.life | W1288 | Agree | Precious Ungco | 2405150629 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-26 07:19 | stmstation4@avive.life | W1288 | Agree | Precious Ungco | 2405150628 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 429 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-26 07:28 | stmstation2@avive.life | W1288 | Agree | Precious Ungco | 2405150085 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-26 07:44 | stmstation4@avive.life | W1288 | Agree | Precious Ungco | 2405150620 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 427 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-26 07:45 | stmstation6@avive.life | W1288 | Agree | Precious Ungco | 2405080467 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 414 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-26 07:46 | stmstation1@avive.life | W1288 | Agree | Precious Ungco | 2404300589 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 424 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-26 07:51 | stmstation2@avive.life | W1288 | Agree | Precious Ungco | 2405150105 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-26 08:00 | stmstation3@avive.life | W1288 | Agree | Precious Ungco | 2405080260 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-26 08:06 | stmstation1@avive.life | W1288 | Agree | Precious Ungco | 2405080779 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-26 08:07 | stmstation3@avive.life | W1288 | Agree | Precious Ungco | 2405150506 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | W1288
2024-06-26 08:39 | stmstation1@avive.life | W1288 | Agree | Precious Ungco | 2405080176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 24.03 | Fail |  | high current | n/a
2024-06-26 09:00 | stmstation3@avive.life | W1288 | Agree | Precious Ungco | 2405080198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 437 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-26 09:01 | stmstation4@avive.life | W1288 | Agree | Precious Ungco | 2405150061 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-26 09:26 | stmstation5@avive.life | W1288 | Agree | Precious Ungco | 2405080396 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 438 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-26 09:26 | stmstation6@avive.life | W1288 | Agree | Precious Ungco | 2405080782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 437 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-26 09:28 | stmstation2@avive.life | W1288 | Agree | Precious Ungco | 2405150142 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 410 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-26 11:06 | stmstation5@avive.life | W1288 | Agree | Christina Cruz | 2405080514 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-26 11:11 | stmstation1@avive.life | W1288 | Agree | Christina Cruz | 2405150604 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-26 11:15 | stmstation2@avive.life | W1288 | Agree | Christina Cruz | 2405150077 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-06-26 13:51 | stmstation6@avive.life | W1288 | Agree | Christina Cruz | 2404300623 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 22.69 | Fail |  | Hi current during sleep | FAIL
2024-06-27 06:18 | stmstation2@avive.life | W1288 | Agree | Christina Cruz | 2405150603 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-27 09:16 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150070 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-27 09:24 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150632 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-27 09:32 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150523 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-27 11:39 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2404300544 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-27 11:49 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150602 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-27 11:55 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2404300479 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-27 12:01 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150084 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-27 12:10 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150502 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-27 12:21 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150120 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-27 12:24 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080735 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-27 12:28 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405080789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 385 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-27 12:31 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150508 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.179 | Pass | Pass |  | N/A
2024-06-27 12:33 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150623 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-27 14:47 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150581 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-27 14:47 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405080658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-27 14:48 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150596 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-27 14:51 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405080784 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-27 14:52 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150219 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-27 14:57 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405080638 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 419 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-27 15:12 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405080703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-27 15:14 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150522 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-27 15:20 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150008 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-27 15:24 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150128 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-27 15:32 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150568 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-27 15:36 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150119 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-27 15:38 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150225 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-27 15:40 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150076 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-27 15:44 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150229 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-27 15:47 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150146 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-06-27 15:50 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150605 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-27 15:56 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150078 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-27 16:05 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150551 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-27 16:07 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150010 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-06-27 16:08 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150009 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-06-27 16:09 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150634 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 386 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-27 16:26 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405080730 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-27 16:27 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405080731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-27 16:31 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150576 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-27 16:33 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150116 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-27 16:35 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150645 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-06-27 16:43 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150578 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-06-27 16:45 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150529 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-06-27 16:49 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150667 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-27 16:49 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150117 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-06-27 16:50 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150664 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-06-27 16:53 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150002 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-06-27 17:00 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-06-27 17:01 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.18 | Pass | Pass |  | N/A
2024-07-01 06:32 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405080702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 06:41 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405080713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 06:43 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080698 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 06:45 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150214 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 07:15 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150507 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 07:16 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405080705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 07:17 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405080786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-01 07:23 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405080293 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-01 07:24 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150138 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 07:27 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405080708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 07:37 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150213 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 368 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 07:38 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150212 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 07:39 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405080707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 07:41 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405080726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 07:55 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150153 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 366 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 07:56 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405080706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.184 | Pass | Pass |  | n/as
2024-07-01 07:56 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150049 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 07:57 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150155 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 354 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 07:58 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150218 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-01 08:00 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150152 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 08:12 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405080729 | Pass | Fail | [11 empty] | screen won't turn on | n/a
2024-07-01 08:17 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150644 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 08:18 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150154 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-01 08:20 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150050 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 08:20 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150055 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 08:22 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150575 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-01 09:17 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405080695 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 09:19 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150114 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 09:19 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405080683 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 09:30 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080261 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 09:30 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150060 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 09:30 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405080727 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-01 09:33 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150589 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 09:34 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150594 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 09:34 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150593 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 09:38 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150555 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-01 09:52 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080783 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 370 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 09:54 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150554 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-01 09:57 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150550 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 09:58 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150510 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 09:58 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405080788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 10:09 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150532 | Pass | Fail | [11 empty] | screen slow response | n/a
2024-07-01 10:12 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150556 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 363 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 10:12 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150624 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-01 10:13 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150143 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 10:34 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405080742 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 10:34 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405080765 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 10:35 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080711 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 403 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 10:36 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405080704 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-01 10:37 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150211 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 351 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 10:38 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2404300425 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 10:43 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080752 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 355 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 10:55 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 10:56 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150217 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 11:01 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2404300201 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.176 | Pass | Pass |  | N/A
2024-07-01 11:28 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405150012 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 11:30 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405150615 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 413 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 11:31 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405150612 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-01 11:34 | stmstation3@avive.life | W1288 | Agree | Liza Ordillo | 2405150001 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 11:37 | stmstation2@avive.life | W1288 | Agree | Liza Ordillo | 2405150505 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 11:38 | stmstation1@avive.life | W1288 | Agree | Liza Ordillo | 2405150147 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-01 11:46 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405150614 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 424 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 11:47 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405150503 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 11:50 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405150158 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 11:56 | stmstation3@avive.life | W1288 | Agree | Liza Ordillo | 2405150215 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 11:58 | stmstation2@avive.life | W1288 | Agree | Liza Ordillo | 2405150500 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-01 11:59 | stmstation1@avive.life | W1288 | Agree | Liza Ordillo | 2405150501 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 12:05 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405080692 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 418 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 12:06 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405080669 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-01 12:07 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405080666 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 405 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 12:18 | stmstation3@avive.life | W1288 | Agree | Liza Ordillo | 2405080671 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 12:23 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405080687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 407 | Pass | Pass | 24.19 | Fail |  | high current on sleep mode also tried to different fixture | n/a
2024-07-01 12:27 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405150524 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 12:29 | stmstation2@avive.life | W1288 | Agree | Liza Ordillo | 2405080665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-01 12:35 | stmstation1@avive.life | W1288 | Agree | Liza Ordillo | 2405150535 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-01 13:30 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405080676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 13:33 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405080686 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 13:37 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405080605 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 24.4 | Fail |  | high current on sleep mode tried on different fixture | n/a
2024-07-01 14:31 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405150145 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-01 14:31 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405150528 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 411 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-01 14:32 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405150661 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-01 14:33 | stmstation1@avive.life | W1288 | Agree | Liza Ordillo | 2405080710 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 374 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 14:34 | stmstation2@avive.life | W1288 | Agree | Liza Ordillo | 2405080761 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 397 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-01 14:35 | stmstation3@avive.life | W1288 | Agree | Liza Ordillo | 2405150006 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-02 06:32 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150101 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 402 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-02 06:37 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080681 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-02 06:44 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150499 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-02 06:45 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080738 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-02 06:45 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150216 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-02 07:05 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150133 | Pass | Fail | [11 empty] | Screen won't scroll | n/a
2024-07-02 07:09 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150531 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 390 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-02 07:09 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150140 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 389 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-02 07:16 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150115 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 377 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-02 07:17 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150141 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-02 07:31 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150509 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 369 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-02 07:32 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405080764 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 450 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-02 07:36 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150067 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-02 07:36 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150625 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 400 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-02 07:36 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405080468 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 364 | Pass | Pass | 0.166 | Pass | Pass |  | n/a
2024-07-02 07:43 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150627 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-02 07:45 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2404300513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 357 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-02 07:45 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405080787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-02 07:55 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150110 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 376 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-02 07:56 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150106 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 404 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-02 07:57 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150007 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-02 07:57 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150004 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 394 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-02 07:58 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150011 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-02 08:09 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150626 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-02 08:09 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405080637 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-02 08:10 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150617 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 356 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-02 08:24 | stmstation6@avive.life | W1288 | Agree | Leonora Cruz | 2405150562 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 409 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-02 08:27 | stmstation5@avive.life | W1288 | Agree | Leonora Cruz | 2405150135 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-02 08:28 | stmstation4@avive.life | W1288 | Agree | Leonora Cruz | 2405150486 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-02 08:48 | stmstation2@avive.life | W1288 | Agree | Leonora Cruz | 2405150208 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 398 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-02 08:49 | stmstation3@avive.life | W1288 | Agree | Leonora Cruz | 2405150102 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 379 | Pass | Pass | 0.178 | Pass | Pass |  | N/A
2024-07-02 09:03 | stmstation1@avive.life | W1288 | Agree | Leonora Cruz | 2405150580 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-02 11:26 | stmstation6@avive.life | W1288 | Agree | Liza Ordillo | 2405080729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 420 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-02 11:27 | stmstation5@avive.life | W1288 | Agree | Liza Ordillo | 2405080633 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-02 11:28 | stmstation4@avive.life | W1288 | Agree | Liza Ordillo | 2405150133 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 384 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-02 11:30 | stmstation1@avive.life | W1288 | Agree | Liza Ordillo | 2405150540 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-02 11:31 | stmstation2@avive.life | W1288 | Agree | Liza Ordillo | 2405150532 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 421 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 05:37 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2404120748 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 433 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-03 05:42 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2404120689 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 05:49 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2404120716 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 05:52 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2405150642 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 05:57 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2404120722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-03 06:00 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2405150514 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-03 06:02 | stmstation4@avive.life | W1291 | Agree | Christina Cruz | 2404300205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-03 06:06 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2404120714 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 503 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-03 06:07 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2405150516 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 06:09 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2404120767 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-03 06:14 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2404120730 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-03 06:16 | stmstation4@avive.life | W1291 | Agree | Christina Cruz | 2404120731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-03 06:18 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2404120781 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-03 06:20 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2405150646 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-03 06:22 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2405080734 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 513 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-03 06:24 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2405150181 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 06:26 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2404120723 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-03 06:28 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2404120691 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 06:33 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2404120726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 06:35 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2405150136 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 06:37 | stmstation4@avive.life | W1291 | Agree | Christina Cruz | 2404120761 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-03 06:39 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2405150062 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-03 06:40 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2404120710 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-03 06:45 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2405080733 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-03 06:47 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2405150139 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 06:48 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2404120786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 497 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-03 06:50 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2404120712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 06:53 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2404120782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-03 06:54 | stmstation4@avive.life | W1291 | Agree | Christina Cruz | 2404120690 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 503 | Pass | Pass | 0.19 | Pass | Pass |  | n/a
2024-07-03 06:57 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2405150643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 06:59 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2404120788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 07:00 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2404120643 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 07:03 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2405150079 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 395 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-03 08:35 | stmstation5@avive.life | W1291 | Agree | Valeria Manzo | 2405080492 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-03 08:35 | stmstation4@avive.life | W1291 | Agree | Valeria Manzo | 2405080716 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 392 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-03 08:35 | stmstation6@avive.life | W1291 | Agree | Valeria Manzo | 2405080265 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-07-03 08:36 | stmstation3@avive.life | W1291 | Agree | Valeria Manzo | 2405080478 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 422 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 08:36 | stmstation2@avive.life | W1291 | Agree | Valeria Manzo | 2405150066 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 08:36 | stmstation1@avive.life | W1291 | Agree | Valeria Manzo | 2405080474 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-03 08:46 | stmstation2@avive.life | W1291 | Agree | Valeria Manzo | 2404120727 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 08:46 | stmstation1@avive.life | W1291 | Agree | Valeria Manzo | 2405080481 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-03 08:46 | stmstation4@avive.life | W1291 | Agree | Valeria Manzo | 2405080791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-03 08:47 | stmstation5@avive.life | W1291 | Agree | Valeria Manzo | 2405150053 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-03 08:47 | stmstation6@avive.life | W1291 | Agree | Valeria Manzo | 2404120737 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-03 08:56 | stmstation2@avive.life | W1291 | Agree | Valeria Manzo | 2405150137 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 09:01 | stmstation5@avive.life | W1291 | Agree | Valeria Manzo | 2405080760 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-07-03 09:06 | stmstation5@avive.life | W1291 | Agree | Valeria Manzo | 2405080793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-03 09:07 | stmstation4@avive.life | W1291 | Agree | Valeria Manzo | 2405080662 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-03 09:08 | stmstation2@avive.life | W1291 | Agree | Valeria Manzo | 2404120729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-03 09:08 | stmstation1@avive.life | W1291 | Agree | Valeria Manzo | 2405150052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-03 09:08 | stmstation6@avive.life | W1291 | Agree | Valeria Manzo | 2405150497 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 09:09 | stmstation3@avive.life | W1291 | Agree | Valeria Manzo | 2404120711 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 09:17 | stmstation4@avive.life | W1291 | Agree | Valeria Manzo | 2405080672 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 09:18 | stmstation5@avive.life | W1291 | Agree | Valeria Manzo | 2405150590 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 09:18 | stmstation6@avive.life | W1291 | Agree | Valeria Manzo | 2405150586 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-03 09:20 | stmstation2@avive.life | W1291 | Agree | Valeria Manzo | 2405080673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-03 09:20 | stmstation1@avive.life | W1291 | Agree | Valeria Manzo | 2404120717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-03 09:21 | stmstation3@avive.life | W1291 | Agree | Valeria Manzo | 2405150498 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-03 09:25 | stmstation4@avive.life | W1291 | Agree | Valeria Manzo | 2404120659 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-03 09:26 | stmstation5@avive.life | W1291 | Agree | Valeria Manzo | 2404120725 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-03 09:28 | stmstation1@avive.life | W1291 | Agree | Valeria Manzo | 2405150587 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-03 09:29 | stmstation2@avive.life | W1291 | Agree | Valeria Manzo | 2404120728 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 23.64 | Fail |  | high current when asleep, tried 3 different fixtures | n/a
2024-07-03 09:31 | stmstation6@avive.life | W1291 | Agree | Valeria Manzo | 2405080719 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.199 | Pass | Pass |  | n/a
2024-07-03 09:44 | stmstation5@avive.life | W1291 | Agree | Valeria Manzo | 2404120780 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 372 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-03 09:44 | stmstation6@avive.life | W1291 | Agree | Valeria Manzo | 2404120715 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-03 10:50 | stmstation6@avive.life | W1291 | Agree | Liza Ordillo | 2405080795 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-03 10:52 | stmstation5@avive.life | W1291 | Agree | Liza Ordillo | 2405150069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-03 10:53 | stmstation4@avive.life | W1291 | Agree | Liza Ordillo | 2404120721 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 10:58 | stmstation1@avive.life | W1291 | Agree | Liza Ordillo | 2404120694 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 436 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-03 10:59 | stmstation2@avive.life | W1291 | Agree | Liza Ordillo | 2405150087 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-03 11:00 | stmstation3@avive.life | W1291 | Agree | Liza Ordillo | 2405150094 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-03 11:04 | stmstation6@avive.life | W1291 | Agree | Liza Ordillo | 2404120768 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-03 11:05 | stmstation4@avive.life | W1291 | Agree | Liza Ordillo | 2404120734 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-03 11:09 | stmstation2@avive.life | W1291 | Agree | Liza Ordillo | 2404120736 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-03 11:10 | stmstation1@avive.life | W1291 | Agree | Liza Ordillo | 2405150098 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-03 11:18 | stmstation3@avive.life | W1291 | Agree | Liza Ordillo | 2405150074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-03 11:19 | stmstation5@avive.life | W1291 | Agree | Liza Ordillo | 2404120720 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-08 07:26 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2404120790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 23.45 | Fail |  | Hi reading during sleep | FAIL
2024-07-08 07:38 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2405150071 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 07:41 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2405150178 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 387 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 07:45 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2405150064 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-08 07:47 | stmstation4@avive.life | W1291 | Agree | Christina Cruz | 2405150057 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-08 07:49 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2405150172 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 24.19 | Fail |  | hi current during sleep | FAIL
2024-07-08 07:54 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2404120789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 07:55 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2405150177 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 07:59 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2404120718 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-08 08:05 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2404300191 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 399 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 08:09 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2404120688 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 23.45 | Fail |  | Hi current during sleep | FAIL
2024-07-08 08:19 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2404120764 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 08:21 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2404120765 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-08 08:31 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2405150065 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 08:33 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2405150606 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-08 08:36 | stmstation4@avive.life | W1291 | Agree | Christina Cruz | 2405150072 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.199 | Pass | Pass |  | n/a
2024-07-08 08:39 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2404300199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 383 | Pass | Pass | 0.187 | Pass | Pass |  | n/a
2024-07-08 08:41 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2404120713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 08:47 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2404120740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 08:49 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2405150176 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 08:52 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2404120762 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-08 08:54 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2404300258 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-08 08:57 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2404120738 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 08:59 | stmstation4@avive.life | W1291 | Agree | Christina Cruz | 2404120709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 09:02 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2405150056 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 09:03 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2404120752 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 09:06 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2404120732 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 408 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 09:09 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2405080682 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-08 09:26 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2405080740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 09:30 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2404120708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 09:32 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2404120747 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 09:43 | stmstation4@avive.life | W1291 | Agree | Christina Cruz | 2405080684 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 09:45 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2404120707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 09:48 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2404120687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 09:52 | stmstation2@avive.life | W1291 | Agree | Christina Cruz | 2405080685 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 391 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 09:52 | stmstation3@avive.life | W1291 | Agree | Christina Cruz | 2404120719 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-07-08 09:57 | stmstation1@avive.life | W1291 | Agree | Christina Cruz | 2405150090 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 10:02 | stmstation6@avive.life | W1291 | Agree | Christina Cruz | 2404120776 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-08 10:03 | stmstation5@avive.life | W1291 | Agree | Christina Cruz | 2404120441 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 24.15 | Fail |  | Hi current during sleep | FAIL
2024-07-08 11:03 | stmstation5@avive.life | W1291 | Agree | Liza Ordillo | 2405150609 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 11:05 | stmstation4@avive.life | W1291 | Agree | Liza Ordillo | 2405150171 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 11:06 | stmstation6@avive.life | W1291 | Agree | Liza Ordillo | 2405150080 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-08 11:15 | stmstation2@avive.life | W1291 | Agree | Liza Ordillo | 2405150097 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 11:16 | stmstation3@avive.life | W1291 | Agree | Liza Ordillo | 2405150173 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 11:18 | stmstation4@avive.life | W1291 | Agree | Liza Ordillo | 2405150599 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 11:25 | stmstation5@avive.life | W1291 | Agree | Liza Ordillo | 2405150607 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.203 | Fail |  | high current tried on diferrent fixture | n/a
2024-07-08 11:31 | stmstation3@avive.life | W1291 | Agree | Liza Ordillo | 2405150182 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 458 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 11:32 | stmstation2@avive.life | W1291 | Agree | Liza Ordillo | 2405150081 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 11:33 | stmstation1@avive.life | W1291 | Agree | Liza Ordillo | 2405150174 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-08 11:40 | stmstation5@avive.life | W1291 | Agree | Liza Ordillo | 2405150512 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 11:42 | stmstation4@avive.life | W1291 | Agree | Liza Ordillo | 2405150095 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-08 11:47 | stmstation2@avive.life | W1291 | Agree | Liza Ordillo | 2405150573 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-08 11:49 | stmstation1@avive.life | W1291 | Agree | Liza Ordillo | 2405150083 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-08 11:51 | stmstation6@avive.life | W1291 | Agree | Liza Ordillo | 2405150144 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 12:00 | stmstation3@avive.life | W1291 | Agree | Liza Ordillo | 2405150520 | Pass | Fail | [11 empty] | already changed with a new screen doesn't turn on | n/a
2024-07-08 12:03 | stmstation4@avive.life | W1291 | Agree | Liza Ordillo | 2405150519 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 12:04 | stmstation5@avive.life | W1291 | Agree | Liza Ordillo | 2405150089 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 12:08 | stmstation6@avive.life | W1291 | Agree | Liza Ordillo | 2405150515 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 380 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-08 12:11 | stmstation1@avive.life | W1291 | Agree | Liza Ordillo | 2405150091 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 12:11 | stmstation2@avive.life | W1291 | Agree | Liza Ordillo | 2405150517 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 12:13 | stmstation5@avive.life | W1291 | Agree | Liza Ordillo | 2405150521 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-08 12:15 | stmstation3@avive.life | W1291 | Agree | Liza Ordillo | 2405150518 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 13:22 | stmstation4@avive.life | W1291 | Agree | Leonora Cruz | 2404120759 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 13:22 | stmstation6@avive.life | W1291 | Agree | Leonora Cruz | 2404120692 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-08 13:23 | stmstation5@avive.life | W1291 | Agree | Leonora Cruz | 2404120770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 415 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 13:24 | stmstation1@avive.life | W1291 | Agree | Leonora Cruz | 2404300052 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.181 | Pass | Pass |  | N/A
2024-07-08 13:33 | stmstation2@avive.life | W1291 | Agree | Leonora Cruz | 2404120773 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-08 13:33 | stmstation1@avive.life | W1291 | Agree | Leonora Cruz | 2404120743 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 13:34 | stmstation3@avive.life | W1291 | Agree | Leonora Cruz | 2404120769 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 13:49 | stmstation2@avive.life | W1291 | Agree | Leonora Cruz | 2404120763 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 428 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 13:50 | stmstation1@avive.life | W1291 | Agree | Leonora Cruz | 2404300193 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 13:50 | stmstation6@avive.life | W1291 | Agree | Leonora Cruz | 2404120779 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 13:51 | stmstation5@avive.life | W1291 | Agree | Leonora Cruz | 2404300074 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 13:52 | stmstation4@avive.life | W1291 | Agree | Leonora Cruz | 2404120754 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 14:09 | stmstation6@avive.life | W1291 | Agree | Leonora Cruz | 2404300192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 14:09 | stmstation5@avive.life | W1291 | Agree | Leonora Cruz | 2404120756 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 435 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 14:10 | stmstation4@avive.life | W1291 | Agree | Leonora Cruz | 2404120758 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 14:10 | stmstation3@avive.life | W1291 | Agree | Leonora Cruz | 2404300257 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-08 14:11 | stmstation2@avive.life | W1291 | Agree | Leonora Cruz | 2404300068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 14:16 | stmstation1@avive.life | W1291 | Agree | Leonora Cruz | 2404300069 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 14:26 | stmstation2@avive.life | W1291 | Agree | Leonora Cruz | 2404300255 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-07-08 14:27 | stmstation1@avive.life | W1291 | Agree | Leonora Cruz | 2404120777 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 14:29 | stmstation4@avive.life | W1291 | Agree | Leonora Cruz | 2404300224 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-08 14:29 | stmstation5@avive.life | W1291 | Agree | Leonora Cruz | 2404120778 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.192 | Pass | Pass |  | n/a
2024-07-08 14:33 | stmstation3@avive.life | W1291 | Agree | Leonora Cruz | 2404300206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.181 | Pass | Pass |  | N/A
2024-07-08 14:44 | stmstation1@avive.life | W1291 | Agree | Leonora Cruz | 2404120799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.193 | Pass | Pass |  | n/a
2024-07-08 14:44 | stmstation2@avive.life | W1291 | Agree | Leonora Cruz | 2404120751 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 452 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 14:51 | stmstation5@avive.life | W1291 | Agree | Leonora Cruz | 2404300190 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 14:51 | stmstation4@avive.life | W1291 | Agree | Leonora Cruz | 2404120784 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-07-08 14:52 | stmstation6@avive.life | W1291 | Agree | Leonora Cruz | 2404120787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 14:54 | stmstation3@avive.life | W1291 | Agree | Leonora Cruz | 2404300223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 15:14 | stmstation6@avive.life | W1291 | Agree | Precious Ungco | 2405080472 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-08 15:14 | stmstation5@avive.life | W1291 | Agree | Precious Ungco | 2405080381 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-08 15:15 | stmstation4@avive.life | W1291 | Agree | Precious Ungco | 2405080361 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 15:16 | stmstation1@avive.life | W1291 | Agree | Precious Ungco | 2405080356 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-08 15:16 | stmstation2@avive.life | W1291 | Agree | Precious Ungco | 2405080482 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-08 15:17 | stmstation3@avive.life | W1291 | Agree | Precious Ungco | 2405080484 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-08 15:30 | stmstation3@avive.life | W1291 | Agree | Precious Ungco | 2405080425 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-08 15:31 | stmstation2@avive.life | W1291 | Agree | Precious Ungco | 2405080285 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 15:31 | stmstation1@avive.life | W1291 | Agree | Precious Ungco | 2405080319 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 440 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 15:32 | stmstation4@avive.life | W1291 | Agree | Precious Ungco | 2405080337 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 15:32 | stmstation5@avive.life | W1291 | Agree | Precious Ungco | 2405080291 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-07-08 15:38 | stmstation6@avive.life | W1291 | Agree | Precious Ungco | 2405080404 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 24.9 | Fail |  | high current | n/a
2024-07-08 15:50 | stmstation3@avive.life | W1291 | Agree | Precious Ungco | 2405150513 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 15:51 | stmstation2@avive.life | W1291 | Agree | Precious Ungco | 2405150591 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 15:51 | stmstation1@avive.life | W1291 | Agree | Precious Ungco | 2405150487 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-08 15:52 | stmstation4@avive.life | W1291 | Agree | Precious Ungco | 2405150179 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 15:52 | stmstation5@avive.life | W1291 | Agree | Precious Ungco | 2405150511 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-08 15:57 | stmstation6@avive.life | W1291 | Agree | Precious Ungco | 2404120693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 23.75 | Fail |  | high current | n/a
2024-07-08 16:06 | stmstation5@avive.life | W1291 | Agree | Precious Ungco | 2404120749 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 16:06 | stmstation4@avive.life | W1291 | Agree | Precious Ungco | 2404120794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 16:09 | stmstation3@avive.life | W1291 | Agree | Precious Ungco | 2404120750 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 16:10 | stmstation2@avive.life | W1291 | Agree | Precious Ungco | 2405150156 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-07-08 16:10 | stmstation1@avive.life | W1291 | Agree | Precious Ungco | 2405150157 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-08 16:12 | stmstation6@avive.life | W1291 | Agree | Precious Ungco | 2404120757 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 22.7 | Fail |  | high current | n/a
2024-07-08 16:25 | stmstation5@avive.life | W1291 | Agree | Precious Ungco | 2404120755 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 16:26 | stmstation4@avive.life | W1291 | Agree | Precious Ungco | 2405150588 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 16:27 | stmstation6@avive.life | W1291 | Agree | Precious Ungco | 2404120753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 22.45 | Fail |  | high current | n/a
2024-07-08 16:28 | stmstation1@avive.life | W1291 | Agree | Precious Ungco | 2404120774 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-08 16:29 | stmstation2@avive.life | W1291 | Agree | Precious Ungco | 2405150068 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-07-08 16:29 | stmstation3@avive.life | W1291 | Agree | Precious Ungco | 2404300251 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-08 16:53 | stmstation3@avive.life | W1291 | Agree | Precious Ungco | 2405150086 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-08 17:00 | stmstation2@avive.life | W1291 | Agree | Precious Ungco | 2405150180 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-09 11:01 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150799 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-09 11:02 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150036 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 11:03 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150722 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-09 11:04 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150801 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 11:05 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150861 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-09 11:17 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150798 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-09 11:22 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150039 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 11:22 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150840 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-09 11:23 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150862 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-09 11:24 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150184 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-09 11:24 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150725 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 430 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-09 11:41 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150756 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.179 | Pass | Pass |  | GPS
2024-07-09 11:42 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150227 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-09 11:45 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150768 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-09 11:46 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150230 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 11:51 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150727 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-09 11:54 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150221 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 11:55 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150729 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.187 | Pass | Pass |  | n/a
2024-07-09 11:59 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150804 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-09 12:13 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150835 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-09 12:13 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150802 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 350 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-09 12:14 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150721 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-09 12:21 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150743 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 500 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-09 12:22 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150420 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 12:28 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150416 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-09 12:28 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150425 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-09 12:29 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150789 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-09 12:31 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150730 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-09 12:35 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150781 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-09 12:39 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150838 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-09 12:46 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150674 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-09 12:52 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150770 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-09 12:53 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150424 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 494 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-09 12:54 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150790 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-09 12:56 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150421 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 378 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-09 13:02 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150726 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 508 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-07-09 13:27 | stmstation5@avive.life | W1292 | Agree | Valeria Manzo | 2405150848 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 13:27 | stmstation6@avive.life | W1292 | Agree | Valeria Manzo | 2405150222 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 13:33 | stmstation4@avive.life | W1292 | Agree | Valeria Manzo | 2405150847 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 13:34 | stmstation1@avive.life | W1292 | Agree | Valeria Manzo | 2405150846 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-09 13:37 | stmstation2@avive.life | W1292 | Agree | Valeria Manzo | 2405150807 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-09 13:37 | stmstation3@avive.life | W1292 | Agree | Valeria Manzo | 2405150766 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-09 13:48 | stmstation6@avive.life | W1292 | Agree | Valeria Manzo | 2405150698 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-09 13:48 | stmstation5@avive.life | W1292 | Agree | Valeria Manzo | 2405150850 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-09 13:48 | stmstation4@avive.life | W1292 | Agree | Valeria Manzo | 2405150849 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-09 13:48 | stmstation1@avive.life | W1292 | Agree | Valeria Manzo | 2405150371 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-09 13:48 | stmstation2@avive.life | W1292 | Agree | Valeria Manzo | 2405150376 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 13:48 | stmstation3@avive.life | W1292 | Agree | Valeria Manzo | 2405150843 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.191 | Pass | Pass |  | n/a
2024-07-09 14:00 | stmstation3@avive.life | W1292 | Agree | Valeria Manzo | 2405150877 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-09 14:00 | stmstation2@avive.life | W1292 | Agree | Valeria Manzo | 2405150805 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-09 14:01 | stmstation1@avive.life | W1292 | Agree | Valeria Manzo | 2405150844 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-09 14:01 | stmstation4@avive.life | W1292 | Agree | Valeria Manzo | 2405150874 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-09 14:01 | stmstation6@avive.life | W1292 | Agree | Valeria Manzo | 2405150728 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-09 14:01 | stmstation5@avive.life | W1292 | Agree | Valeria Manzo | 2405150705 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-10 09:48 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150677 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-10 09:49 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150742 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-10 09:50 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150741 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-10 09:50 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150682 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 09:51 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150679 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 09:51 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150683 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-10 10:02 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150782 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 10:03 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150760 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 10:03 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150759 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-10 10:04 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150755 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-10 10:05 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150758 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.197 | Pass | Pass |  | n/a
2024-07-10 10:05 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150757 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-10 10:16 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150769 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 10:17 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150763 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 10:17 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150771 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-10 10:18 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150598 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 10:18 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150767 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-10 10:20 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150426 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 10:34 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150888 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 10:34 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150422 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 10:34 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150415 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.191 | Pass | Pass |  | n/a
2024-07-10 10:35 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150868 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-10 10:36 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150878 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 454 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-10 10:36 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150855 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-10 10:48 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150800 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-07-10 10:49 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150695 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-10 10:50 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150684 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 10:50 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150746 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 10:51 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150747 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-10 10:51 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150739 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.174 | Pass | Pass |  | n/a
2024-07-10 11:08 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150189 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-10 11:08 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150693 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 11:08 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150830 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 11:08 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150754 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 11:09 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150731 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-10 11:09 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150750 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 11:42 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150724 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 11:43 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150373 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-07-10 11:44 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150740 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-10 11:45 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150894 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 11:46 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150776 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 502 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 11:49 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150900 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-10 11:59 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150735 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-10 12:01 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150897 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 12:01 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150046 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-10 12:02 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150038 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 12:08 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150374 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-10 12:16 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150896 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-07-10 12:19 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150372 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-10 12:21 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150773 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 12:21 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150678 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 12:21 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150732 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 449 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-07-10 12:29 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150377 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-10 12:34 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150845 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-10 12:37 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150736 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-10 12:38 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150774 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-07-10 12:38 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150872 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-10 12:39 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150419 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-10 12:49 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150734 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.192 | Pass | Pass |  | n/a
2024-07-10 12:50 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150777 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-10 12:50 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150775 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-10 12:53 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150780 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 500 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-10 13:02 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150851 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 13:02 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150378 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-10 13:02 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150866 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-10 13:13 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150370 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 13:14 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150025 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.195 | Pass | Pass |  | n/a
2024-07-10 13:15 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150794 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-10 13:15 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150037 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-10 13:16 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150733 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-10 13:20 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150899 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 511 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-10 13:22 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150898 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-11 12:58 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150040 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 12:58 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150045 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 466 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-11 12:59 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150772 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-11 13:00 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150676 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 13:01 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150884 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-11 13:02 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150882 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-11 13:13 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150870 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 13:14 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150827 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-11 13:14 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150839 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 491 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-07-11 13:15 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150826 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-11 13:16 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150879 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-11 13:17 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150820 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 13:27 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150818 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 13:27 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150824 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.186 | Pass | Pass |  | n/a
2024-07-11 13:28 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150834 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 13:29 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150706 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-11 13:29 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150714 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.186 | Pass | Pass |  | n/a
2024-07-11 13:30 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150812 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 13:52 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150226 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-11 13:52 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150791 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-11 13:55 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150186 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.198 | Pass | Pass |  | n/a
2024-07-11 13:56 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150685 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 13:56 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150207 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 14:04 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150204 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 14:06 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150027 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-11 14:08 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150223 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 495 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-11 14:11 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150687 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.187 | Pass | Pass |  | n/a
2024-07-11 14:21 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150220 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 14:22 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150185 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-11 14:25 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150034 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-11 14:25 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150785 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-11 14:26 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150202 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 506 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 14:39 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150787 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-11 14:41 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150028 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-11 14:42 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150035 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 501 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 14:44 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150026 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-11 14:45 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150738 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-07-11 14:53 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150032 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-11 15:16 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150885 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 434 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 15:16 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150883 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 388 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 15:17 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150031 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 396 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-11 15:18 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150673 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-11 15:19 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150672 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 456 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 15:19 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150819 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-11 15:30 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150660 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 15:30 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150663 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 15:30 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150681 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 15:31 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150871 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 15:32 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150822 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 15:32 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150823 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-11 15:45 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150809 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 15:46 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150796 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-11 15:47 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150797 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 15:47 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150831 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 451 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 15:48 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150814 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-11 15:48 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150833 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-11 16:02 | stmstation4@avive.life | W1292 | Agree | Precious Ungco | 2405150821 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 462 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 16:03 | stmstation5@avive.life | W1292 | Agree | Precious Ungco | 2405150810 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 447 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 16:03 | stmstation6@avive.life | W1292 | Agree | Precious Ungco | 2405150597 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-11 16:04 | stmstation3@avive.life | W1292 | Agree | Precious Ungco | 2405150828 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-11 16:04 | stmstation2@avive.life | W1292 | Agree | Precious Ungco | 2405150042 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-11 16:05 | stmstation1@avive.life | W1292 | Agree | Precious Ungco | 2405150825 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 453 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-12 10:04 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150203 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-12 10:05 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150691 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 375 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-12 10:06 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150890 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 430 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-12 10:08 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150187 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-12 10:08 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150195 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-12 10:19 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150668 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 493 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-12 10:20 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150059 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 463 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-12 10:21 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150892 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 513 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-12 10:24 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150719 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-12 10:24 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150859 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 430 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-12 10:35 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150720 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-12 10:42 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150201 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.265 | Fail |  | value out of range/nogps | n/a
2024-07-12 10:45 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.187 | Pass | Fail | GPS | GPS
2024-07-12 10:52 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150793 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.185 | Pass | Pass |  | n/a
2024-07-12 10:53 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150889 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 504 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-12 10:54 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150891 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-12 11:01 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150783 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-12 11:01 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150686 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-12 11:02 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150694 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-12 11:02 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150716 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-12 11:12 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150196 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-12 11:12 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150893 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-12 11:13 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150710 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 455 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-12 11:15 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150651 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-12 11:23 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150205 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-12 11:23 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150689 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-12 11:30 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150058 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 352 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-12 11:30 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150656 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 490 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-07-12 11:31 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150188 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-12 11:48 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150786 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 371 | Pass | Pass | 0.183 | Pass | Pass |  | n/a
2024-07-12 11:49 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150688 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-12 11:50 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150030 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-12 12:01 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150033 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-12 12:01 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150753 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 461 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-12 12:01 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150206 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-12 12:05 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150788 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-12 12:06 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150690 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 509 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-12 12:09 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150680 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.189 | Pass | Pass |  | n/a
2024-07-12 12:14 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150675 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-12 12:17 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150029 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-07-12 12:21 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150194 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-12 12:30 | stmstation5@avive.life | W1292 | Agree | Leonora Cruz | 2405150737 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 499 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-12 12:30 | stmstation6@avive.life | W1292 | Agree | Leonora Cruz | 2405150200 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 431 | Pass | Pass | 0.188 | Pass | Pass |  | n/a
2024-07-12 12:32 | stmstation2@avive.life | W1292 | Agree | Leonora Cruz | 2405150224 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-12 12:33 | stmstation1@avive.life | W1292 | Agree | Leonora Cruz | 2405150198 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.182 | Pass | Pass |  | n/a
2024-07-12 12:35 | stmstation3@avive.life | W1292 | Agree | Leonora Cruz | 2405150784 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 485 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-12 12:36 | stmstation4@avive.life | W1292 | Agree | Leonora Cruz | 2405150795 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-15 06:30 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150199 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 489 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-15 06:31 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150228 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 483 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-15 06:35 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150163 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-15 06:37 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150044 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 446 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-15 06:46 | stmstation4@avive.life | W1292 | Agree | Liza Ordillo | 2405150418 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-15 06:49 | stmstation3@avive.life | W1292 | Agree | Liza Ordillo | 2405150837 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 381 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-15 06:54 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150860 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 06:56 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150043 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 07:03 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150708 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 486 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-15 07:04 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150863 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 492 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-15 07:07 | stmstation4@avive.life | W1292 | Agree | Liza Ordillo | 2405150858 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 480 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-15 07:14 | stmstation3@avive.life | W1292 | Agree | Liza Ordillo | 2405150699 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 457 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 07:15 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150857 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 444 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-15 07:17 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150697 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 07:22 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150875 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-15 07:23 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150669 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-15 07:29 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150854 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 468 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-15 07:30 | stmstation3@avive.life | W1292 | Agree | Liza Ordillo | 2405150713 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-15 07:31 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150806 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 401 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 07:35 | stmstation4@avive.life | W1292 | Agree | Liza Ordillo | 2405150654 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 465 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-15 07:37 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150715 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 464 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 07:40 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150869 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 496 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 07:48 | stmstation3@avive.life | W1292 | Agree | Liza Ordillo | 2405150704 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-15 07:49 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150749 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 367 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-15 07:51 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150659 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 365 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-15 07:52 | stmstation4@avive.life | W1292 | Agree | Liza Ordillo | 2405150867 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 07:55 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150701 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 477 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 07:56 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150647 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 474 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 08:11 | stmstation3@avive.life | W1292 | Agree | Liza Ordillo | 2405150808 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-15 08:12 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150707 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-15 08:14 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150764 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-15 08:16 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150665 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 476 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-15 08:17 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150712 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 487 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-15 08:18 | stmstation4@avive.life | W1292 | Agree | Liza Ordillo | 2405150887 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-15 08:26 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150864 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 448 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-15 08:27 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150829 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-15 08:28 | stmstation3@avive.life | W1292 | Agree | Liza Ordillo | 2405150702 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 445 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 08:31 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150811 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 470 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-15 08:32 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150865 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-15 08:33 | stmstation4@avive.life | W1292 | Agree | Liza Ordillo | 2405150696 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 08:43 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150703 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-15 08:44 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150881 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 441 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-15 08:46 | stmstation3@avive.life | W1292 | Agree | Liza Ordillo | 2405150745 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 459 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-15 08:50 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150880 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 478 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 08:51 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150841 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 488 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-15 08:52 | stmstation4@avive.life | W1292 | Agree | Liza Ordillo | 2405150718 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 498 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-15 08:59 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150658 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-15 09:00 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150666 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 438 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 09:00 | stmstation3@avive.life | W1292 | Agree | Liza Ordillo | 2405150700 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 09:07 | stmstation4@avive.life | W1292 | Agree | Liza Ordillo | 2405150803 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 484 | Pass | Pass | 0.187 | Pass | Pass |  | n/a
2024-07-15 09:09 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150886 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 437 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 09:10 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150709 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 09:15 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150671 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 439 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-15 09:17 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150856 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 481 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 09:18 | stmstation3@avive.life | W1292 | Agree | Liza Ordillo | 2405150168 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 471 | Pass | Pass | 0.181 | Pass | Pass |  | n/a
2024-07-15 09:22 | stmstation4@avive.life | W1292 | Agree | Liza Ordillo | 2405150183 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 479 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 09:24 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150842 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 469 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-15 09:25 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150048 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 472 | Pass | Pass | 0.175 | Pass | Pass |  | n/a
2024-07-15 09:31 | stmstation3@avive.life | W1292 | Agree | Liza Ordillo | 2405150423 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 443 | Pass | Pass | 0.179 | Pass | Pass |  | n/a
2024-07-15 09:32 | stmstation2@avive.life | W1292 | Agree | Liza Ordillo | 2405150231 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 460 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 09:33 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150164 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 436 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 09:35 | stmstation4@avive.life | W1292 | Agree | Liza Ordillo | 2405150853 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 475 | Pass | Pass | 0.18 | Pass | Pass |  | n/a
2024-07-15 09:35 | stmstation5@avive.life | W1292 | Agree | Liza Ordillo | 2405150852 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 467 | Pass | Pass | 0.178 | Pass | Pass |  | n/a
2024-07-15 09:36 | stmstation6@avive.life | W1292 | Agree | Liza Ordillo | 2405150192 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 482 | Pass | Pass | 0.176 | Pass | Pass |  | n/a
2024-07-15 09:42 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150159 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 473 | Pass | Pass | 0.177 | Pass | Pass |  | n/a
2024-07-15 10:51 | stmstation1@avive.life | W1292 | Agree | Liza Ordillo | 2405150717 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 360 | Pass | Pass | 0.184 | Pass | Pass |  | n/a
2024-04-23 10:01 | stmstation3@avive.life | W1253 | Agree | Valeria Manzo | 2404020480 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 393 | Pass | Pass | 0.173 | Pass | Pass |  | n/a
2024-04-23 09:30 | stmstation1@avive.life | W1253 | Agree | Valeria Manzo | 2404020776 | Pass | Pass | Pass | Pass | Pass | Pass | Pass | 406 | Pass | Pass | 0.173 | Pass | Pass |  | n/a